# LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T0900055-v1	ADVANCED LIGO	13 March 2009

Rehang of the Quad NP with glass masses but metal wires (January 12-23 2009)

# Mark Barton, Doug Cook, Ken Mailand, Rich Mittleman, Joe O’Dell, Russell Jones, Janeen Romie, Brett Shapiro, Bob Taylor, Phil Willems

Distribution of this document:

DCC

# This is an internal working note

of the LIGO Project.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Table of Contents

1	Introduction	3

1.1	Purpose and Scope	3

1.2	References	3

1.3	Version history	3

2	Main procedure	4

3	Sub-procedures	15

3.1	Aligning the Brunson transit	15

3.2	Using the ergo-arm	17

3.3	Preparing the CP	18

# Introduction

## Purpose and Scope

This document describes the procedure used in going from the all-metal build of the quad noise prototype at LASTI to the build with glass masses but wires. It is written in the style of an instruction manual for ease of recycling the text into future procedure documents but except where noted is actually historical.

## References

T080101-00, D. Bridges et al., Notes on Lower Quad Installation at LASTI.

T080165-00, B. Shapiro, Metal Quad Noise Prototype Balancing and Alignment Procedure.

T060040-06, I. Wilmut, Noise prototype Assembly procedure.

G070359-00, K. Mailand, LASTI Tooling (instructional DVD)

E070292-00, H. Armandula, Optics Cleaning Specification - First Contact™

## Version history

1/18/09: First pre-v1 draft covering activities as far as having swapped out the masses in the reaction chain.

1/28/09: Second pre-v1 draft covering getting the system back in the tank.

2/12/09: Third pre-v1 draft with feedback from Brett on main and ESD procedures. Added Doug Cook as author. Changed installation tooling terminology to match Ken Mailand’s (“rails” to “conveyor” etc). Got doc number: T0900055.

3/13/09: Corrected dates to 09 in version history. Cited E070292-00 for First Contact. Added timing and tools data.

# Main procedure

| Step | What | Where | Time | People | Tools |
| --- | --- | --- | --- | --- | --- |
|  | Lock all masses. Stops need only be done up gently – touching plus 1/8 turn. Putting divots in the metal is unnecessary. (The “2+1” in the “People” column refers to the the fact that all in-tank operations need at least one support person outside.) | in BSC | 10 min | 2+1 | hex keys |
|  | Disconnect electrical wiring to lower structure and masses and tuck ends safely out of the way. | in BSC | 5 min | 2+1 |  |
|  | Lock top stage, lock top mass blades. | in BSC | 5 min | 2+1 | hex keys |
|  | Raise stops under UIM to touch and overload UIM blade tip by 3mm. | in BSC | 5 min | 2+1 | hex keys |
|  | Raise UIM until top to UIM wires are slack. | in BSC | 5 min | 2+1 | hex keys |
|  | Remove UIM pitch mass (to allow access to screws in next step). | in BSC | 5 min | 2+1 | hex keys |
|  | Disconnect wires from top mass to UIM at UIM. | in BSC | 5 min | 2+1 | hex keys |
|  | Remove face earthquake stops. | in BSC | 5 min | 2+1 | hex keys |
|  | Remove X braces and sleeve first, then ¼-20 bolts and spacers at the bottom of the sleeve, then 3/8-16 bolts up top. Loosen all the ¼-20 bolts at the bottom of the sleeve before removing them so that removing the spacers behind them is easier with less risk of dropping them on the optics. | in BSC | 30 min | 3+1 | hex keys |
|  | Install conveyor, five-axis table and elevator (this was done first in reality which made it not quite impossible but very difficult to get the sleeve off). [Rich: conveyor through-holes to the door flange should be opened up so not so tight.] Make sure elevator is installed on 5 axis table such that you have +- 180o of rotation. Wrap electrical cabling for elevator in UHV foil. | in, near BSC | 1 day | 4 | 5AT, conveyor |
|  | Remove side plates on elevator on five-axis tooling and rotate elevator so that it can be slid around structure on the conveyor. When elevator is around structure, rotate so that side labeled “2” is on main chain side. | in BSC | 10 min | 2+1 | wrench |
|  | Insert halves of (lower structure assembly tooling) LSAT around the structure, hooking them on the pins on top of the fully retracted pushers of the five-axis table. To make this easier, start with some half-inch spacers at the bottom of the elevator to lift the LSAT above the pusher pins. These spacers will allow the LSAT to be pushed together before lowering it onto the pusher pins. | in BSC | 5 min | 2+1 | ½” spacers |
|  | Lock the halves of the LSAT together with plates and bolts, except at the bottom level (where this will cause interference with the sides of the elevator). | in BSC | 5 min |  | LSAT side plates |
|  | Reinstall side plates on five-axis table, aligning “1” on corner of first plate to matching “1” on table and “2” on second plate to “2” on other side of table. Bolts are best inserted from centre outwards. Finger-tight is fine. (Doug says: need better tolerancing on holes.) | in BSC | 5 min | 2+1 | 5AT side plates, bolts |
|  | Raise LSAT on pushers until it almost engages with the structure. | in BSC | 5 min | 2+1 |  |
|  | Adjust the various DOFs of five-axis table as appropriate until the top of the LSAT is well-aligned in position and angle with the lower structure. | in BSC | 5 min | 2+1 |  |
|  | Raise pushers until the LSAT has fully engaged with the lower structure and the weight is off the implementation ring. | in BSC | 5 min | 2+1 |  |
|  | Replace face stops on both chains. | in BSC | 10 min | 2+1 | hex keys, face stops |
|  | Unbolt implementation ring from upper structure and remove bolts from upper to lower structure. | in BSC | 10 min | 2+1 | hex keys |
|  | Retract pushers fully, lowering lower structure. | in BSC | 5 min | 2+1 |  |
|  | Insert translational locking pin into 5 axis table. (This must happen after disconnecting the lower structure because the table position may need to be adjusted to get the pin in place.) | in BSC | 5 min | 2+1 | pin |
|  | Unbolt implementation ring from lower structure (because otherwise it won’t get through the door). | in BSC | 10 min | 2+1 | hex keys |
|  | Move lower structure to door of tank on conveyor. | in BSC | 5 min | 2+1 |  |
|  | Remove side plate on five-axis table on the main chain side (“2”). | near BSC | 5 min | 2 | hex keys |
|  | Bring in Genie (avoiding bumping the HEPI) so that the forks are below the uppermost of the side plates on the LSAT. Use spacers (approximately ½” thick) on top of the forks near the tips to allow for sag of the forks under load. | near BSC | 5 min | 4 | Genie, ½” spacers |
|  | Lift the LSAT off the five-axis table onto the cart, install remaining (bottom) plates holding together the two haves and clamp it down. | near BSC | 10 min | 4 | cart, LSAT side plates, dogs |
|  | Cover the LSAT on the cart with a door cloth and wheel to the assembly area. | in transit | 5 min | 4 | door cloth |
|  | Manhandle Genie past solid stack to assembly area. | in transit | 5 min | 4 |  |
|  | Bring in Genie as before and take LSAT off cart using ½” spacers for sag (but don’t set it down yet). | assembly area | 5 min | 4 | Genie, ½” spacers |
|  | Disconnect lowest set of bolts joining the two halves of the structure (these are difficult to remove later). | assembly area | 5 min | 4 | hex keys |
|  | Set LSAT down on foil-covered pallet with ¼” spacers. (Use enough spacers that each half is stable independently.) | assembly area | 5 min | 2 | pallet, (lots of) ¼” spacers |
|  | Remove plates holding halves of LSAT together. | assembly area | 5 min | 2 | hex keys |
|  | Remove remaining bolts holding halves of structure together. | assembly area | 5 min | 2 | hex keys |
|  | Bring in Genie as before except placing ¼” spacers near the base of the forks so as to enable just the reaction chain to be picked up. | assembly area | 5 min | 4 | Genie, ¼” spacers |
|  | Take weight of reaction chain with Genie and remove ¼” spacers under it. Withdraw straight back so as to cleanly disengage pins aligning two halves. | assembly area | 5 min | 4 | Genie |
|  | Set reaction chain on a second foil-covered pallet (or piece of foil on the floor), remove spacers and pick up the reaction chain again closer to the centre of the forks. | assembly area | 5 min | 4 | Genie, pallet #2, foil |
|  | Move the reaction chain to the turntable. | assembly area | 5 min | 4 | breadboard on 10” blocks, turntable |
|  | Rotate reaction chain 180°, apply safety chocks under turntable to minimize rocking. | assembly area | 5 min | 4 | chocks |
|  | Disconnect suspension wires between UIM and PM and between PM and (dummy) ERM. | assembly area | 5 min | 2 |  |
|  | Attach ergo-arm adapter plates to PM. | assembly area | 10 min | 2 | adapter plates, bolts, hex-keys: |
|  | Bring in ergo-arm (see Section 3.2), move PM to work table (double-check weight etc). | assembly area | 20 min | 4 | ergo-arm, vacuum pump, hose, reservoir |
|  | Remove dummy ERM with adapter plates and ergo-arm. | assembly area | 30 min | 4 | adapter plates |
|  | Weigh CP and prepare replacement PRM with appropriate mass. | assembly area | 30 min | 2 | CP |
|  | Install replacement PM using adapter plates and ergo-arm. | assembly area | 30 min | 4 | adapter plates, ergo-arm |
|  | Install new wire assemblies incorporating UIM-PRM wires and wire loops for CP. | assembly area | 10 min | 2 | pre-made wire assembly, hex-keys |
|  | Set up the Brunson transit on opposite side from ergo-arm access direction. | assembly area | 30 min | 1 | transit |
|  | Using the transit, clock the PRM until the clamps on either side are at the same height. | assembly area | 10 min | 2 | transit |
|  | Prepare the CP, soldering on the tabs and electrical wires per Section 3.3. (This step could be done at any convenient time prior.) | clean room | 1 day | 1 | tabs, coax, indium, Al buttons, soldering iron, nitrogen blower |
|  | Bring the CP into the assembly area and stand it up in a V-block with the non-ESD face accessible. (This was done manually, which is not ideal. On the other hand it’s not clear that it could have been done with the ergo-arm because once the tabs and wiring are added, the CP can only be laid with the non-ESD face down, and one might worry that the tabs would prevent the suction plate from sealing on the ESD face – this needs to be tested.) | assembly area | 10 min | 2 | V-block |
|  | Arrange a foil pouch (about 9”x9”) near the bottom of the structure on the back (transit) side for receiving and keeping safe/clean the CP electrical wires. | assembly area | 5 min | 1 | foil |
|  | Fit the silica-tips to upper earthquake stops but leave them fully retracted. | assembly area | 10 min | 2 | silica [PTFE] tips |
|  | Put type “D” chocks under the lower pads. | assembly area | 10 min | 2 | type “D” chocks |
|  | Bring the suction plate of the ergo-arm up to the non-ESD face of the CP with the T-piece between the valve and gauge at about 22.5° (so that in a later step it will fit into the corner of the octagon of struts defining the CP’s position in the structure). | assembly area | 5 min | 4 | ergo-arm |
|  | Pick up the CP with the ergo-arm, assigning a person to support the ESD wires throughout the next few steps. | assembly area | 5 min | 4 | ergo-arm, vacuum pump, hose, reservoir |
|  | Bring the ergo-arm with the CP near the structure, and adjust the angle of the base and the height, lateral position, pitch and yaw of the head so that the CP is ready to pass through the structure neatly, with approximately equal clearance on all sides. Lock the castors of the ergo-arm. | assembly area | 5 min | 4 | ergo-arm |
|  | Roll the CP so that the T-piece is at 12 o’clock. (This will put the prisms and ESD wiring tabs in line with corners of the octagon of struts.) | assembly area | 1 min | 4 | ergo-arm |
|  | Pull the two loops of suspension wire gently out of the octagon at the far side of the structure and assign a person to hold them up out of the way for the next few steps. | assembly area | 1 min | 4 | ergo-arm |
|  | Pass the ESD wiring carefully through the structure and place it in the foil pouch. | assembly area | 1 min | 4 | foil pouch , ergo-arm |
|  | Move the CP carefully into the structure until the leading face, the ESD wiring tabs and the prisms have all fully entered the structure volume and are clear of the front-face struts. | assembly area | 5 min | 4 | ergo-arm |
|  | Roll the head back to the original position with the prisms at 3 o’clock and 9 o’clock, and the T-piece at 22.5°. | assembly area | 1 min | 4 | ergo-arm |
|  | Continue to move the CP into the structure either until it has reached its design position with the ESD face level with the far side of the structure or until the T-piece is about to foul on the near side of the structure. | assembly area | 5 min | 4 | ergo-arm |
|  | Thread the pouch of ESD wiring up through the loops of suspension wire and lower the suspension wires against the far side of the structure. | assembly area | 1 min | 4 | ergo-arm |
|  | Stretch the suspension wire loops around and under the mass, being especially careful not to snag them on the ESD wiring tabs. Route them between the pairs of upper earthquake stops on each side. | assembly area | 1 min | 4 | ergo-arm |
|  | Lower the CP gradually until the wires are pulled straight but not taut. If the T-piece on the suction plate fouled against the structure earlier and additional insertion is required, it will tend to become possible at this stage – alternate small downward and inward steps as appropriate. | assembly area | 1 min | 4 | ergo-arm |
|  | Nudge the wire loops into the grooves on the prisms and around the bottom of the CP. A mirror placed facing up on the turntable below the structure will make it easier to check that this has been done correctly. | assembly area | 1 min | 4 | mirror, ergo-arm |
|  | Using the transit, adjust the CP in roll until the prisms are at the same height. | assembly area | 10 min | 4 | transit |
|  | Continue to lower (and if necessary, further insert) the CP until the wires are taut and/or it rests on the bottom earthquake stops. | assembly area | 1 min | 4 | ergo-arm |
|  | Screw in the upper earthquake stops to a 0.5 mm gap. | assembly area | 5 min | 4 |  |
|  | Position a person to restrain the CP in case it swings front to back. Have another person remove the vacuum hose, open the valve on the suction plate and withdraw the ergo-arm. | assembly area | 5 min | 4 | ergo-arm |
|  | Back off the upper earthquake stops, insert Teflon pads and screw them back in to touch. | assembly area | 5 min | 2 | Teflon pads |
|  | Install the face stops and screw them in to a 0.5 mm gap. | assembly area | 5 min | 2 | face stops |
|  | Take out the chocks under the turntable and rotate the reaction chain by 180°. | assembly area | 5 min | 4 |  |
|  | Using the Genie, move the reaction chain off the turntable and place on pallet with approx ¼” spacers underneath. [In fact we used some of the plates for holding the halves of the LSAT together, which was a bad idea because we had to muck around later getting them out. Purpose-made would be better.] | assembly area | 30 min | 4 | ¼” spacers |
|  | Check reaction chain for balance per Section 3.4. | assembly area | 45 min | 2 | triple-hang tooling, hex keys |
|  | Route the ESD wiring up the chain. There need to be wire clamps at top and bottom of the back face of the PRM. | assembly area | 30 min | 2 | ESD wire clamps |
|  | Using the Genie, put the main chain on the turntable with the outside of the LSAT towards the transit and the centre side toward the access side. | assembly area | 20 min | 4 | Genie, turntable |
|  | Remove the UIM-PM and PM-TM wires. | assembly area | 5 min | 2 |  |
|  | Using the ergo-arm and adapter plates, remove the PM and TM and double-check weights etc. | assembly area | 30 min | 4 | balance |
|  | Prepare new PM with appropriate mass. | assembly area | 30 min | 2 | balance |
|  | Install the new PM with adapter plates and ergo-arm. | assembly area | 30 min | 2 | adapter plates, ergo-arm |
|  | Install the wire assembly as for the reaction chain. | assembly area | 10 min | 2 | pre-made wire assembly, hex-keys |
|  | Release overload on UIM blades until PM suspends and check for balance. Restore overload. | assembly area | 10 min | 2 | hex-keys |
|  | Clock PM with transit. | assembly area | 10 min | 2 | transit |
|  | Put silica tips on upper TM earthquake stops but leave well retracted. Install silica tipped face stops. | assembly area | 5 min | 2 | hex-keys |
|  | Put type “A” adjustable pad spacers under lower earthquake pads. | assembly area | 10 min | 2 | type “A” spacers |
|  | Install ring heater, leaving halves as far apart as possible (for the biggest possible gap in the centre). | assembly area | 20 min | 2 | ring heater, bolts, hex-keys |
|  | Pick up TM with ergo-arm from AR side. (Angle of T-piece doesn’t much matter.) | assembly area | 10 min | 4 | ergo-arm, pump, hose reservoir |
|  | Insert TM into structure, being extremely careful not to bump prisms on ring heater, stopping when ergo-arm side is level with face of structure. Position wire loops around TM as it passes through. (Joe says: could try tilting to put prisms at 6 and 12 o’clock.) | assembly area | 10 min | 4 | ergo-arm |
|  | Lower TM until wires are straight but not taut. | assembly area | 5 min | 4 | ergo-arm |
|  | Dress wire loops to pass through prism grooves and then neatly under TM. Again, a mirror is convenient. | assembly area | 5 min | 4 | mirror |
|  | Lower TM onto pads, remove ergo-arm. | assembly area | 5 min | 4 |  |
|  | Put Teflon pads between upper earthquake stops and mass for both PM and TM and tighten stops. | assembly area | 5 min | 2 | Teflon pads, hex keys |
|  | Pick up main chain with tips of Genie forks, using thin spacers (≈1/4”) on tip side to counteract sag. | assembly area | 10 min | 4 | Genie, ¼” spacers |
|  | Bring main chain over to pallet #2. | assembly area | 5 min | 4 | Genie |
|  | Check main chain for balance per Section 3.4. | assembly area | 45 min | 2 | triple-hang tooling |
|  | Remove First Contact from AR face of TM. | assembly area | 5 min | 2 | razor blade |
|  | Pick up main chain with tips of Genie forks, using thin spacers (≈1/4”) on tip side to counteract sag. Bring over to reaction chain on pallet #1. | assembly area | 10 min | 4 | Genie, ¼” spacers |
|  | Match height first then bring in horizontally to mate locating pins. Insert ¼” spacers under main chain and then set down. | assembly area | 10 min | 4 | Genie, (more) ¼” spacers |
|  | Attach plates holding LSAT together. | assembly area | 5 min | 2 | hex-keys |
|  | Pick up whole LSAT with Genie, using thick spacers (≈1/2”) to counteract sag. | assembly area | 5 min | 4 | Genie, ½” spacers |
|  | Transfer LSAT to cart, clamp down, cover, and wheel to tank. | in transit | 15 min | 4 | cart, hex keys, dogs |
|  | Manhandle Genie past solid stack to tank area. | in transit | 5 min | 2 | Genie |
|  | Pick up LSAT with tips of forks using ½” spacers | near BSC | 5 min | 4 | Genie, ½” spacers |
|  | Take off lower plates connecting halves of LSAT. (These will foul on the elevator of the five-axis table.) | near BSC | 5 min | 2 | hex keys |
|  | Remove one side-plate (“2”) from elevator of five-axis table. | near BSC | 5 min | 2 | wrench |
|  | Place LSAT into elevator of five-axis table (avoiding bumping the HEPI with the Genie). The LSAT should be positioned vertically a few mm above the floor of the elevator and horizontally with locating holes directly above locating pins on pushers. | near BSC | 10 min | 4 | 5AT, elevator |
|  | Raise pushers to engage pins and support LSAT. Remove Genie. | near BSC | 5 min | 2 |  |
|  | Install side plate on elevator on five-axis table (matching corner “1” or “2” and inserting bolts from centre as before). | near BSC | 5 min | 2 | wrench |
|  | Lower LSAT until it is almost sitting on floor of elevator (so that it will go through door). | near BSC | 5 min | 2 |  |
|  | Roll table into chamber on conveyor. | near BSC | 5 min | 2+1 |  |
|  | Attach implementation ring to lower structure. | in BSC | 10 min | 2+1 | hex keys |
|  | Raise pushers until lower structure is about to contact, adjusting five-axis table DOFs as necessary to achieve a good mate. | near BSC | 5 min | 2+1 |  |
|  | Connect implementation ring to upper structure and install through bolts, using washers under the through bolts. | in BSC | 10 min | 2+1 | implementation ring, hex keys |
|  | Remove remaining plates holding halves of LSAT together. | in BSC | 5 min | 2+1 | hex keys |
|  | Remove upper two face stops on PM on both main and reaction chains. | in BSC | 5 min | 2+1 | hex keys |
|  | Put two long spacers (1” thick) on floor of elevator. | in BSC | 1 min | 2+1 | 1” spacers |
|  | Lower pushers till LSAT comes to rest on spacers. | in BSC | 2 min | 2+1 |  |
|  | Using two people per side simultaneously, pull the halves of the LSAT out and remove. | in BSC | 2 min | 4+1 |  |
|  | Remove UIM upper pitch mass (both chains). | in BSC | 10 min | 2+1 | hex keys |
|  | Connect wires from top mass to UIM (both chains). | in BSC | 5 min | 2+1 | hex keys |
|  | Install UIM upper pitch mass (both chains). | in BSC | 10 min | 2+1 | hex keys |
|  | Remove five-axis table and conveyor. | in BSC | 1 hour | 4 | ?? |
|  | Install two extra planks of flooring in the tank so there are two on each side and bring in the table to give easy access to top mass. | in BSC | 1 hour | 4 | ?? |
|  | [This shouldn’t be necessary! The ends of the glass rod in the upper half of the ring heater stuck out too far on each side and fouled on the sleeve. Assuming the radius of the curved section in the prototype was right, the welded on sections should be 20-25 mm long to serve their purpose but not foul.] Loosen the set screws holding the glass rod in the top half of the ring heater. Push one end of the rod in about 3 mm and tighten the set screw. Carefully push the other end of the rod in twice as far (bending the glass rod to a tighter radius!) and tighten that set screw. | in BSC | 10 min | 2+1 | hex keys |
|  | Remove all face earthquake stops on the bottom two levels of both chains. | in BSC | 10 min | 2+1 | hex keys |
|  | Bring in the sleeve, lift it up around the structure and bolt it on. [Need to rethink washers.] | in BSC | 10 min | 2+1 | hex keys |
|  | Install X braces. | in BSC | 10 min | 2+1 | hex keys |
|  | [This shouldn’t be necessary!] Restore the upper half of the ring heater to its original condition. | in BSC | 5 min | 2+1 |  |
|  | Reinstall face stops and set at 0.5 mm. | in BSC | 10 min | 2+1 | hex keys |
|  | Finish routing ESD and OSEM wiring. | in BSC | 20 min | 2+1 |  |
|  | Suspend per Section TBD. | in BSC | ?? | 2+1 |  |
|  | Drag wipe per Section 3.6. | in BSC | 20 min | 2+1 | foil, lens tissue, methanol |

# Sub-procedures

## Aligning the Brunson transit

| Step | What | Where | Time | People | Tools |
| --- | --- | --- | --- | --- | --- |
|  | Set up the Brunson about 10’-15’ from the structure, with the telescope at very roughly the height of the mass to be clocked. If you get too close you won’t be able to see both ears/prisms/clamps and if you get too far away, the ears will be too small in the viewfinder to have their height read accurately. If there is a very large difference in height then you need to be careful that the structure is facing the telescope accurately (so that the ears/prisms/clamps are the same distance away), but this is not at all critical. Midway in height between the bottom mass and the penultimate mass is probably good enough, and gets you two clockings for the one setup. |  | 30 min | 2 |  |
|  | Make sure the lock on the vertical height adjustment is tight and that upper mechanism is firm against moderate horizontal pressure. |  | 5 min | 1 | wrench: ¾” open-ended |
|  | Level the upper section as accurately as possible using the circular bubble level in the base of the rotating section. |  | 5 min | 1 |  |
|  | Turn the telescope pitch adjustment screw until it is roughly in the middle of its range. |  | 1 min | 1 |  |
|  | Unlock the telescope pitch clamp screw, roughly level the barrel of the telescope, and relock the clamp screw. |  | 1 min | 1 |  |
|  | Using the pitch adjustment screw, level the telescope as accurately as possible looking by eye at the barrel. |  | 1 min | 1 |  |
|  | Pick an opposing pair of the brass leveling discs in the leveling section and rotate the upper section until telescope is parallel with the line between the discs. |  | 1 min | 1 |  |
|  | Rotate the prism in the knurled housing near the top bubble level so that the aperture is at right angles to the telescope. |  | 1 min | 1 |  |
|  | Look into the prism aperture and adjust the long mirror to reflect the most ambient light into the side of the bubble level as indicated by the brightest view in the prism. |  | 1 min | 1 |  |
|  | Adjust the pitch adjustment screw until both ends of the bubble can be seen in the prism and are aligned with each other. |  | 1 min | 1 |  |
|  | Rotate the telescope by 180°, and then rotate the prism by a further 180° to bring the aperture back to the original direction. Readjust the long mirror if necessary. |  | 1 min | 1 |  |
|  | Grip the telescope pitch adjustment screw knob and note its position. Keep careful track of the amount of adjustment required in the next step, either by keeping a grip on the knob (if the amount is not too great), or counting the number of quarter turns of adjustment. |  | 1 min | 1 |  |
|  | Adjust the pitch adjustment screw until both ends of the bubble are aligned in the prism. |  | 1 min | 1 |  |
|  | Back the pitch adjustment screw off to a point as near as possible to halfway between the initial and final positions. |  | 1 min | 1 |  |
|  | Redo the second half of the levelling using the the two brass discs identified earlier, rotating them in opposite directions, so as to tighten one as the other is loosened. |  | 1 min | 1 |  |
|  | Rotate the telescope another 180° and readjust the prism and long mirror. Hopefully the ends of the bubble will be very nearly aligned. Repeat the previous six steps until convergence is achieved. |  | 5 min | 1 |  |
|  | Rotate the telescope by 90° to align with the other pair of brass discs and repeat the previous seven steps. |  | 15 min | 1 |  |
|  | Rotate the telescope back to the line of the first pair of brass disks and check that the alignment in that direction has not been disturbed. |  | 10 min | 1 |  |

## Using the ergo-arm

| Step | What | Where | Time | People | Tools |
| --- | --- | --- | --- | --- | --- |
|  | Connect ergo-arm reservoir to vacuum pump with hose. [According to Mike Gerfen, the hose should be permanently band-clamped to the reservoir, with the quick release fitting at other end connecting alternately to pump and suction plate. We were doing this backwards, and the following procedure has been revised to reflect what we should have done.] |  | 5 min | 1 | reservoir with hose, pump |
|  | Start pump, open valve at reservoir, evacuate reservoir to 30 psi, close valve, stop pump, disconnect hose. |  | 1 min | 1 | reservoir with hose pump |
|  | Connect hose to ergo-arm suction plate. |  | 1 min | 1 |  |
|  | Close valve at suction plate, open valve at reservoir, monitor reservoir gauge for short time (e.g., 1 min) to check for stable pressure (i.e., no leaks in hose or connections). |  | 2 min | 1 |  |
|  | Bring suction plate near to mass and use horizontal, vertical, pitch and yaw DOFs to match position and angle. |  | 5 min | 4 |  |
|  | Hold suction plate firmly against mass and open valve at plate. |  | 1 min | 4 |  |
|  | Check that good suction has been achieved (reservoir pressure should still be around 23 psi). If the alignment was poor there will likely be no vacuum at all, in which case, repeat from the beginning, being more careful in Step 5. |  | 1 min | 1 |  |
|  | Close the valve at the suction plate, and then the valve at the reservoir. (The suction plate has a very slight leak and a small volume, so closing it requires constant attention to the pressure at the suction plate. If it drops it can be topped up by opening both valves momentarily. But if both valves are open and someone trips over the reservoir and pulls the hose off one of the connectors it’s an instant catastrophe.) |  | 1 min | 1 |  |
|  | Raise mass, checking pressures at suction plate and reservoir regularly, and keeping a hand on the crankhandle at all times. |  | 1 min | 4 |  |

## Preparing the CP

| Step | What | Where | Time | People | Tools |
| --- | --- | --- | --- | --- | --- |
|  | Check the electrical continuity of the ESD cables. (The cables are extremely prone to failure at the end where the gold connectors have been crimped on.) |  | 1 day for all steps |  | ohmmeter |
|  | Take the CP out of its case, remove the face-plate from the ESD side, and lay it with the ESD side up in a clean room. [It was very difficult to remove the face plates because they were quite tight and there were no vent grooves in them.] |  |  |  |  |
|  | Carefully wipe the face and sides of the optic with lint-free wipes moistened with methanol [acetone?] to remove dust and dirt. |  |  |  | methanol [??acetone], wipes |
|  | Cut gold tabs to appropriate size: width about the same as the traces in the ESD mask, length sufficient to protrude about 5 mm off the edge of the optic. (This will be different for different traces.) [Brett: 5 mm turned out to be too much given the narrow clearance between the CP and the structure – it should be more like 2-3 mm.) |  |  |  | tabs |
|  | Crimp a furrow across the end of the gold tab which will be used to support the coax cable at a later step. |  |  |  | pliers |
|  | Set up a bottle of clean, dry nitrogen with a regulator and nozzles to direct a flow of nitrogen across the work area. |  |  |  | N2 bottle, regulator, nozzles |
|  | Repeat the next few steps for each tab to be soldered: |  |  |  |  |
|  | Point the nozzles at the end of the ESD trace that the tab is to be attached to. |  |  |  |  |
|  | Place a small bead of indium on the end of the trace, lay the flat end of the tab on top and cover with an aluminium button. |  |  |  | indium solder, Al button |
|  | Press a soldering iron heated to 600 degrees F onto the button and keep it there until 10 seconds after the indium melts. The button will visibly sag when the solder melts. Remove the button and inspect the joint. Too much heat can damage the pattern, so do not keep the iron there longer than necessary. |  |  |  | soldering iron, button |
|  | Remove a length of shield approximately 1.5” long from the end of the coax, exposing the (very delicate) central conductor and inner insulation. |  |  |  | wire stripper |
|  | Carefully strip the inner insulation exposing 2 to 3 mm of the central conductor. |  |  |  | wire stripper |
|  | Lay the end of the intact section of shield into the groove in the tab and roll the end of the tab over so that the shield is gripped. |  |  |  | pliers |
|  | Carefully bend the inner conductor around towards the body of the tab and solder it there, using the same procedure as for the tab. Maintain slight pressure on the tab at all times so that it does not move if it the solder behind it should melt. |  |  |  | soldering iron, solder, button |
|  | Test the electrical continuity from the pattern to the end of the cable. |  |  |  | ohmmeter |

## Using the triple-hang tooling

| Step | What | Where | Time | People | Tools |
| --- | --- | --- | --- | --- | --- |
|  | Start with main or reaction chain lower structure with all masses and wires in place, with the UIM approximately 4 mm high of nominal on its stops, and with UIM blades overloaded by 5 mm. |  |  |  |  |
|  | Check UIM is level and if not, adjust earthquake stops till it is. |  | 5 min | 2 | bubble level |
|  | Retract upper earthquake stops on bottom mass. |  | 5 min | 2 | hex keys |
|  | Screw in lifting screws on lower earthquake stops a tiny amount to ease weight on pad spacers. |  | 2 min | 2 | hex keys |
|  | Remove pad spacers. |  | 1 min | 2 |  |
|  | Retract lifting screws on lower earthquake stops until optic is suspended. |  | 5 min | 2 | hex keys |
|  | Check that optic is level relative to structure by eye – debug if not. |  | 1 min++ | 2 |  |
|  | Retract upper earthquake stops on PM. |  | 5 min | 2 | hex keys |
|  | Retract overload screws on UIM blades, monitoring lower masses. If blade strength is matched to payload, PM should be about 4 mm off lower stops (same as UIM was high to begin with). |  | 5 min | 2 | hex keys |
|  | Place 12 mm slip gauge on top of each UIM blade in turn and adjust blade height until top of slip gauge is level with reference notch in upright of UIM blade stop bridge (D060399). |  | 5 min | 2 | slip gauge |
|  | Check that PM is level relative to structure by eye – debug if not. |  | 1 min++ | 2 |  |
|  | On reaction chain, remove pitch adjuster, remove cable clamp, refit pitch adjuster. |  | 10 min | 2 | hex keys |
|  | Fit wire assemblies from triple-hang tooling to UIM. |  | 5 min | 2 | hex keys |
|  | Fit triple-hang tooling spacer blocks to top of lower structure. |  | 5 min | 2 | hex keys |
|  | Fit triple-hang tooling top plate to spacer blocks. |  | 5 min | 2 | hex keys |
|  | Connect wire assemblies to blades on triple hang tooling. |  | 5 min | 2 | hex keys |
|  | Release overload screws on triple-hang tooling. |  | 2 min | 2 | hex keys |
|  | Check that all three masses are level relative to structure by eye – debug if not. |  | 1 min++ | 2 |  |
|  | Reapply overload screws on triple-hang tooling until tension is off wire assemblies. |  | 5 min | 2 | hex keys |
|  | Disconnect wire assemblies at blades triple-hang tooling. |  | 5 min | 2 | hex keys |
|  | Remove triple-hang tooling top plate and spacer blocks. |  | 5 min | 2 | hex keys |
|  | Disconnect wire assemblies at UIM. |  | 5 min | 2 | hex keys |
|  | On reaction chain, remove pitch adjuster, fit cable clamp, replace pitch adjuster. |  | 10 min | 2 | hex keys |

## Applying/removing First Contact

### Applying

| Step | What | Where | Time | People | Tools |
| --- | --- | --- | --- | --- | --- |
|  | See E070292-00. |  | 3 hours | 1 | See E070292-00. |

### Removing

| Step | What | Where | Time | People | Tools |
| --- | --- | --- | --- | --- | --- |
|  | Carefully shave the entire bevel with a sharp single-sided razor blade to remove traces of First Contact that may have spilled there off the face. |  | 5 min | 1 | razor blade |
|  | With the edge of the razor blade leading, scrape from the bevel toward the face to prise up a corner of the First Contact on the face. The corner between the straight and curved sections is a particularly good place to start. |  | 1 min | 1 | razor blade |
|  | Grab the prised-up corner with gloved fingers and carefully pull the whole sheet off the face, avoiding tears as much as possible. |  | 1 min | 1 |  |
|  | If any small patches of First Contact remain, very carefully scrape them off with a razor blade and clean up the area with spectroscopic grade methanol and a lens tissue. (This should not happen if the First Contact was applied thickly enough originally.) |  | 5 min | 1 | razor blade, methanol, lens tussue |

## Drag-wiping

| Step | What | Where | Time | People | Tools |
| --- | --- | --- | --- | --- | --- |
|  | Pour a little spectroscopic grade methanol into a small foil boat or dish. |  | 5 min |  | foil, methanol |
|  | Repeatedly, bend a sheet of lens tissue (3”x5” is good) in half without creasing it, dip the bend in the methanol and drag slowly across the optic. (Doug: This bend technique is particularly good for vertical surfaces.) |  | 1 min |  | lens tissue |
|  | Work by strips, using a fresh sheet each time. If the lens tissue does not stick to the optic with surface tension, it is too dry. If it leaves streaks of liquid methanol behind (especially from the corners), it is too wet. |  | 10 min |  | lens tissue |