| California Institute of Technology MC 18-34, 1200 E. California Blvd. Pasadena CA 91125 USA TEL: 617.395.2129 FAX: 617.304.9834 www.ligo.caltech.edu | LIGO Livingston Observatory P.O. Box 940 Livingston LA 70754 USA TEL: 225.686.3100 FAX: 225.686.7189 www.ligo-la.caltech.edu | LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 USA TEL: 509.372.8106 FAX: 509.372.8137 www.ligo-wa.caltech.edu | Massachusetts Institute of Technology MIT NW22 – 295, 185 Albany St. Cambridge MA 02139 USA TEL: 617.235.4824 FAX: 617.253.7014 www.ligo.mit.edu |
| --- | --- | --- | --- |

| Date: | 24 March 2009 | Refer to: | T0900120-v2 |
| --- | --- | --- | --- |
| Subject: | Dynamics of the Damper Suspension for the ETM Transmission Telescope/Monitor |
| To: | Mike Smith |
| From: | Dennis Coyne |

Revision History:

v1	Initial release which simply demonstrates that the  ring (upper) stage yaw (torsional) mode has a high frequency. There were 3 errors in v1: (1) the magnet plate mass was only 3 kg, (2) the ring was comprised of a stainless steel tubes, not aluminum as intended and (3) the pendulum length was set to the full TransMon length (optics table to optical axis length of 1742 mm), but should be shorter in order to interface to the top of the main TransMon assembly.

v2	Some preliminary sizing to achieve high frequency modes.

## Introduction

A suspended magnet plate serves to eddy current damp the adjacent suspended ETM Telescope and transmitted beam monitor assembly (aka TransMon). Eddy current damping requires a different frequency for the reaction/damping suspension than the main suspension, in order to develop relative velocities between the suspension chains. A recent concept is to effectively half the length of the reaction/damping suspension by using a stiff “wire structure” which is about half the full pendulum length and then a compliant suspension for the magnet plate hanging from this stiff “wire structure”. The wire structure was proposed to support a stiff triangular ring with two wires splayed out from each corner of the ring to the optics table.

It is likely that all modes of the magnet plate suspension must either be < ~10 Hz or > 150 Hz, in order to stay sufficiently below or above the upper unity gain frequency of the BSC ISI control. It is possible that modal frequencies in the range of 10 to 150 Hz could be acceptable if the associated modal mass is low and/or the interaction of the mode with the compliance of the BSC optics table is slight; This aspect is not discussed in this memo.

The intent of this memo is to roughly determine the parameters of a magnet plate suspension which meets (or comes close to) the desired dynamics.

## Suspension Geometry and Properties

The payload for the ETM Transmission Telescope and monitor (aka “TransMon”) is depicted in Figure 1. The optics bench is approximately 1.05 m by 0.6 m (from the Zemax file, D0900446-v2). The depth of the off-axis, parabolic reflective telescope is 0.3 m. (also from D0900446-v2). The distance from the optics table to the center of the ETM is 1742 mm and to the center of the aperture of the large telescope is ~1742 mm.

The TransMon is suspended from the BSC optics table with a single pendulum suspension. The intent is to passively damp this suspension with another suspension which supports a plate with a set of magnets (no concept sketches were found in the DCC).

Figure 1: Transmon Table, Zemax Layout (2008-12-04 email from M. Smith, “ETM Transmon”)
[approximately consistent with zemax file D0900446-v2]

I choose to represent the reaction/damper suspension in ANSYS with:

- an aluminum ring in the form of an equilateral triangle with a circumscribed radius of ~0.25 m. Each bar in the ring has an outer radius of ~15 mm and an inner radius of ~13 mm (hollow circular bar with ~2 mm wall thickness). The resulting mass is ~1.2 kg for the 3 bars total. (These are nominal ring dimensions – the dimensions were varied.)

- 0.1 to 1 kg mass was added at each of the 3 vertices of the ring to represent wire clamps & tube conections

- Two steel wires with ~1.5 mm radius support each of the 3 ring vertices. The wires connect to points on the optics table that are also the vertices of an equilateral triangle circumscribed by a radius of 0.5 m, but rotated 30 degrees from the ring.

- The ring was set to a distance of half (1742 -300) mm, or 721 mm, below the optics table

- At each of the three vertices of the ring, a vertical wire extends down to 1442 mm (the full magnet plate pendulum length). These vertical wires also have ~1.5 mm radius.

- The payload (the magnet plate) was represented as a point mass of 10 kg connected by rigid links (constraint equations) to the ends of the vertical wires. In order to effectively damp the main suspension, the magnet plate mass must be ~1/3 of the main suspension mass, which is guessed to be ~30 kg.

- The wires and ring were meshed as beam elements.

A depiction of the geometry is shown in Figure 2.

| Isometric view | Top view |
| --- | --- |

Figure 2: Simple ANSYS beam model of the reaction/damper suspension for the TransMon

The ANSYS script that generates the model and solves for the mode shapes and frequencies is listed in the appendix.

## Frequencies

The modal frequencies were determined as some of the key parameters of the suspension were varied. The “nominal” case was for the ring circumscribed radius to be the same as the radius that circumscribes the points of attachment at the optics table, 0.25 m (i.e. vertical upper stage wires). The modal frequencies are listed in Table 1. The ring radius was then varied to determine an approximate optimum. At 0.17 m, the pendulum modes of the ring are greater than 150 Hz, but the upper wire violin modes decreased to 130 Hz. Note that due to highly coupled nature of the modes (it seems), the modal frequency variation is not very uniform (or perhaps the FEA is incorrect?).

Using a ring circumscribed radius of 0.17 m (with the optics table attachment point circumscribed radius still 0.25 m), the ring added vertex mass and wire radii were varied. For a ring added vertex mass of only 0.1 kg (unreasonable?), upper wire radii of 1.7 mm and lower radii of 1.3 mm, the frequencies (just barely) meet the stated frequency exclusion zone.

Note that the magnet plate roll & pitch frequencies can be lowered by making the circumscribed radius of the attachment points at the magnet plate smaller (while keeping the lower stage wires vertical). Note also that lower stage pitch mode is not modeled (no moments of inertia set for the magnet plate mass, or pitch offset distance for the wire attachment). Note also that there are a number of frequencies below 150 Hz in this particular design.

Table 1: Modal Frequency variation with suspension parameters

## Selected mode shapes

## Appendix: ANSYS Macro Listing

!***********************************************************************

! pend4.mac

!***********************************************************************

! macro to analyse natural frequencies of a double pendulum suspension

! proposed for the reaction/damper suspension used on the TransMon assy

!

!Dennis Coyne 2009-03-23

!

finish

/CLEAR,START

/COM,ANSYS MODEL OF A SIMPLE PENDULUM

/PREP7

/TITLE,PENDULUM

PSTRES,ON

!***********************************************************************

!*     GEOMETRIC PARAMETERS

!***********************************************************************

! values of parameters

! SI units (m,N)

! r1 = radius of circle circumscribing equilateral triangle which forms

!       wire termination points at the optics table interface

! r2 = radius of circle circumscribing equilateral triangle ring of 1st stage

! r3 = radius of circle circumscribing equilateral triangle which forms

!       wire termination points at the magnet plate

! l1 = distance from optics table to ring

! l2 = distance from ring to magnet plate

! rw1 = radius of upper wires

! rw2 = radius of lower wires

! rro = outer radius of rods forming 1st stage triagular ring

! rri= inner radius of rods

! mr = added mass of the ring (added to the vertices)

! mmp = mass of the magnet plate

! Inx, Iny, Inz = mass moments of inertia for the magnet plate

pi=3.1416

r1=0.5

r2=0.17

r3=r2

l1=(1.742-0.300)/2

l2=l1

rw1=0.0017

rw2=0.0013

rro=0.015

rri=0.013

!***********************************************************************

!*     MASS PARAMETERS

!***********************************************************************

mmp=10

mr=0.1

Inx=0

Iny=0

Inz=0

!***********************************************************************

!*     MATERIAL PROPERTIES

!***********************************************************************

ringyoung=68.9e9

ringpoiss=0.33

ringdens=2700

wireyoung=2.0e11

wirepoiss=0.3

wiredens=7800

MP,EX,1,ringyoung

MP,EY,1,ringyoung

MP,EZ,1,ringyoung

MP,PRXY,1,ringpoiss

MP,DENS,1,ringdens

MP,EX,2,wireyoung

MP,EY,2,wireyoung

MP,EZ,2,wireyoung

MP,PRXY,2,wirepoiss

MP,DENS,2,wiredens

!***********************************************************************

!*     GENERATE GEOMETRY

!***********************************************************************

K,1,0,r1,0,

K,2,-r1*cos(pi/6),-r1*sin(pi/6),0

K,3, r1*cos(pi/6),-r1*sin(pi/6),0

K,4,-r2*cos(pi/6), r2*sin(pi/6),-l1

K,5,0,-r2,-l1

K,6, r2*cos(pi/6), r2*sin(pi/6),-l1

K,7,-r3*cos(pi/6), r3*sin(pi/6),-l1-l2

K,8,0,-r3,-l1-l2

K,9, r3*cos(pi/6), r3*sin(pi/6),-l1-l2

K,10,0,0,-l1-l2

!***********************************

!*          Wires

!***********************************

LSTR,1,4     ! line 1

LSTR,1,6     ! line 2

LSTR,2,4     ! line 3

LSTR,2,5     ! line 4

LSTR,3,5     ! line 5

LSTR,3,6     ! line 6

LSTR,4,7     ! line 7

LSTR,5,8     ! line 8

LSTR,6,9     ! line 9

!***********************************

!*          Ring

!***********************************

LSTR,4,5     ! line 10

LSTR,5,6     ! line 11

LSTR,6,4     ! line 12

!***********************************

!*          Rigid links

!***********************************

LSTR,10,7    ! line 13

LSTR,10,8    ! line 14

LSTR,10,9    ! line 15

!***********************************************************************

!*     MESH GEOMETRY

!***********************************************************************

!***********************************

!           Suspension Masses

!***********************************

ET,2,MASS21,0,0,0

TYPE,2     ! discrete mass elements

!     Ring Mass

R,4,mr/3,mr/3,mr/3,0,0,0

REAL,4

KMESH,4,6,1

!     Magnet Plate Mass

R,5,mmp,mmp,mmp,Inx,Iny,Inz

REAL,5

KMESH,10

!***********************************

!           Wires

!***********************************

ET,3,BEAM4,,0,,,,0

TYPE,3     ! beam elements

MAT,2

wAn=3.14159*rw1**2

wIn=(3.14159*rw1**4)/4

wJn=(3.14159*rw1**4)/2

R,9,wAn,wIn,wIn,2*rw1,2*rw1,0

RMORE,0,wJn

REAL,9

ESIZE,0.005

LMESH,1,6

wAn=3.14159*rw2**2

wIn=(3.14159*rw2**4)/4

wJn=(3.14159*rw2**4)/2

R,9,wAn,wIn,wIn,2*rw2,2*rw2,0

RMORE,0,wJn

REAL,9

ESIZE,0.005

LMESH,7,9

!***********************************

!           Ring

!***********************************

ET,3,BEAM4,,0,,,,0

TYPE,3     ! beam elements

MAT,1

wAn=3.14159*(rro**2-rri**2)

wIn=(3.14159*(rro**4-rri**4))/4

wJn=(3.14159*(rro**4-rri**4))/2

R,10,wAn,wIn,wIn,2*rro,2*rro,0

RMORE,0,wJn

REAL,10

ESIZE,0.01

LMESH,10,12

!***********************************

!           Rigid Links

!***********************************

ET,4,MPC184,1,0

TYPE,4      ! rigid body constraint elements

ESIZE,,1

LMESH,13,15

!***********************************************************************

!*     BOUNDARY CONDITIONS

!***********************************************************************

DK, 1,ALL,0

DK, 2,ALL,0

DK, 3,ALL,0

EPLOT

FINISH

!***********************************************************************

!   Static Preload Analysis

!***********************************************************************

/SOL

ANTYPE,STATIC

PSTRES,ON

G=9.84

ACEL,0,0,G

SOLVE

FINISH

!***********************************************************************

!   Eigenvalue Analysis

!***********************************************************************

/SOL

ANTYPE,MODAL

nmodes=50

initialshiftfreq=0

MODOPT,LANB,nmodes

LUMPM,OFF

PSTRES,ON

SOLVE

FINISH

# Footnotes

Note that T060360-01, AOS: PO Mirror Assembly & Telescope, and OMMT Conceptual Design Requirements, is obsolete. The telescope is no longer refractive, the aperture is larger, the mass is likely much greater than 9.5 kg, and the ETM wedge angle is smaller (0.8 deg), so the distance is no longer 1744 mm.