Geotech GS-13

Alternative Flexure Installation Instructions

LIGO-T0900583-V1

Geotech GS-13 seismometer beryllium-copper flexure installation instructions as used in aLIGO ISI systems.

Graham Allan

Daniel Clark

Brian Lantz

2009.11.18

# Table of Contents

Table of Contents	2

Table of Figures	2

1. Introduction	4

2. Pre-Assembly	4

3. Assembly Procedure	5

4. Installing the preamp board	12

5. Installing the anti-resonant offload spring absorbers	17

6. BeCu flexure installation	19

7. Testing	30

8. Reassembly	34

# Table of Figures

Figure 1: Before initial disassembly showing the recommended tools and equipment	5

Figure 2: Close shot of recommended tools	5

Figure 3: Ensure that the GS-13 is locked	6

Figure 4: Turn the GS-13 upside down in preparation for can removal	6

Figure 5: Remove the factory feet and lock ring from the three studs	7

Figure 6: Factory feet removed - these will not be reassembled on the instrument	7

Figure 7: Using the 9/16 combination wrench, loosen the three nuts to about halfway up the stud	8

Figure 8: Slide the aluminum washer cap and rubber grommet up slightly on each stud	8

Figure 9: Push down on the studs while gripping and pulling up on the can until…	9

Figure 10:  It slides off the sealing o-ring at the base of the instrument	9

Figure 11: Next remove the nuts, cap washers, and grommets and set them aside for reassembly later	10

Figure 12: Removed hardware for can removal	10

Figure 13: Carefully lift the can off of the instrument	11

Figure 14: GS-13 with can removed	11

Figure 15: Using the P1 screwdriver remove the existing calibration board	12

Figure 16: Pull the board out and make sure that pin 1 is marked on the connector - if not mark it	12

Figure 17: Unplug the connector from the circuit board	13

Figure 18: The LIGO preamp ready to install	13

Figure 19: Plug the connector into the new board observing correct orientation	14

Figure 20: Tuck the wires back in making sure that the wires do not contact offload springs, cantilevers, or flexures	14

Figure 21: Visually ensure the wires are not interfering with the offload springs	15

Figure 22: Attach the circuit board with original screws	15

Figure 23: Turn the GS-13 over and set it on its studs	16

Figure 24:  Make sure that all of the jumpers on the board are positioned as shown - in the lower positions	16

Figure 25: Locate the small Viton o-ring on the top of each offload spring (vertical models only)	17

Figure 26: Slip the o-ring around the corner by about 1 cm.	17

Figure 27: Ensure that the spring exits the rod-end that supports it by right angles – if not – rotate it into alignment (this one is not aligned)	18

Figure 28: This one is aligned - repeat for all three offload springs	18

Figure 29: Using the 3/32 hex wrench, remove the screws and plates that hold the lower delta rods in place	19

Figure 30: While supporting the free end remove the other side	19

Figure 31: The removed delta rod, SHCS, and plates (top) will not be reused, the new flexure and 4-40 x 3/8 in. SHCS bottom will be installed	20

Figure 32: Slide the new BeCe flexure in place	20

Figure 33: While holding the end in the correct position tighten the screw	21

Figure 34: Start the screw by about 4 threads	21

Figure 35: Place one clamp on the arm holding the top plate in place	22

Figure 36: Place the second clamp holding the end up against the temporary plate	22

Figure 37: Showing the temporary clamps and plate	23

Figure 38: Attach by tightening the screw	23

Figure 39: Remove the top delta rods	24

Figure 40: Slip the new BeCu flexure in place	24

Figure 41: BeCu flexure partially inserted	25

Figure 42: Lift up the end and start sliding it in place from the back	25

Figure 43: Slightly push the cantilever to the right to help slide the end forward	26

Figure 44: Lift the other end into place	26

Figure 45: Attach both ends	27

Figure 46: This is the trickiest flexure installation	27

Figure 47: Flexure slipped in	28

Figure 48: The flexure will bind up at this point, push the BeCu flexure forward bending the flexure in order to enable it to slip in place	28

Figure 49:  Hold it in place and attach with hardware	29

Figure 50: Unlock the proof mass and look at the red indicator - it should fill the center ring completely	29

Figure 51: Remove the outer two jumpers from the pre-amp board	30

Figure 52: Attach the 10X probe and ground clip to the center conductor on each side of the pre-amp	30

Figure 53: Take a ring-down by holding the mass down against its stop and releasing it	31

Figure 54: Typical ring-down. In this case the natural frequency was low	31

Figure 55: To raise the natural frequency loosen the lock collar with the 7/64 hex wrench and turn the knob counterclockwise to increase the frequency	32

Figure 56: Corrected ring-down showing good response	32

Figure 57: Replace the jumpers on the pre-amp board	33

Figure 58: Relock the proof mass	33

Figure 59: Turn the seismometer over	34

Figure 60: Carefully lower the lid over the instrument	34

Figure 61: A fairly large downward force is required to slip the cover over the o-ring	35

Figure 62: Install the grommet,  cap, and nut on the studs. Tighten with the 9/16 wrench	35

Figure 63: The correct sequence of installation	36

Figure 64: Installed can	36

Figure 65: Completed GS-13 flexure retrofit ready for podding in aLIGO	37

# 1. Introduction

The Geotech GS-13 seismometer is supplied to aLIGO with the standard delta-rods. These need to be replaced with the alternative beryllium-copper flexures before being placed inside the GS-13 vacuum pods and placed in service in the aLIGO system. Along with the flexure installation, a new aLIGO pre-amp circuit board must be installed, anti-resonance o-rings slipped in place on the offload springs on the vertical instruments, and the instrument tested.

There are some differences in the horizontal and vertical GS-13 configurations. Namely, the horizontal GS-13 seismometers as used in aLIGO do not have the three offload springs, cantilevers, and supporting flexures. In the case that the instrument being fitted is a horizontal unit the installer may disregard the steps that only pertain to the vertical GS-13.

The following instructions assume that the required flexures and GS-13s have already passed their acceptance tests.

# 2. Pre-Assembly

The following tools are required for the assembly and testing as described in this document:

- 3 approved top flexures

- 3 approved bottom flexures

- aLIGO pre-amp board

- 12 4-40 x 3/8 in. stainless steel Socket Head Cap Screws (SHCS)

- 9/16 in. combination wrench

- 3/32 hex wrench

- 7/64 hex wrench

- Bottle Loctite 290 green wicking grade threadlocker

- Paint pen or marker

- Screwdriver – P1

- Miniature clamps (we used two plastic tweezers with o-rings and a stainless steel strap)

- Oscilloscope with 10X probe

# 3. Assembly Procedure

The entire assembly procedure should be done in a relatively clean environment. A cleanroom is not necessary but care should be taken to avoid contamination.

Figure 1: Before initial disassembly showing the recommended tools and equipment

Figure 2: Close shot of recommended tools

Figure 3: Ensure that the GS-13 is locked

Verify that the proof-mass internal to the GS-13 is locked by both cranking the “LOCK” knob clockwise and the red flower in the center of the view glass fills the view glass to the outer ring.

Figure 4: Turn the GS-13 upside down in preparation for can removal

Figure 5: Remove the factory feet and lock ring from the three studs

Figure 6: Factory feet removed - these will not be reassembled on the instrument

Figure 7: Using the 9/16 combination wrench, loosen the three nuts to about halfway up the stud

Figure 8: Slide the aluminum washer cap and rubber grommet up slightly on each stud

Figure 9: Push down on the studs while gripping and pulling up on the can until…

Figure 10:  It slides off the sealing o-ring at the base of the instrument

Figure 11: Next remove the nuts, cap washers, and grommets and set them aside for reassembly later

Figure 12: Removed hardware for can removal

Figure 13: Carefully lift the can off of the instrument

Figure 14: GS-13 with can removed

# 4. Installing the preamp board

Figure 15: Using the P1 screwdriver remove the existing calibration board

Figure 16: Pull the board out and make sure that pin 1 is marked on the connector - if not mark it

Figure 17: Unplug the connector from the circuit board

Figure 18: The LIGO preamp ready to install

Figure 19: Plug the connector into the new board observing correct orientation

Figure 20: Tuck the wires back in making sure that the wires do not contact offload springs, cantilevers, or flexures

Figure 21: Visually ensure the wires are not interfering with the offload springs

Figure 22: Attach the circuit board with original screws

Figure 23: Turn the GS-13 over and set it on its studs

Figure 24:  Make sure that all of the jumpers on the board are positioned as shown - in the lower positions

# 5. Installing the anti-resonant offload spring absorbers

Figure 25: Locate the small Viton o-ring on the top of each offload spring (vertical models only)

Figure 26: Slip the o-ring around the corner by about 1 cm.

Figure 27: Ensure that the spring exits the rod-end that supports it by right angles – if not – rotate it into alignment (this one is not aligned)

Figure 28: This one is aligned - repeat for all three offload springs

# 6. BeCu flexure installation

Figure 29: Using the 3/32 hex wrench, remove the screws and plates that hold the lower delta rods in place

Figure 30: While supporting the free end remove the other side

Figure 31: The removed delta rod, SHCS, and plates (top) will not be reused, the new flexure and 4-40 x 3/8 in. SHCS bottom will be installed

Figure 32: Slide the new BeCe flexure in place

Figure 33: While holding the end in the correct position tighten the screw

Figure 34: Start the screw by about 4 threads

Figure 35: Place one clamp on the arm holding the top plate in place

Figure 36: Place the second clamp holding the end up against the temporary plate

Figure 37: Showing the temporary clamps and plate

Figure 38: Attach by tightening the screw

Figure 39: Remove the top delta rods

Figure 40: Slip the new BeCu flexure in place

Figure 41: BeCu flexure partially inserted

Figure 42: Lift up the end and start sliding it in place from the back

Figure 43: Slightly push the cantilever to the right to help slide the end forward

Figure 44: Lift the other end into place

Figure 45: Attach both ends

Figure 46: This is the trickiest flexure installation

Figure 47: Flexure slipped in

Figure 48: The flexure will bind up at this point, push the BeCu flexure forward bending the flexure in order to enable it to slip in place

Figure 49:  Hold it in place and attach with hardware

Figure 50: Unlock the proof mass and look at the red indicator - it should fill the center ring completely

If the red indicator is not centered, or not filling completely the center indicator ring, adjust the proof mass offload springs by the three adjustment knobs. Clockwise raises the mass, counterclockwise lowers the mass. Try to turn all three the same amount to fill the indicator evenly.

# 7. Testing

Figure 51: Remove the outer two jumpers from the pre-amp board

Figure 52: Attach the 10X probe and ground clip to the center conductor on each side of the pre-amp

Figure 53: Take a ring-down by holding the mass down against its stop and releasing it

Figure 54: Typical ring-down. In this case the natural frequency was low

Figure 55: To raise the natural frequency loosen the lock collar with the 7/64 hex wrench and turn the knob counterclockwise to increase the frequency

Figure 56: Corrected ring-down showing good response

Figure 57: Replace the jumpers on the pre-amp board

Figure 58: Relock the proof mass

# 8. Reassembly

Figure 59: Turn the seismometer over

Figure 60: Carefully lower the lid over the instrument

Figure 61: A fairly large downward force is required to slip the cover over the o-ring

Figure 62: Install the grommet,  cap, and nut on the studs. Tighten with the 9/16 wrench

Figure 63: The correct sequence of installation

Figure 64: Installed can

Figure 65: Completed GS-13 flexure retrofit ready for podding in aLIGO