## Info
A analog
E external (from PSL to …)
D digital
W water
L light
Version control:
2006-06-06 | Benno | first draft with basic structure, send to Maik and Peter for editing their sections
2008-02-06 | Christian/Maik | new stucture with A, D, X etc interfaces
2009-01-14 | Benno | new interface definition picture, major revision of document for PSL FDR
The purpose of this document is to define the interfaces between "subsystems".

## modules
| front | [7 empty] | back
PSL control chassis 2 | FB ILS | ILS |  | monitor FB |  | servo PMC | FB PMC |  | FB ILS | servo ILS |  | monitor FB |  | servo PMC | FB PMC
|  |  |  | AMP |  |  |  |  | power | power |  | power |  | power | power | power in
ISO
|  | from PD |  | BREW |  | from PD
|  | LO |  | INT |  | LO
| resonant | [5 empty] | resonant
| trigger |  |  | PMC refl |  |  | trigger
| locked |  |  | PMC trans |  |  | locked
| HV mon |  |  | HPL_DC |  |  | HV mon
HPL_AC
[9 empty] | ILS bus | [5 empty] | PMC bus
| to Beckhoff |  |  | MIS-Flow
| from Beckhoff |  |  | PWR NPRO
NPRO RRO
|  | HV (375V) |  |  |  | HV (375V)
|  | to PZT |  |  |  | to PZT
| to AI |  |  | to AI |  |  | to AI
| to AA |  |  | to AA |  |  | to AA
PSL control chassis 1 | servo ISS |  |  |  |  | VCO FB | TTFSS FB |  | servo ISS |  |  |  |  | VCO FB | TTFSS FB
[9 empty] | power | power | power |  | power | power | power | power in
| outer loop in |  |  |  |  | to VCO | to rf-PD
| PD A | [5 empty] | to trans-PD
PD B
| QPD | [5 empty] | to table top
TT FSS
AOM
HV 150V
| to AI |  |  |  |  | to AI | to AI
| to AA |  |  |  |  | to AA | to AA
PSL control chassis 3 | HV modul | DBB module | DBB module | DBB module | DBB module | DBB module | DBB FB | table top boxes | TT FSS |  | 80MHz LO |  | AOM Driver
[11 empty] | to FB FSS LO |  | 80 MHz in
[9 empty] | to front of |  |  |  | from ISS
[9 empty] | FB FSS |  |  |  | to AOM
to LO FSS
to AI
to AA
bus lines connected only between communicating modules (via ribbon cable)
backplane bus connects all modules BUT the HV modue
attention:
18V for ISS module and left part of chasis
24V for TTFSSFB and VCO FBB and right part of chassis
both voltages are connected to same bus line number (the one that was formerly 24V in iLIGO)

## Analog (A)
| No | sensing / control | bandwidth | bits | Chassis # | ADC/DAC # | Comments | Variable (Beckhoff) | Variable (EPICS)
A1 / PSL EPICS control
A1.1 pre modecleaner PMC
DB9 Connector 1 (In1)
| 1 |  |  | 1 | 2 | DAC_2_0
sw2 | 1 |  | NA | 1 | 2 | DAC_2_1 | enables summation of TEST2 input |  | {ifo}:PSL-PMC_TF_IN_ON
Enable locking (lock on) | 1 |  | NA | 1 | 2 | DAC_2_2 | turns lock automation on |  | {ifo}:PSL-PMC_LOCK_ON
Ramp on/off (ramp on) | 1 |  | NA | 1 | 2 | DAC_2_3 | turns ramp on |  | {ifo}:PSL-PMC_RAMP_ON
DB9 Connector 2 (In3)
Set gain | 1 |  | NA | 16 | 2 | DAC_2_8 | set gain of control loop |  | {ifo}:PSL-PMC_GAIN
Set power threshold (ref) | 1 |  | NA | 16 | 2 | DAC_2_9 | sets threshold which lock acquisition automation assumes PMC is locked |  | {ifo}:PSL-PMC_REF
Set error point offset (in offset) | 1 |  | NA | 16 | 2 | DAC_2_10 | correction of electronic offset in control loop |  | {ifo}:PSL-PMC_INOFFSET
heater |  |  | NA |  | 2 | DAC_2_11 | set offset of heater and hence operation point of PMC temp. loop |  | {ifo}:PSL-PMC_HEATER
DB9 Connector 3 (In2)
blancking | 1 |  | NA | 16 | 2 | DAC_2_4 | opens feedback loop |  | {ifo}:PSL-PMC_BLANKING
gain boost | 1 |  | NA | 16 | 2 | DAC_2_5 | activates integrator in control loop |  | {ifo}:PSL-PMC_BOOST
calibration line in (Tfin) | 1 |  | NA | 16 | 2 | DAC_2_6 | summation point for calibration lines, used for alignment ramp |  | {ifo}:PSL-PMC_TF_IN
ramp |  |  | NA |  | 2 | DAC_2_7 | not used (would be ramp input in case digital lock acquisition is used) |  | {ifo}:PSL-PMC_RAMP
Lemo 00
PTZ voltage monitor | 1 |  |  | 16 | 2
Lemo 00 (prelim)
Relock counter | 1 |  |  | 1 | 2
DB9 Connector 4 (Out 2)
HV monitor | 1 |  |  | 1 | 2 | ADC_2_4 | monitor of HV signal send to PZT actuator of PMC |  | {ifo}:PSL-PMC_HV_MON
Mixer out monitor | 1 |  |  | 1 | 2 | ADC_2_5 | error signal of control loop |  | {ifo}:PSL-PMC_MIXER
Local oscilllator power (LODET) | 1 |  |  | 1 | 2 | ADC_2_6 | power of local oscillator used in demodulation electronic |  | {ifo}:PSL-PMC_LO_POWER_MON
"resonant" trigger (TRIGGER) | 1 |  |  | 1 | 2 | ADC_2_7 | DC signal of locking diode, used by autolock to dertermine if PMC is resonant |  | {ifo}:PSL-PMC_TRIGGER
DB9 Connector 5 (Out1)
TF B | 1 |  |  | 16 | 2 | ADC_2_0 | test point for transfer function measurement |  | {ifo}:PSL-PMC_TF_B
Temp | 1 |  |  | 16 | 2 | ADC_2_1 | temperature sensor on PMC |  | {ifo}:PSL-PMC_TEMP
TF A | 1 |  |  | 16 | 2 | ADC_2_2 | test point for transferfunction measurement |  | {ifo}:PSL-PMC_TF_A
Resonant | 1 |  |  |  | 2 | ADC_2_3 | binary signals, 1 if auto lock thinks PMC is resonant |  | {ifo}:PSL-PMC_RESONANT_MON
A1.2 intensity stabilization ISS
DB9 Connector 1 (Out3)
AOM | 1 |  |  | 16 | 1 | ADC_0_8 | ISS actuator signal send to AOM |  | {ifo}:PSL-ISS_AOM
QPD dX | 1 |  |  | 16 | 1 | ADC_0_9 | quadrant photodiode dx signal |  | {ifo}:PSL-ISS_QPDDX
QPD dY | 1 |  |  | 16 | 1 | ADC_0_10 | quadrant photodiode dy signal |  | {ifo}:PSL-ISS_QPDDY
DB9 Connector 2 (In1)
reference signal | 1 |  |  | 16 | 1 | DAC_0_0 | reference to which the ISS stabilizes the power |  | {ifo}:PSL-ISS_REFSIGNAL
transfer 1 in | 1 |  |  | 16 | 1 | DAC_0_1 | summation into error point |  | {ifo}:PSL-ISS_TRANSFER1_INJ
transfer 2 in | 1 |  |  | 16 | 1 | DAC_0_2 | summation into outer-loop signal |  | {ifo}:PSL-ISS_TRANSFER2_INJ
set gain | 1 |  |  | 16 | 1 | DAC_0_3 | ISS control loop gain |  | {ifo}:PSL-ISS_GAIN
DB9 Connector 3 (In2)
AOM offset | 1 |  |  | 16 | 1 | DAC_0_4 | offest added to control signal to enable AOM to increase power |  | {ifo}:PSL-ISS_CTRL_OFFSET
interchange of photodiodes | 1 |  |  | 16 | 1 | DAC_0_5 | selects if photodiode A of B is used as in-loop sensor |  | {ifo}:PSL-ISS_INLOOP_PD_SELECT
close outer loop | 1 |  |  | 16 | 1 | DAC_0_6 | switch to open/close second ISS feedback loop |  | {ifo}:PSL-ISS_SECONDLOOP_CLOSED
close feedback control loop | 1 |  |  | 16 | 1 | DAC_0_7 | switch to open/close feedback loop |  | {ifo}:PSL-ISS_LOOP_STATE
DB9 Connector 4 (Out1)
PD A monitor | 1 |  |  | 16 | 1 | ADC_0_0 | monitor signal of photodiode A, (at output of signal conditioning filter) |  | {ifo}:PSL-ISS_PDAMON
PD B monitor | 1 |  |  | 16 | 1 | ADC_0_1 | monitor signal of photodiode B, (at output of signal conditioning filter) |  | {ifo}:PSL-ISS_PDBMON
Outer PD monitor | 1 |  |  | 16 | 1 | ADC_0_2 | monitor signal of second loop photodiode |  | {ifo}:PSL-ISS_OUTERPDMON
reference signal out | 1 |  |  | 16 | 1 | ADC_0_3 | low pass filtered readback of ref. signal |  | {ifo}:PSL-ISS_REFSIGNALOUT
DB9 Connector 5 (Out2)
transfer 1A | 1 |  |  | 16 | 1 | ADC_0_4 | signal infront of summation point in first loop |  | {ifo}:PSL-ISS_TRANSFER1A
transfer 1B | 1 |  |  | 16 | 1 | ADC_0_5 | signal behind summation point in first loop |  | {ifo}:PSL-ISS_TRANSFER1B
transfer 2A | 1 |  |  | 16 | 1 | ADC_0_6 | signal infront of summation point in second loop |  | {ifo}:PSL-ISS_TRANSFER2A
transfer 2B | 1 |  |  | 16 | 1 | ADC_0_7 | signal behind summation point in second loop |  | {ifo}:PSL-ISS_TRANSFER2B
A.1.2.1 second loop whitening
DB9 Connector 1 (Out1)
PD1 | 1 |  |  | 16 | floor | ADC_0_20 |  |  | {ifo}:PSL-ISS_SECONDLOOP_PD1
PD2 | 1 |  |  | 16 | floor | ADC_0_21 |  |  | {ifo}:PSL-ISS_SECONDLOOP_PD2
PD3 | 1 |  |  | 16 | floor | ADC_0_22 |  |  | {ifo}:PSL-ISS_SECONDLOOP_PD3
PD4 | 1 |  |  | 16 | floor | ADC_0_23 |  |  | {ifo}:PSL-ISS_SECONDLOOP_PD4
DB9 Connector 2 (Out2)
PD5 | 1 |  |  | 16 | floor | ADC_0_24 |  |  | {ifo}:PSL-ISS_SECONDLOOP_PD5
PD6 | 1 |  |  | 16 | floor | ADC_0_25 |  |  | {ifo}:PSL-ISS_SECONDLOOP_PD6
PD7 | 1 |  |  | 16 | floor | ADC_0_26 |  |  | {ifo}:PSL-ISS_SECONDLOOP_PD7
PD8 | 1 |  |  | 16 | floor | ADC_0_27 |  |  | {ifo}:PSL-ISS_SECONDLOOP_PD8
A.1.2.2 second loop QPD
DB9 Connector 1 (Out1)
SEG1 | 1 |  |  | 16 | ISC | ADC_0_16 | PD array QPD Segment 1 |  | {ifo}:PSL-ISS_SECONDLOOP_QPD_SEG1
SEG2 | 1 |  |  | 16 | ISC | ADC_0_17 | PD array QPD Segment 2 |  | {ifo}:PSL-ISS_SECONDLOOP_QPD_SEG2
SEG3 | 1 |  |  | 16 | ISC | ADC_0_18 | PD array QPD Segment 3 |  | {ifo}:PSL-ISS_SECONDLOOP_QPD_SEG3
SEG4 | 1 |  |  | 16 | ISC | ADC_0_19 | PD array QPD Segment 4 |  | {ifo}:PSL-ISS_SECONDLOOP_QPD_SEG4
A.1.2.3 second loop servo
DB9 Connector 1 (Out1)
PD 5-8 SUM | 1 |  |  | 16 | 2 | ADC_0_12 | sum of PDs 5-8 |  | {ifo}:PSL-ISS_SECONDLOOP_PD_58_SUM
ref signal mon | 1 |  |  | 16 | 2 | ADC_0_13 | reference voltage readback |  | {ifo}:PSL-ISS_SECONDLOOP_REFSIGNAL_MON
spare | 1 |  |  | 16 | 2 | ADC_0_14
PD 1-4 SUM | 1 |  |  | 16 | 2 | ADC_0_15 | sum of PDs 1-4 |  | {ifo}:PSL-ISS_SECONDLOOP_PD_14_SUM
DB9 Connector 2 (In1)
ref signal | 1 |  |  | 16 | 2 | DAC_0_8 | reference voltage |  | {ifo}:PSL-ISS_SECONDLOOP_REF_SIGNAL
PD switch | 1 |  |  | 16 | 2 | DAC_0_9 | in loop PD switch |  | {ifo}:PSL-ISS_SECONDLOOP_PD_SW
servo on/off | 1 |  |  | 16 | 2 | DAC_0_10 | servo on/off |  | {ifo}:PSL-ISS_SECONDLOOP_SERVO_ON
gain | 1 |  |  | 16 | 2 | DAC_0_11 | variable gain |  | {ifo}:PSL-ISS_SECONDLOOP_GAIN
DB9 Connector 3 (In2)
boost on/off | 1 |  |  | 16 | 2 | DAC_0_12 | boost on/off |  | {ifo}:PSL-ISS_SECONDLOOP_BOOST_ON
int on/off | 1 |  |  | 16 | 2 | DAC_0_13 | integrator on/off |  | {ifo}:PSL-ISS_SECONDLOOP_INT_ON
spare | 1 |  |  | 16 | 2 | DAC_0_14
spare | 1 |  |  | 16 | 2 | DAC_0_15
A1.3 frequency stabilization FSS
TTFSS fieldbox
DB9 Connector (Out1)
IN2 |  |  |  |  | 1 | ADC_1_0 | FSS loop errorpoint |  | {ifo}:PSL-FSS_MIXER
IN1 |  |  |  |  | 1 | ADC_1_1 | LEMO input IN1 at TTFSS fieldbox |  | {ifo}:PSL-FSS_IN1
FastM_OUT |  |  |  | 16 | 1 | ADC_1_2 | monitor of NPRO PZT signal |  | {ifo}:PSL-FSS_FAST_MON
PCM_OUT |  |  |  | 12 | 1 | ADC_1_3 | monitor of signal to phase correcting pockels cell |  | {ifo}:PSL-FSS_PC_MON
DB9 Connector (Out2)
LOMOUT |  |  |  | 12 | 1 | ADC_1_4 | strength of local oscilator signal |  | {ifo}:PSL-FSS_LO_POWER_MON
TEST2 |  |  |  |  | 1 | ADC_1_5 | LEMO input TEST2 at TTFSS fieldbox |  | {ifo}:PSL-FSS_TEST2_MON
RFPDDC |  |  |  | 16 | 1 | ADC_1_6 | DC signal of photodiode in reflection of reference cavity |  | {ifo}:PSL-FSS_RFPD_DC
TPD |  |  |  | 16 | 1 | ADC_1_7 | DC signal of photodiode in transmission of reference cavity |  | {ifo}:PSL-FSS_TPD_DC
DB9 Connector (In1)
CGI |  |  |  | 16 | 1 | DAC_1_0 | common gain |  | {ifo}:PSL-FSS_COMMON_GAIN
OSI |  |  |  | 16 | 1 | DAC_1_1 | signal to be added into error point |  | {ifo}:PSL-FSS_MIXER_OFS
FGI |  |  |  | 16 | 1 | DAC_1_2 | fast path gain (NPRO PZT) |  | {ifo}:PSL-FSS_FAST_GAIN
SLOWI |  |  |  | 16 | 1 | DAC_1_3 | signal given to the SLOW BNC output of TTFSS (connected to RAMP IN) |  | {ifo}:PSL-FSS_FAST_RAMP_INJECTION
DB9 Connector (In2)
TST1E |  |  |  |  | 1 | DAC_1_4 | enables BNC TEST1 IN (summation at output of MIXER (at input of LP filter) |  | {ifo}:PSL-FSS_TEST1_ON
TST2E |  |  |  |  | 1 | DAC_1_5 | enables BNC TEST2 IN (summation at error point, at output of MIXER LP filter) |  | {ifo}:PSL-FSS_TEST2_ON
RMPE |  |  |  |  | 1 | DAC_1_6 | open/close control loop after error point and infront of ramp injection in fast path |  | {ifo}:PSL-FSS_LOOP_CLOSED
TEMP |  |  |  |  | 1 | DAC_1_7 | signal to LEMO TEMP at TTFSS FB |  | {ifo}:PSL-FSS_NPRO_TEMP
VCO fieldbox | only installed at AEI (low noise VCO at the sites communicates via Beckhoff)
DB9 Connector (Out1)
VCO_DETPWR |  |  |  |  | 1 | ADC_1_8 | VCO power monitor |  | {ifo}:PSL-FSS_VCO_POWER_MON
[5 empty] | 1 | ADC_1_9
[5 empty] | 1 | ADC_1_10
[5 empty] | 1 | ADC_1_11
DB9 Connector (In2)
MODLEVEL |  |  |  |  | 1 | DAC_1_12 | sets the level of the rf-signal send from the VCO to the FSS-AOM |  | {ifo}:PSL-FSS_VCO_MODLEVEL
TEST_SW |  |  |  |  | 1 | DAC_1_13 | enables the TEST input (BNC at VCO box) |  | {ifo}:PSL-FSS_VCO_TEST_ON
WIDE_SW |  |  |  |  | 1 | DAC_1_14 | enables the WIDE input (BNC at VCO box, wideband actuator signal from MC loop) |  | {ifo}:PSL-FSS_WIDE_ON
A1.4 diagnostic breadboard DBB
DB25 Connector 1
QPD1-QX | 1 |  |  |  | 2 | ADC_3_0 | quadrant phtotodiode 1 dx -DC |  | {ifo}:PSL-DBB_QPD1_QX
QPD1-QY | 1 |  |  |  | 2 | ADC_3_1 | quadrant phtotodiode 1 dy -DC |  | {ifo}:PSL-DBB_QPD1_QY
QPD1-QS | 1 |  |  |  | 2 | ADC_3_2 | quadrant phtotodiode 1 sum -DC |  | {ifo}:PSL-DBB_QPD1_QS
QPD2-QX | 1 |  |  |  | 2 | ADC_0_3 | quadrant phtotodiode 1 dx -DC |  | {ifo}:PSL-DBB_QPD2_QX
QPD2-QY | 1 |  |  |  | 2 | ADC_3_4 | quadrant phtotodiode 1 dy -DC |  | {ifo}:PSL-DBB_QPD2_QY
QPD2-QS | 1 |  |  |  | 2 | ADC_3_5 | quadrant phtotodiode 1 sum -DC |  | {ifo}:PSL-DBB_QPD2_QS
RIN-DC | 1 |  |  |  | 2 | ADC_3_6 | power noise sensing photodiode DC |  | {ifo}:PSL-DBB_RPD_DC
TPD-0DB | 1 |  |  |  | 2 | ADC_3_7 | photodiode in transmission of cavity (not amplified) |  | {ifo}:PSL-DBB_TPD_0DB
DB25 Connector 2 | [8 empty] | DB25 Connector 2
QPD1-DX | 1 |  |  |  | 2 | ADC_3_8 | quadrant phtotodiode 1 dx - demodulated |  | {ifo}:PSL-DBB_QPD1_DX
QPD1-DY | 1 |  |  |  | 2 | ADC_3_9 | quadrant phtotodiode 1 dy - demodulated |  | {ifo}:PSL-DBB_QPD1_DY
QPD-DS | 1 |  |  |  | 2 | ADC_3_10 | quadrant phtotodiode 1 sum - demodulated |  | {ifo}:PSL-DBB_QPD_DS
QPD2-DX | 1 |  |  |  | 2 | ADC_0_11 | quadrant phtotodiode 1 dx - demodulated |  | {ifo}:PSL-DBB_QPD2_DX
QPD2-DY | 1 |  |  |  | 2 | ADC_3_12 | quadrant phtotodiode 1 dy - demodulated |  | {ifo}:PSL-DBB_QPD2_DY
PZTMON | 1 |  |  |  | 2 | ADC_3_13 | actuator signal of locking loop with signal conditioning |  | {ifo}:PSL-DBB_PZTMON
HVMON | 1 |  |  |  | 2 | ADC_3_14 | actuator signal of locking loop |  | {ifo}:PSL-DBB_HVMON
RPD_AC | 1 |  |  |  | 2 | ADC_3_15 | power noise sensing photodiode AC |  | {ifo}:PSL-DBB_RPD_AC
DB25 Connector 3 | [8 empty] | DB25 Connector 3
TPD-40DB | 1 |  |  |  | 2 | ADC_3_16 | photodiode in transmission of cavity (40dB gain output) |  | {ifo}:PSL-DBB_TPD_40DB
TPD-80DB | 1 |  |  |  | 2 | ADC_3_17 | photodiode in transmission of cavity (80db gain output) |  | {ifo}:PSL-DBB_TPD_80DB
SHUTTER-REQ | 1 |  |  |  | 2 | ADC_3_18 | close shutter request |  | {ifo}:PSL-DBB_SHUTTER_REQ
AUX-INPUT1 | 1 |  |  |  | 2 | ADC_0_19 |  |  | {ifo}:PSL-DBB_AUX_INPUT1
AUX-INPUT2 | 1 |  |  |  | 2 | ADC_3_20 |  |  | {ifo}:PSL-DBB_AUX_INPUT2
AUX-INPUT3 | 1 |  |  |  | 2 | ADC_3_21 |  |  | {ifo}:PSL-DBB_AUX_INPUT3
AUX-INPUT4 | 1 |  |  |  | 2 | ADC_3_22 |  |  | {ifo}:PSL-DBB_AUX_INPUT4
RMT-READBACK | 1 |  |  |  | 2 | ADC_3_23 |  |  | {ifo}:PSL-DBB_RMT_READBACK
DB25 Connector 4 | [8 empty] | DB25 Connector 4
CTRL0 | 1 |  |  |  | 2 | DAC_3_0 | control signal length control |  | {ifo}:PSL-DBB_CTRL0
CTRL1X | 1 |  |  |  | 2 | DAC_3_1 | control signal alignment control |  | {ifo}:PSL-DBB_CTRL1X
CTRL1Y | 1 |  |  |  | 2 | DAC_3_2 | control signal alignment control |  | {ifo}:PSL-DBB_CTRL1Y
CTRL2X | 1 |  |  |  | 2 | DAC_0_3 | control signal alignment control |  | {ifo}:PSL-DBB_CTRL2X
CTRL2Y | 1 |  |  |  | 2 | DAC_3_4 | control signal alignment control |  | {ifo}:PSL-DBB_CTRL2Y
PHASE | 1 |  |  |  | 2 | DAC_3_5 |  |  | {ifo}:PSL-DBB_PHASE
LENS1 | 1 |  |  |  | 2 | DAC_3_6 | actuates mode matching lens 1 |  | {ifo}:PSL-DBB_LENS1
LENS2 | 1 |  |  |  | 2 | DAC_3_7 | actuates mode matching lens 2 |  | {ifo}:PSL-DBB_LENS2
DB25 Connector 5 | [8 empty] | DB25 Connector 5
LENS-LATCH | 1 |  |  |  | 2 | DAC_3_8 |  |  | {ifo}:PSL-DBB_LENS-LATCH
SHUTTER_DBB | 1 |  |  |  | 2 | DAC_3_9 |  |  | {ifo}:PSL-DBB_SHUTTER_DBB
RMT | 1 |  |  |  | 2 | DAC_3_10 |  |  | {ifo}:PSL-DBB_RMT
MODOFF | 1 |  |  |  | 2 | DAC_0_11 |  |  | {ifo}:PSL-DBB_MODOFF
SCANMODE | 1 |  |  |  | 2 | DAC_3_12 |  |  | {ifo}:PSL-DBB_SCANMODE
SHUTTER-MULTIPLEX | 1 |  |  |  | 2 | DAC_3_13 |  |  | {ifo}:PSL-DBB_SHUTTER_MULTIPLEX
FILTER-SWITCH1 | 1 |  |  |  | 2 | DAC_3_14 |  |  | {ifo}:PSL-DBB_FILTER_SWITCH1
FILTER-SWITCH2 | 1 |  |  |  | 2 | DAC_3_15 |  |  | {ifo}:PSL-DBB_FILTER_SWITCH2
injection locking is analog loop, lock acquisition is guided by Beckhoff (see A2.3), EPICS used for fast sensing and signal injection (via PMC RT-model)
A1.5 injection locking IJL (EPICS part)
DB9 Connector 4 (Out2)
HV monitor | 1 |  |  | 1 | 2 | ADC_2_12 | monitor of HV signal send to PZT actuator of ILS |  | {ifo}:PSL-ILS_HV_MON
Mixer out monitor | 1 |  |  | 1 | 2 | ADC_2_13 | error signal of injection locking control loop |  | {ifo}:PSL-ILS_MIXER
Local oscilllator power (LODET) | 1 |  |  | 1 | 2 | ADC_2_14 | power of local oscillator used in demodulation electronic |  | {ifo}:PSL-ILS_LO_POWER_MON
"resonant" trigger (TRIGGER) | 1 |  |  | 1 | 2 | ADC_2_15 | DC signal of locking diode, used by autolock to dertermine if HPL is resonant |  | {ifo}:PSL-ILS_TRIGGER
DB9 Connector 5 (Out1)
TF B | 1 |  |  | 16 | 2 | ADC_2_8 | test point for transfer function measurement |  | {ifo}:PSL-ILS_TF_B
Temp | [5 empty] | ADC_2_9 | temperature sensor on PMC
TF A | [5 empty] | ADC_2_10 | test point for transferfunction measurement |  | {ifo}:PSL-ILS_TF_A
Resonant | [5 empty] | ADC_2_11 | binary signals, 1 if auto lock thinks PMC is resonant |  | {ifo}:PSL-ILS_RESONANT_ON
module for HPL photodiodes, RT transferfunction measurements and channels missed in design
A1.6 PSL monitoring Field Box
DB9 Connector 1
OSC photodiode - amplifier power in | 1 |  |  | 16 | 2 | ADC_2_16 | monitoring photodiode in HPL: signal from front-end at output of Faraday Isolator |  | {ifo}:PSL-OSC_PDAMP_DC
OSC photodiode - reverse power (FI reflected) | 1 |  |  | 16 | 2 | ADC_2_17 | monitoring photodiode in HPL: power in reverse direction split of at FI |  | {ifo}:PSL-OSC_PDISO_DC
OSC photodiode - Brewster Plate reflected | 1 |  |  | 16 | 2 | ADC_2_18 | monitoring photodiode in HPL: reflection at Brewster Plate in HPL oscillator |  | {ifo}:PSL-OSC_PDBP_DC
OSC photodiode - oscillator internal power | 1 |  |  | 16 | 2 | ADC_2_19 | monitoring photodiode in HPL: circulating light power in HPL oscillator |  | {ifo}:PSL-OSC_PDINT_DC
DB9 Connector 2
OSC photodiode - amplifier power in -AC | 1 |  |  | 16 | 2 | ADC_2_20 | monitoring photodiode in HPL: signal from front-end at output of Faraday Isolator |  | {ifo}:PSL-OSC_PDAMP_AC
OSC photodiode - reverse power (FI reflected) -AC | 1 |  |  | 16 | 2 | ADC_2_21 | monitoring photodiode in HPL: power in reverse direction split of at FI |  | {ifo}:PSL-OSC_PDISO_AC
OSC photodiode - Brewster Plate reflected -AC | 1 |  |  | 16 | 2 | ADC_2_22 | monitoring photodiode in HPL: reflection at Brewster Plate in HPL oscillator |  | {ifo}:PSL-OSC_PDBP_AC
OSC photodiode - oscillator internal power -AC | 1 |  |  | 16 | 2 | ADC_2_23 | monitoring photodiode in HPL: circulating light power in HPL oscillator |  | {ifo}:PSL-OSC_PDINT_AC
DB9 Connector 3
PMC - reflected port power meter | 1 |  |  | 16 | 2 | ADC_2_24 | power in reflection of PMC (power meter used for installation and maintenance) |  | {ifo}:PSL-PWR_PMC_REFL
PMC - transmitted port power meter (installation only) | 1 |  |  | 16 | 2 | ADC_2_25 | power in transmission of PMC (power meter used for installation and maintenance) |  | {ifo}:PSL-PWR_PMC_TRANS
HPL-power monitoring diode - DC | 1 |  |  | 16 | 2 | ADC_2_26 | monitor photodiode (PD001) at output of HPL, main HPL power monitor |  | {ifo}:PSL-PWR_HPL_DC
HPL-power monitoring diode - AC | 1 |  |  | 16 | 2 | ADC_2_27 | monitor photodiode (PD001) at output of HPL, main HPL power monitor |  | {ifo}:PSL-PWR_HPL_AC
[10 empty] | {ifo}:PSL-PWR_PMC_SUM | transmitted plus reflected power
DB9 Connector 4
water flow sensor for power meters and beamdumps on PSL table | 1 |  |  | 16 | 2 | ADC_2_28 | used to detect too low water flow through beam dumps and power meters |  | {ifo}:PSL-MIS_FLOW
NPRO RRO monintoring PD - DC | 1 |  |  | 16 | 2 | ADC_2_29 | monitor photodiode at input of 35W amplifier: senses NPRO power |  | {ifo}:PSL-PWR_NPRO
NPRO RRO monintoring PD - RRO detect | 1 |  |  | 16 | 2 | ADC_2_30 | binary signal for NPRO rf- power (in band around relaxation oscillation frequency) |  | {ifo}:PSL-MIS_NPRO_RRO
spare | 1 |  |  | 16 | 2 | ADC_2_31 |  |  | {ifo}:PSL-MIS_SPARE
module for auxiliary inputs and outputs
A1.7 PSL input/output Field Box
DB9 Connector 1 (Out1)
ADC1 | 1 |  |  | 16 | 2 | ADC_1_12
ADC2 | 1 |  |  | 16 | 2 | ADC_1_13
ADC3 | 1 |  |  | 16 | 2 | ADC_1_14
ADC4 | 1 |  |  | 16 | 2 | ADC_1_15
DB9 Connector 2 (Out2)
ADC5 | 1 |  |  | 16 | 2 | ADC_1_16
ADC6 | 1 |  |  | 16 | 2 | ADC_1_17
ADC7 | 1 |  |  | 16 | 2 | ADC_1_18
ADC8 | 1 |  |  | 16 | 2 | ADC_1_19
DB9 Connector 3 (In1)
DAC1 | 1 |  |  | 16 | 2 | DAC_1_16
DAC2 | 1 |  |  | 16 | 2 | DAC_1_17
DAC3 | 1 |  |  | 16 | 2 | DAC_1_18
DAC4 | 1 |  |  | 16 | 2 | DAC_1_19
DB9 Connector 4 (In2)
DAC5 | 1 |  |  | 16 | 2 | DAC_2_12
DAC6 | 1 |  |  | 16 | 2 | DAC_2_13
DAC7 | 1 |  |  | 16 | 2 | DAC_2_14
DAC8 | 1 |  |  | 16 | 2 | DAC_2_15
A2 / Beckhoff interface LVEA | for information on acquired channels available in EPICS see interface D1.1 in "Digital (D)" spreadsheet
A2.0 / NPRO Diagnostic | for information on acquired channels available in EPICS see interface D1.1 in "Digital (D)" spreadsheet
A2.1 / Frontend Control | for information on acquired channels available in EPICS see interface D1.1 in "Digital (D)" spreadsheet
A2.2 / HPL Control | injection locking is analog loop, lock acquisition is guided by Beckhoff , EPICS used for fast sensing and test point injection (see A1.5)
A2.3 / Injection locking (Beckhoff part)
DB9 Connector 1 | 1 | c / 24V |  | 1
Enable locking (lock on) | 1 | c / 24V |  | 1
Ramp on/off (ramp on) | 1 | c / 24V |  | 1 |  |  | high gain, only relevant for PMC locking, switches between full and reduced power mode
sw1 | 1 | c / 24V |  | 1 |  |  | enables summation of TEST2 input
sw2
DB9 Connector 2 | 1 | c / -10 ..10V |  | 16
Set power threshold (ref) | 1 | c / -10 ..10V |  | 16
Set error point offset (in offset) | 1 | c / -10 ..10V |  | 16
Set gain
Lemo 00 | 1 | s / -10..10V |  | 16
PTZ voltage monitor
Lemo 00 (prelim) | 1 | s / 5V |  | 1
Relock counter
A3 / Beckhoff - Diodeboxes | for information on acquired channels available in EPICS see interface D1.1 in "Digital (D)" spreadsheet
A3.0 / Laser Diode Box Front End | 7 |  |  | 12 |  |  | Set Temperature LD 1-4 | rAmpDiodeTempSet[1-4]
ST_LD0.1-4 | 7 | s / PT1000 | 20 Hz | 16 |  |  | Input Temperature LD 1-4 | rAmpDiodeTempAct[1-4]
IT_LD0.1-4 | 7 | c / 0-10V | 20 Hz | 12 |  |  | Output Temperature to Pelt.Amp. LD 1-4 | rAmpDiodeTempCtrl[1-4]
OT_LD0.1-4 | 2 | c |  | 12 |  |  | Set Current LD 1-2/3-4 | rAmpDiodeCurrentSet[1-2]
SC_LD0.1-2/3-4 | 2 | s / PT1000 |  | 12 |  |  | Input Current LD 1-2/3-4 | rAmpDiodeCurrentAct[1-2]
IC_LD0.1-2/3-4 | 2 | c / 0-10V | 20 Hz | 12 |  |  | Output Current to Diode Supply LD 1-2/3-4 | rAmpDiodeCurrentCtrl[1-2]
OC_LD0.1-2/3-4 | 2 | s | 20 Hz | 12 |  |  | Input, Modulation Voltage frontend LD1-4 | rAmpModulationVoltage[1-2]
IV_LD0.1-2/3-4 | 1 | s / PT1000 | 20 Hz | 16 |  |  | Input Heatsink Temp frontend diode box | rAmpDBHeatsinkTempAct
IT_DBHeasinkTemp | 1 | s | 20 Hz | 1 |  |  | Input Diode TempGuard | bAmpTempGuard
I_TempGuards | 1 | s | 20 Hz | 1 |  |  | Input | bAmpDBInterlock
I_DBInterlock | 1 | c | 20 Hz | 1 |  |  | Output | bAmpEnablePowerSupplies
O_PowerSuppliesFrontEnd | for information on acquired channels available in EPICS see interface D1.1 in "Digital (D)" spreadsheet
A3.1 / Laser Diode Box 1 | 7 |  |  | 12 |  |  | Set Temperature LD 1-7 | rOscDiodeTempSet[1,1-7]
Set temp DB1_LD1-7 | 7 | s / PT1000 | 100 Hz | 16 |  |  | Input Temperature LD 1-7 | rOscDiodeTempAct[1,1-7]
Act temp DB1_LD 1-7 | 7 | c / 0-10V | 100 Hz | 12 |  |  | Output Temperature to Pelt.Amp. LD 1-7 | rOscDiodeTempCtrl[1,1-7]
Output TEC DB1_LD 1-7 | 1 | c |  | 12 |  |  | Set Current LD 1-7 | rOscDiodeCurrentSet[1]
set current DB1 | 1 |  | 100 Hz | 12 |  |  | Input Current LD 1-7 | rOscDiodeCurrentAct[1]
act current DB1 | 1 | c / 0-5V | 100 Hz | 12 |  |  | Output Current from Diode Supply LD 1-7 | rOscDiodeCurrentCtrl[1]
Output current DB1 | 2 | s / PT1000
Act temp heatsink DB1_1-2 | for information on acquired channels available in EPICS see interface D1.1 in "Digital (D)" spreadsheet
A3.2 / Laser Diode Box 2 | 7 | [5 empty] | Set Temperature LD 8-14 | rOscDiodeTempSet[2,1-7]
Set temp_DB2_LD1-7 | 7 | s / PT1000 | 100 Hz | 16 |  |  | Input Temperature LD 8-14 | rOscDiodeTempAct[2,1-7]
Act temp DB2_LD 2-8 | 7 | c / 0-10V | 100 Hz | 12 |  |  | Output Temperature to Pelt.Amp. LD 8-14 | rOscDiodeTempCtrl[2,1-7]
Output TEC DB2_LD 2-8 | 1 | c |  |  |  |  | Set Current LD 8-14 | rOscDiodeCurrentSet[2]
set current DB2 | 1 | [5 empty] | Input Current LD 8-14 | rOscDiodeCurrentAct[2]
act current DB2 | 1 | c / 0-5V | 100 Hz |  |  |  | Output Current to Diode Supply LD 8-14 | rOscDiodeCurrentCtrl[2]
Output current DB2 | 2 | s / PT1000
Act temp heatsink DB2_1-2 | for information on acquired channels available in EPICS see interface D1.1 in "Digital (D)" spreadsheet
A3.3 / Laser Diode Box 3 | 7 | [5 empty] | Set Temperature LD 15-21 | rOscDiodeTempSet[3,1-7]
Set temp_DB3_LD1-7 | 7 | s / PT1000 | 100 Hz | 16 |  |  | Input Temperature LD 15-21 | rOscDiodeTempAct[3,1-7]
Act temp DB3_LD 2-8 | 7 | c / 0-10V | 100 Hz | 12 |  |  | Output Temperature to Pelt.Amp. LD 15-21 | rOscDiodeTempCtrl[3,1-7]
Output TEC DB3_LD 2-8 | 1 | c |  |  |  |  | Set Current LD 15-21 | rOscDiodeCurrentSet[3]
set current DB3 | 1 | [5 empty] | Input Current LD 15-21 | rOscDiodeCurrentAct[3]
act current DB3 | 1 | c / 0-5V | 100 Hz |  |  |  | Output Current to Diode Supply LD 15-21 | rOscDiodeCurrentCtrl[3]
Output current DB3 | 2 | s / PT1000
Act temp heatsink DB3_1-2 | for information on acquired channels available in EPICS see interface D1.1 in "Digital (D)" spreadsheet
A3.4 / Laser Diode Box 4 | 7 | [5 empty] | Set Temperature LD 22-28 | rOscDiodeTempSet[4,1-7]
Set temp_DB4_LD1-7 | 7 | s / PT1000 | 100 Hz | 16 |  |  | Input Temperature LD 22-28 | rOscDiodeTempAct[4,1-7]
Act temp DB4_LD 2-8 | 7 | c / 0-10V | 100 Hz | 12 |  |  | Output Temperature to Pelt.Amp. LD 22-28 | rOscDiodeTempCtrl[4,1-7]
Output TEC DB4_LD 2-8 | 1 | c |  |  |  |  | Set Current LD 22-28 | rOscDiodeCurrentSet[4]
set current DB4 | 1 | [5 empty] | Input Current LD 22-28 | rOscDiodeCurrentAct[4]
act current DB4 | 1 | c / 0-5V | 100 Hz |  |  |  | Output Current to Diode Supply LD 22-28 | rOscDiodeCurrentCtrl[4]
Output current DB4 | 2 | s / PT1000
Act temp heatsink DB4_1-2 | see "Interlock (I) " spreadsheet
A3.5 / Interlock Box
A4 / Beckhoff - chiller | for information on acquired channels available in EPICS see interface D1.1 in "Digital (D)" spreadsheet
A4.1 / Xtal chiller
DB9 connector | 1 | s / 24V | 1 | 1
Xtal chiller pump status | 1 | c / relais contact | 1 | 1
Enable Xtal chiller
DB9 connector |  | s / RS232 | 1
Xtal chiller readout | for information on acquired channels available in EPICS see interface D1.1 in "Digital (D)" spreadsheet
A4.1 / Diode chiller
DB9 connector | 1 | c / relais contact | 1 | 1
Enable diode chiller
DB9 connector |  | s / RS232 | 1
diode chiller readout
A5 / PSL EPICS in laser diode room | optional RT module for fast sensing, fast LD control and transferfunction measurements in laser diode room
A5.1 RT control chassis #3
DB9 Connector 1 | 1 |  |  | 16 | 3
TPout 1 | 1 |  |  | 16 | 3
TPout 2 | 1 |  |  | 16 | 3
TPout 3 | 1 |  |  | 16 | 3
TPout 4
DB9 Connector 1 | 1 |  |  | 16 | 3
TPin 1 | 1 |  |  | 16 | 3
TPin 2 | 1 |  |  | 16 | 3
TPin 3 | 1 |  |  | 16 | 3
TPin 4

## Digital (D)
EPICS Name | Laser computer name [Opc-Server].[Beckhoff variable] | Type | OPC group name | Polling speed [ms]
[5 empty] | # Channels | Global control | bandwidth | bits | units | calibration | ADC # | DAC # | DIO # | Comments | Comments
D1
D1.1 / front end (Beckhoff)
{ifo}:PSL-AMP_DBHTSNKOVRTEMP | AdL-REF-OPC1.bAmpDBHeatsinkOvertempError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Heatsink Temperature Error
{ifo}:PSL-AMP_DTEMPERR | AdL-REF-OPC1.bAmpDiodeTempError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Diode Temperature Error
{ifo}:PSL-AMP_ENABLEPWRDOG | AdL-REF-OPC1.bAmpEnablePowerWatchdog | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Enable Power Watchdog (>15% deviation)
{ifo}:PSL-AMP_ERR | AdL-REF-OPC1.bAmpError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Error state positive
{ifo}:PSL-AMP_LIDCLOSED | AdL-REF-OPC1.bAmpLidClosed | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Lid Closed
{ifo}:PSL-AMP_LIDERR | AdL-REF-OPC1.bAmpLidError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Lid opened without Lid override active
{ifo}:PSL-AMP_MANUALMODE | AdL-REF-OPC1.bAmpManualMode | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Manual Mode (enables individual control of NPRO and amplifier currents)
{ifo}:PSL-AMP_NPROOK | AdL-REF-OPC1.bAmpNPRORunning | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. NPRO Running (from NPRO internal diagnostic)
{ifo}:PSL-AMP_LIDERR_OVR | AdL-REF-OPC1.bAmpOverrideLidInterlock | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Lid Interlock Override acitve
{ifo}:PSL-AMP_WATCHDOG | AdL-REF-OPC1.bAmpPowerWatchdogError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Power Watchdog triggered
{ifo}:PSL-AMP_SHUTTERCLSD | AdL-REF-OPC1.bAmpShutterClosed | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp.Shutter Closed INPUT signal
{ifo}:PSL-AMP_SHUTTEROPEN | AdL-REF-OPC1.bAmpShutterOpen | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Shutter Open INPUT Signal
{ifo}:PSL-AMP_SYSTEMOK | AdL-REF-OPC1.bAmpSystemRunning | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. System in automatic mode - Running
{ifo}:PSL-AMP_WARN | AdL-REF-OPC1.bAmpWarning | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Warning NPRO or water flow
{ifo}:PSL-AMP_LOWFLOW | AdL-REF-OPC1.bAmpWaterFlowWarning | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Water Flow Warning (< x l/min)
{ifo}:PSL-AMP_FLOWERR | AdL-REF-OPC1.bAmpWaterFlowError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Water Flow Error (< x l/min)
{ifo}:PSL-AMP_XTALOVRTEMP | AdL-REF-OPC1.bAmpXtalOvertempError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. crystal temperature error
{ifo}:PSL-AMP_LIDCOUNT | AdL-REF-OPC1.iAmpLidCounter | INT | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. Lid # of opening
{ifo}:PSL-AMP_HRS | AdL-REF-OPC1.iAmpOpHrs | INT | AdL-REF_Read.db | 125 | 1 | [8 empty] | Amp. # of operation hours
{ifo}:PSL-AMP_DBHTSNKTEMP | AdL-REF-OPC1.rAmpDBHeatsinkTempAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | Amp. Heatsink Temperature
{ifo}:PSL-AMP_DCUR1 | AdL-REF-OPC1.rAmpDiodeCurrentAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | Amp. Diode current (Diode 1+2)
{ifo}:PSL-AMP_DCUR2 | AdL-REF-OPC1.rAmpDiodeCurrentAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | Amp. Diode current (Diode 3+4)
{ifo}:PSL-AMP_DCURSET1 | AdL-REF-OPC1.rAmpDiodeCurrentSet[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | Amp. Diode current set point (Diode 1+2)
{ifo}:PSL-AMP_DCURSET2 | AdL-REF-OPC1.rAmpDiodeCurrentSet[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | Amp. Diode current set point (Diode 3+4)
{ifo}:PSL-AMP_D1PWR | AdL-REF-OPC1.rAmpDiodePowerAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | Amp. Diode power (internal monitor diode)
{ifo}:PSL-AMP_D2PWR | AdL-REF-OPC1.rAmpDiodePowerAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | Amp. Diode power (internal monitor diode)
{ifo}:PSL-AMP_D3PWR | AdL-REF-OPC1.rAmpDiodePowerAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | Amp. Diode power (internal monitor diode)
{ifo}:PSL-AMP_D4PWR | AdL-REF-OPC1.rAmpDiodePowerAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | Amp. Diode power (internal monitor diode)
{ifo}:PSL-AMP_D1TEMP | AdL-REF-OPC1.rAmpDiodeTempAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | Amp. Diode temperature
{ifo}:PSL-AMP_D2TEMP | AdL-REF-OPC1.rAmpDiodeTempAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | Amp. Diode temperature
{ifo}:PSL-AMP_D3TEMP | AdL-REF-OPC1.rAmpDiodeTempAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | Amp. Diode temperature
{ifo}:PSL-AMP_D4TEMP | AdL-REF-OPC1.rAmpDiodeTempAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | Amp. Diode temperature
{ifo}:PSL-AMP_D1TEC | AdL-REF-OPC1.rAmpDiodeTempCtrl[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | Amp. Diode temp. control signal
{ifo}:PSL-AMP_D2TEC | AdL-REF-OPC1.rAmpDiodeTempCtrl[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | Amp. Diode temp. control signal
{ifo}:PSL-AMP_D3TEC | AdL-REF-OPC1.rAmpDiodeTempCtrl[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | Amp. Diode temp. control signal
{ifo}:PSL-AMP_D4TEC | AdL-REF-OPC1.rAmpDiodeTempCtrl[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | Amp. Diode temp. control signal
{ifo}:PSL-AMP_D1TEMPSET | AdL-REF-OPC1.rAmpDiodeTempSet[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | Amp. Diode temp. set point
{ifo}:PSL-AMP_D2TEMPSET | AdL-REF-OPC1.rAmpDiodeTempSet[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | Amp. Diode temp. set point
{ifo}:PSL-AMP_D3TEMPSET | AdL-REF-OPC1.rAmpDiodeTempSet[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | Amp. Diode temp. set point
{ifo}:PSL-AMP_D4TEMPSET | AdL-REF-OPC1.rAmpDiodeTempSet[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | Amp. Diode temp. set point
{ifo}:PSL-AMP_DVOLT1 | AdL-REF-OPC1.rAmpDiodeVoltageAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | Amp. Diode voltage (Diode 1+2)
{ifo}:PSL-AMP_DVOLT2 | AdL-REF-OPC1.rAmpDiodeVoltageAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | Amp. Diode voltage (Diode 3+4)
{ifo}:PSL-AMP_PWR1 | AdL-REF-OPC1.rAmpLaserPowerAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | Amp. power monitor (behind stage 1)
{ifo}:PSL-AMP_PWR2 | AdL-REF-OPC1.rAmpLaserPowerAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | Amp. power monitor (behind stage 2)
{ifo}:PSL-AMP_PWR3 | AdL-REF-OPC1.rAmpLaserPowerAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | Amp. power monitor (behind stage 3)
{ifo}:PSL-AMP_FLOW | AdL-REF-OPC1.rAmpWaterFlowAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | l/min |  |  |  |  | Amp. cooling water flow rate
{ifo}:PSL-AMP_XTALTEMP | AdL-REF-OPC1.rAmpXtalTempAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | Amp. crystal temperature
{ifo}:PSL-NPRO_TEMPGUARD | AdL-REF-OPC1.bAmpNPROTempGuardError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | NPRO Temperature Error
{ifo}:PSL-NPRO_HRS | AdL-REF-OPC1.iNproOpHrs | INT | AdL-REF_Read.db | 125 | 1 |  |  |  | h |  |  |  |  | NPRO operation hours
{ifo}:PSL-NPRO_D1PWR | AdL-REF-OPC1.rAmpNPRODiag[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | NPRO LD1 power monitor, sensed
{ifo}:PSL-NPRO_D2PWR | AdL-REF-OPC1.rAmpNPRODiag[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | NPRO LD2 power monitor, sensed
{ifo}:PSL-NPRO_D1TEMPERR | AdL-REF-OPC1.rAmpNPRODiag[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | NPRO LD1 Temperatures , error signal
{ifo}:PSL-NPRO_D2TEMPERR | AdL-REF-OPC1.rAmpNPRODiag[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | NPRO LD2 Temperatures , error signal
{ifo}:PSL-NPRO_NEMON | AdL-REF-OPC1.rAmpNPRODiag[5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | NPRO Noise Eater monitor (NE DC signal)
{ifo}:PSL-NPRO_XTALTEMPERR | AdL-REF-OPC1.rAmpNPRODiag[6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | NPRO Xtal Temperature, error signal
{ifo}:PSL-NPRO_D1TEMPSET | AdL-REF-OPC1.rAmpNPRODiag[7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | NPRO LD1 Temperatures, set
{ifo}:PSL-NPRO_D2TEMPSET | AdL-REF-OPC1.rAmpNPRODiag[8] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | NPRO LD2 Temperatures, set
{ifo}:PSL-NPRO_D1TEMP | AdL-REF-OPC1.rAmpNPRODiag[9] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | NPRO LD Temperatures, sensed
{ifo}:PSL-NPRO_D2TEMP | AdL-REF-OPC1.rAmpNPRODiag[10] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | NPRO LD Temperatures, sensed
{ifo}:PSL-NPRO_XTALTEMPSET | AdL-REF-OPC1.rAmpNPRODiag[11] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | NPRO Xtal Temperature, set
{ifo}:PSL-NPRO_XTALTEMP | AdL-REF-OPC1.rAmpNPRODiag[12] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | NPRO Xtal Temperature, sensed
{ifo}:PSL-NPRO_CURSET | AdL-REF-OPC1.rAmpNPRODiag[13] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | NPRO LD current, set
{ifo}:PSL-NPRO_CUR | AdL-REF-OPC1.rAmpNPRODiag[14] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | NPRO LD current, sensed
D1.2 / HPL (Beckhoff)
{ifo}:PSL-EPICSALARM | AdL-REF-OPC1.bEpicsExternalAlarm | BOOL | AdL-REF_Write.db | 125 | 1 | [8 empty] | EPICS signal to close the high power laser shutter
{ifo}:PSL-IL_DBTEMPGUARDS | AdL-REF-OPC1.bILDBTempGuards | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | hardware HPO laser diode temperature interlock
{ifo}:PSL-IL_DCHILFLOW | AdL-REF-OPC1.bILDChilFlow | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | hardware diode chiller flow interlock
{ifo}:PSL-IL_DIODERMEMSTOP | AdL-REF-OPC1.bILDiodeRoomEmergencyStop | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | hardware laser diode room interlock box safety button
{ifo}:PSL-IL_DIODERMFACILITY | AdL-REF-OPC1.bILDiodeRoomFacility | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | hardware laser diode room facility interlock
{ifo}:PSL-IL_DIODERMKEYLOCK | AdL-REF-OPC1.bILDiodeRoomSafetyKeylock | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | hardware laser diode room, interlock-box key interlock
{ifo}:PSL-IL_AMPTEMPGUARDS | AdL-REF-OPC1.bILFrontendTempGuards | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | hardware amplifier laser diode temperature interlock
{ifo}:PSL-IL_LVEAEMSTOP | AdL-REF-OPC1.bILLveaEmergencyStop | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | hardware laser area enclosure safety buttom interlock
{ifo}:PSL-IL_LVEAFACILITY | AdL-REF-OPC1.bILLveaFacility | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | hardware LVEA facility interlock
{ifo}:PSL-IL_LVEAKEYLOCK | AdL-REF-OPC1.bILLveaSafetyKeylock | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | hardware LVEA control box key interlock
{ifo}:PSL-IL_OK | AdL-REF-OPC1.bILOK | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | no interlock
{ifo}:PSL-IL_TECSTATUS | AdL-REF-OPC1.bILTECStatus | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | HPO laser diode temperature contol power supplies on
{ifo}:PSL-IL_TWINSAFEFBERR | AdL-REF-OPC1.bILTwinSafeFBError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | ??? Not clear ???
{ifo}:PSL-IL_XCHILFLOW | AdL-REF-OPC1.bILXChilFlow | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | hardware crystal chiller flow interlock
{ifo}:PSL-OSC_COUPLERSTATEERR | AdL-REF-OPC1.bCouplerStateError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | TwinCad interface error
{ifo}:PSL-OSC_DCHILALARM | AdL-REF-OPC1.bDChilAlarm | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | Alarm Diode Chiller
{ifo}:PSL-OSC_LRARANGEWARN | AdL-REF-OPC1.bOscActuatorRangeWarning | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | LRA out of range
{ifo}:PSL-OSC_DBHTSNKOVRTEMP | AdL-REF-OPC1.bOscDBHeatsinkOvertempError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscDBHeatsinkOvertemp
{ifo}:PSL-OSC_DTEMPERR | AdL-REF-OPC1.bOscDiodeTempError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscDiodeOvertemp
{ifo}:PSL-OSC_LRAON | AdL-REF-OPC1.bOscEnableActuator | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscEnableLRA
{ifo}:PSL-OSC_LOCKON | AdL-REF-OPC1.bOscEnableLock | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscEnableLocking
{ifo}:PSL-OSC_PWRDOGON | AdL-REF-OPC1.bOscEnablePowerWatchdog | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscEnablePowerWatchdog
{ifo}:PSL-OSC_FLOWDOGON | AdL-REF-OPC1.bOscEnableFlowWatchdog | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscEnableFlowWatchdog
{ifo}:PSL-OSC_RAMPON | AdL-REF-OPC1.bOscEnableRamp | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscEnableRamp (injection locking)
{ifo}:PSL-OSC_SW2ON | AdL-REF-OPC1.bOscEnableSW2 | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | ??? Not longer used…?
{ifo}:PSL-OSC_ERROR | AdL-REF-OPC1.bOscError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscError
{ifo}:PSL-OSC_VBNOTALIVE | AdL-REF-OPC1.bOscFiberSwitchVBAliveWarning | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscFiberSwitchVB not Alive
{ifo}:PSL-OSC_HPSCLSD | AdL-REF-OPC1.bOscHighPowerShutterClosed | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscHighPowerShutterClosed
{ifo}:PSL-OSC_LIDCLSD | AdL-REF-OPC1.bOscLidClosed | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscLidClosed
{ifo}:PSL-OSC_LIDERR | AdL-REF-OPC1.bOscLidError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscLidError (were opened)
{ifo}:PSL-OSC_LOCKED | AdL-REF-OPC1.bOscLocked | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscLocked
{ifo}:PSL-OSC_LPSCLSD | AdL-REF-OPC1.bOscLowPowerShutterClosed | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscLowPowerShutterClosed
{ifo}:PSL-OSC_LIDERR_OVR | AdL-REF-OPC1.bOscOverrideLidInterlock | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscOverrideLidInterlock
{ifo}:PSL-OSC_PWRMETERLOWERR | AdL-REF-OPC1.bOscPowermeterWaterFlowError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscPowermeterWaterFlowError
{ifo}:PSL-OSC_PWRMETERLOWFLOW | AdL-REF-OPC1.bOscPowermeterWaterFlowWarning | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscPowermeterWaterFlowWarning
{ifo}:PSL-OSC_WATCHDOG | AdL-REF-OPC1.bOscPowerwatchdogError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscPowerwatchdogError
{ifo}:PSL-OSC_WARNING | AdL-REF-OPC1.bOscWarning | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscWarning one other warnings active
{ifo}:PSL-OSC_FLOWERR | AdL-REF-OPC1.bOscWaterFlowError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscWaterFlowError
{ifo}:PSL-OSC_PRESS1ERR | AdL-REF-OPC1.bOscPressure1MaxError | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscPressure1MaxError
{ifo}:PSL-OSC_LOWFLOW | AdL-REF-OPC1.bOscWaterFlowWarning | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscWaterFlowWarning
{ifo}:PSL-OSC_SERVICEMODE | AdL-REF-OPC1.bServiceMode | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | OscServiceMode
{ifo}:PSL-OSC_XCHILALARM | AdL-REF-OPC1.bXChilAlarm | BOOL | AdL-REF_Read.db | 125 | 1 | [8 empty] | XChilAlarm
{ifo}:PSL-OSC_DCHILHRS | AdL-REF-OPC1.iDChilOpHrs | INT | AdL-REF_Read.db | 125 | 1 |  |  |  | h |  |  |  |  | DChilOpHrs
{ifo}:PSL-OSC_DB1HRS | AdL-REF-OPC1.iOscDB1OpHrs | INT | AdL-REF_Read.db | 125 | 1 |  |  |  | h |  |  |  |  | AdL-REF-OPC1.iOscDB1OpHrs
{ifo}:PSL-OSC_DB2HRS | AdL-REF-OPC1.iOscDB2OpHrs | INT | AdL-REF_Read.db | 125 | 1 |  |  |  | h |  |  |  |  | AdL-REF-OPC1.iOscDB2OpHrs
{ifo}:PSL-OSC_DB3HRS | AdL-REF-OPC1.iOscDB3OpHrs | INT | AdL-REF_Read.db | 125 | 1 |  |  |  | h |  |  |  |  | AdL-REF-OPC1.iOscDB3OpHrs
{ifo}:PSL-OSC_DB4HRS | AdL-REF-OPC1.iOscDB4OpHrs | INT | AdL-REF_Read.db | 125 | 1 |  |  |  | h |  |  |  |  | AdL-REF-OPC1.iOscDB4OpHrs
{ifo}:PSL-OSC_LIDCOUNT | AdL-REF-OPC1.iOscLidCounter | INT | AdL-REF_Read.db | 125 | 1 | [8 empty] | AdL-REF-OPC1.iOscLidCounter
{ifo}:PSL-OSC_RELOCKCOUNT | AdL-REF-OPC1.iOscRelockCounter | INT | AdL-REF_Read.db | 125 | 1 | [8 empty] | AdL-REF-OPC1.iOscRelockCounter
{ifo}:PSL-OSC_XCHILHRS | AdL-REF-OPC1.iXChilOpHrs | INT | AdL-REF_Read.db | 125 | 1 |  |  |  | h |  |  |  |  | HPO crystal chiller operating hours
{ifo}:PSL-OSC_DCHILCND | AdL-REF-OPC1.rDChilConductAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | uSiemens |  |  |  |  | HPO diode chiller, water conductivity
{ifo}:PSL-OSC_DCHILFLOW | AdL-REF-OPC1.rDChilFlowAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | l/min |  |  |  |  | HPO diode chiller, water flow
{ifo}:PSL-OSC_DCHILTEMP | AdL-REF-OPC1.rDChilTempAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode chiller, water temperature
{ifo}:PSL-OSC_DCHILTEMPSET | AdL-REF-OPC1.rDChilTempSet | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode chiller, water temperature, set value
{ifo}:PSL-OSC_LRAPOS | AdL-REF-OPC1.rOscActuatorPositionAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | um |  |  |  |  | HPO postion of long range actuator
{ifo}:PSL-OSC_LRAPZTVOLTMAX | AdL-REF-OPC1.rOscActuatorPztVoltageMax | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO trigger LRA movement (upper boundary)
{ifo}:PSL-OSC_LRAPZTVOLTMIN | AdL-REF-OPC1.rOscActuatorPztVoltageMin | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO trigger LRA movement (lower boundary)
{ifo}:PSL-OSC_LRARANGE | AdL-REF-OPC1.rOscActuatorRangeSet | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | um |  |  |  |  | HPO set maximum range for LRA with respect to reference point
{ifo}:PSL-OSC_LRAREFPOINT | AdL-REF-OPC1.rOscActuatorRefPointSet | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | um |  |  |  |  | HPO LRA refence point
{ifo}:PSL-OSC_LRASTEPSIZE | AdL-REF-OPC1.rOscActuatorStepSizeSet | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | um |  |  |  |  | HPO LRA step size
{ifo}:PSL-OSC_BASEPL_REARLTEMP | AdL-REF-OPC1.rOscBasePlateTempAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO Base Plate Temperature, sensor 1
{ifo}:PSL-OSC_BASEPL_CENTTEMP | AdL-REF-OPC1.rOscBasePlateTempAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO Base Plate Temperature, sensor 2
{ifo}:PSL-OSC_BASEPL_REARRTEMP | AdL-REF-OPC1.rOscBasePlateTempAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO Base Plate Temperature, sensor 3
{ifo}:PSL-OSC_BASEPL_FRONTLTEMP | AdL-REF-OPC1.rOscBasePlateTempAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO Base Plate Temperature, sensor 4
{ifo}:PSL-OSC_BASEPL_FRONTRTEMP | AdL-REF-OPC1.rOscBasePlateTempAct[5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO Base Plate Temperature, sensor 5
{ifo}:PSL-OSC_DB1_CNTWAVELENGTH | AdL-REF-OPC1.rOscDBCenterWavelengthAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO center wavelength of light from diode box 1
{ifo}:PSL-OSC_DB2_CNTWAVELENGTH | AdL-REF-OPC1.rOscDBCenterWavelengthAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO center wavelength of light from diode box 2
{ifo}:PSL-OSC_DB3_CNTWAVELENGTH | AdL-REF-OPC1.rOscDBCenterWavelengthAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO center wavelength of light from diode box 3
{ifo}:PSL-OSC_DB4_CNTWAVELENGTH | AdL-REF-OPC1.rOscDBCenterWavelengthAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO center wavelength of light from diode box 4
{ifo}:PSL-OSC_DB1_FWHM | AdL-REF-OPC1.rOscDBFWHMAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO full width half max. of light from diode box 1
{ifo}:PSL-OSC_DB2_FWHM | AdL-REF-OPC1.rOscDBFWHMAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO full width half max. of light from diode box 2
{ifo}:PSL-OSC_DB3_FWHM | AdL-REF-OPC1.rOscDBFWHMAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO full width half max. of light from diode box 3
{ifo}:PSL-OSC_DB4_FWHM | AdL-REF-OPC1.rOscDBFWHMAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO full width half max. of light from diode box 4
{ifo}:PSL-OSC_DB1_SMOMENT | Adl-REF-OPC1.rOscDBSMoment[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO first moment of spectral distribution of light from diode box 1
{ifo}:PSL-OSC_DB2_SMOMENT | Adl-REF-OPC1.rOscDBSMoment[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO first moment of spectral distribution of light from diode box 2
{ifo}:PSL-OSC_DB3_SMOMENT | Adl-REF-OPC1.rOscDBSMoment[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO first moment of spectral distribution of light from diode box 3
{ifo}:PSL-OSC_DB4_SMOMENT | Adl-REF-OPC1.rOscDBSMoment[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO first moment of spectral distribution of light from diode box 4
{ifo}:PSL-OSC_DB1_FMOMENT | Adl-REF-OPC1.rOscDBFMoment[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO second moment of spectral distribution of light from diode box 1
{ifo}:PSL-OSC_DB2_FMOMENT | Adl-REF-OPC1.rOscDBFMoment[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO second moment of spectral distribution of light from diode box 2
{ifo}:PSL-OSC_DB3_FMOMENT | Adl-REF-OPC1.rOscDBFMoment[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO second moment of spectral distribution of light from diode box 3
{ifo}:PSL-OSC_DB4_FMOMENT | Adl-REF-OPC1.rOscDBFMoment[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | nm |  |  |  |  | HPO second moment of spectral distribution of light from diode box 4
{ifo}:PSL-OSC_DB1_HTSNK1TEMP | AdL-REF-OPC1.rOscDBHeatsinkTempAct[1][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO heat sink temperature diode box #, sensor #
{ifo}:PSL-OSC_DB1_HTSNK2TEMP | AdL-REF-OPC1.rOscDBHeatsinkTempAct[1][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO heat sink temperature diode box #, sensor #
{ifo}:PSL-OSC_DB2_HTSNK1TEMP | AdL-REF-OPC1.rOscDBHeatsinkTempAct[2][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO heat sink temperature diode box #, sensor #
{ifo}:PSL-OSC_DB2_HTSNK2TEMP | AdL-REF-OPC1.rOscDBHeatsinkTempAct[2][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO heat sink temperature diode box #, sensor #
{ifo}:PSL-OSC_DB3_HTSNK1TEMP | AdL-REF-OPC1.rOscDBHeatsinkTempAct[3][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO heat sink temperature diode box #, sensor #
{ifo}:PSL-OSC_DB3_HTSNK2TEMP | AdL-REF-OPC1.rOscDBHeatsinkTempAct[3][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO heat sink temperature diode box #, sensor #
{ifo}:PSL-OSC_DB4_HTSNK1TEMP | AdL-REF-OPC1.rOscDBHeatsinkTempAct[4][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO heat sink temperature diode box #, sensor #
{ifo}:PSL-OSC_DB4_HTSNK2TEMP | AdL-REF-OPC1.rOscDBHeatsinkTempAct[4][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO heat sink temperature diode box #, sensor #
{ifo}:PSL-OSC_DB1_PWR | AdL-REF-OPC1.rOscDBPowerAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | HPO pump light monitor at laser head 1
{ifo}:PSL-OSC_DB2_PWR | AdL-REF-OPC1.rOscDBPowerAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | HPO pump light monitor at laser head 2
{ifo}:PSL-OSC_DB3_PWR | AdL-REF-OPC1.rOscDBPowerAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | HPO pump light monitor at laser head 3
{ifo}:PSL-OSC_DB4_PWR | AdL-REF-OPC1.rOscDBPowerAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | HPO pump light monitor at laser head 4
{ifo}:PSL-OSC_DB1_SPECPWR | AdL-REF-OPC1.rOscDBSpectralPowerAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | counts
{ifo}:PSL-OSC_DB2_SPECPWR | AdL-REF-OPC1.rOscDBSpectralPowerAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | counts
{ifo}:PSL-OSC_DB3_SPECPWR | AdL-REF-OPC1.rOscDBSpectralPowerAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | counts
{ifo}:PSL-OSC_DB4_SPECPWR | AdL-REF-OPC1.rOscDBSpectralPowerAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | counts
{ifo}:PSL-OSC_DB1_WPD | AdL-REF-OPC1.rOscDBWPDAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | HPO pump light monitor at laser head 1, percentage of initial power
{ifo}:PSL-OSC_DB2_WPD | AdL-REF-OPC1.rOscDBWPDAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | HPO pump light monitor at laser head 2, percentage of initial power
{ifo}:PSL-OSC_DB3_WPD | AdL-REF-OPC1.rOscDBWPDAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | HPO pump light monitor at laser head 3, percentage of initial power
{ifo}:PSL-OSC_DB4_WPD | AdL-REF-OPC1.rOscDBWPDAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | HPO pump light monitor at laser head 4, percentage of initial power
{ifo}:PSL-OSC_DB1_CUR | AdL-REF-OPC1.rOscDiodeCurrentAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO Diode current, diode box 1
{ifo}:PSL-OSC_DB2_CUR | AdL-REF-OPC1.rOscDiodeCurrentAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO Diode current, diode box 2
{ifo}:PSL-OSC_DB3_CUR | AdL-REF-OPC1.rOscDiodeCurrentAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO Diode current, diode box 3
{ifo}:PSL-OSC_DB4_CUR | AdL-REF-OPC1.rOscDiodeCurrentAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO Diode current, diode box 4
{ifo}:PSL-OSC_DB1_CURSET | AdL-REF-OPC1.rOscDiodeCurrentSet[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO Diode current set point, diode box 1
{ifo}:PSL-OSC_DB2_CURSET | AdL-REF-OPC1.rOscDiodeCurrentSet[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO Diode current set point, diode box 2
{ifo}:PSL-OSC_DB3_CURSET | AdL-REF-OPC1.rOscDiodeCurrentSet[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO Diode current set point, diode box 3
{ifo}:PSL-OSC_DB4_CURSET | AdL-REF-OPC1.rOscDiodeCurrentSet[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO Diode current set point, diode box 4
{ifo}:PSL-OSC_DB1_D1PWR | AdL-REF-OPC1.rOscDiodePowerAct[1][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB1_D2PWR | AdL-REF-OPC1.rOscDiodePowerAct[1][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB1_D3PWR | AdL-REF-OPC1.rOscDiodePowerAct[1][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB1_D4PWR | AdL-REF-OPC1.rOscDiodePowerAct[1][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB1_D5PWR | AdL-REF-OPC1.rOscDiodePowerAct[1][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB1_D6PWR | AdL-REF-OPC1.rOscDiodePowerAct[1][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB1_D7PWR | AdL-REF-OPC1.rOscDiodePowerAct[1][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB2_D1PWR | AdL-REF-OPC1.rOscDiodePowerAct[2][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB2_D2PWR | AdL-REF-OPC1.rOscDiodePowerAct[2][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB2_D3PWR | AdL-REF-OPC1.rOscDiodePowerAct[2][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB2_D4PWR | AdL-REF-OPC1.rOscDiodePowerAct[2][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB2_D5PWR | AdL-REF-OPC1.rOscDiodePowerAct[2][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB2_D6PWR | AdL-REF-OPC1.rOscDiodePowerAct[2][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB2_D7PWR | AdL-REF-OPC1.rOscDiodePowerAct[2][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB3_D1PWR | AdL-REF-OPC1.rOscDiodePowerAct[3][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB3_D2PWR | AdL-REF-OPC1.rOscDiodePowerAct[3][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB3_D3PWR | AdL-REF-OPC1.rOscDiodePowerAct[3][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB3_D4PWR | AdL-REF-OPC1.rOscDiodePowerAct[3][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB3_D5PWR | AdL-REF-OPC1.rOscDiodePowerAct[3][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB3_D6PWR | AdL-REF-OPC1.rOscDiodePowerAct[3][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB3_D7PWR | AdL-REF-OPC1.rOscDiodePowerAct[3][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB4_D1PWR | AdL-REF-OPC1.rOscDiodePowerAct[4][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB4_D2PWR | AdL-REF-OPC1.rOscDiodePowerAct[4][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB4_D3PWR | AdL-REF-OPC1.rOscDiodePowerAct[4][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB4_D4PWR | AdL-REF-OPC1.rOscDiodePowerAct[4][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB4_D5PWR | AdL-REF-OPC1.rOscDiodePowerAct[4][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB4_D6PWR | AdL-REF-OPC1.rOscDiodePowerAct[4][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB4_D7PWR | AdL-REF-OPC1.rOscDiodePowerAct[4][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO diode power internal monitor, diode box #, diode #
{ifo}:PSL-OSC_DB1_D1TEMP | AdL-REF-OPC1.rOscDiodeTempAct[1][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB1_D2TEMP | AdL-REF-OPC1.rOscDiodeTempAct[1][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB1_D3TEMP | AdL-REF-OPC1.rOscDiodeTempAct[1][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB1_D4TEMP | AdL-REF-OPC1.rOscDiodeTempAct[1][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB1_D5TEMP | AdL-REF-OPC1.rOscDiodeTempAct[1][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB1_D6TEMP | AdL-REF-OPC1.rOscDiodeTempAct[1][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB1_D7TEMP | AdL-REF-OPC1.rOscDiodeTempAct[1][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB2_D1TEMP | AdL-REF-OPC1.rOscDiodeTempAct[2][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB2_D2TEMP | AdL-REF-OPC1.rOscDiodeTempAct[2][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB2_D3TEMP | AdL-REF-OPC1.rOscDiodeTempAct[2][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB2_D4TEMP | AdL-REF-OPC1.rOscDiodeTempAct[2][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB2_D5TEMP | AdL-REF-OPC1.rOscDiodeTempAct[2][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB2_D6TEMP | AdL-REF-OPC1.rOscDiodeTempAct[2][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB2_D7TEMP | AdL-REF-OPC1.rOscDiodeTempAct[2][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB3_D1TEMP | AdL-REF-OPC1.rOscDiodeTempAct[3][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB3_D2TEMP | AdL-REF-OPC1.rOscDiodeTempAct[3][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB3_D3TEMP | AdL-REF-OPC1.rOscDiodeTempAct[3][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB3_D4TEMP | AdL-REF-OPC1.rOscDiodeTempAct[3][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB3_D5TEMP | AdL-REF-OPC1.rOscDiodeTempAct[3][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB3_D6TEMP | AdL-REF-OPC1.rOscDiodeTempAct[3][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB3_D7TEMP | AdL-REF-OPC1.rOscDiodeTempAct[3][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB4_D1TEMP | AdL-REF-OPC1.rOscDiodeTempAct[4][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB4_D2TEMP | AdL-REF-OPC1.rOscDiodeTempAct[4][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB4_D3TEMP | AdL-REF-OPC1.rOscDiodeTempAct[4][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB4_D4TEMP | AdL-REF-OPC1.rOscDiodeTempAct[4][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB4_D5TEMP | AdL-REF-OPC1.rOscDiodeTempAct[4][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB4_D6TEMP | AdL-REF-OPC1.rOscDiodeTempAct[4][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB4_D7TEMP | AdL-REF-OPC1.rOscDiodeTempAct[4][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature, diode box #, diode #
{ifo}:PSL-OSC_DB1_D1TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[1][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB1_D2TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[1][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB1_D3TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[1][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB1_D4TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[1][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB1_D5TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[1][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB1_D6TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[1][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB1_D7TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[1][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB2_D1TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[2][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB2_D2TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[2][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB2_D3TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[2][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB2_D4TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[2][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB2_D5TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[2][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB2_D6TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[2][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB2_D7TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[2][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB3_D1TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[3][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB3_D2TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[3][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB3_D3TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[3][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB3_D4TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[3][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB3_D5TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[3][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB3_D6TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[3][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB3_D7TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[3][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB4_D1TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[4][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB4_D2TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[4][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB4_D3TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[4][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB4_D4TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[4][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB4_D5TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[4][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB4_D6TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[4][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB4_D7TEC | AdL-REF-OPC1.rOscDiodeTempCtrl[4][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | A |  |  |  |  | HPO diode temperature control signal, diode box #, diode #
{ifo}:PSL-OSC_DB1_D1TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[1][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB1_D2TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[1][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB1_D3TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[1][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB1_D4TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[1][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB1_D5TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[1][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB1_D6TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[1][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB1_D7TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[1][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB2_D1TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[2][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB2_D2TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[2][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB2_D3TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[2][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB2_D4TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[2][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB2_D5TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[2][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB2_D6TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[2][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB2_D7TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[2][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB3_D1TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[3][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB3_D2TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[3][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB3_D3TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[3][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB3_D4TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[3][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB3_D5TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[3][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB3_D6TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[3][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB3_D7TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[3][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB4_D1TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[4][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB4_D2TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[4][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB4_D3TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[4][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB4_D4TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[4][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB4_D5TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[4][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB4_D6TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[4][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB4_D7TEMPSET | AdL-REF-OPC1.rOscDiodeTempSet[4][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO diode temperature set point, diode box #, diode #
{ifo}:PSL-OSC_DB1_VOLT | AdL-REF-OPC1.rOscDiodeVoltageAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO supply voltage diode box #
{ifo}:PSL-OSC_DB2_VOLT | AdL-REF-OPC1.rOscDiodeVoltageAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO supply voltage diode box #
{ifo}:PSL-OSC_DB3_VOLT | AdL-REF-OPC1.rOscDiodeVoltageAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO supply voltage diode box #
{ifo}:PSL-OSC_DB4_VOLT | AdL-REF-OPC1.rOscDiodeVoltageAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO supply voltage diode box #
{ifo}:PSL-OSC_DB1_D1WPD | AdL-REF-OPC1.rOscDiodeWPDAct[1][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB1_D2WPD | AdL-REF-OPC1.rOscDiodeWPDAct[1][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB1_D3WPD | AdL-REF-OPC1.rOscDiodeWPDAct[1][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB1_D4WPD | AdL-REF-OPC1.rOscDiodeWPDAct[1][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB1_D5WPD | AdL-REF-OPC1.rOscDiodeWPDAct[1][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB1_D6WPD | AdL-REF-OPC1.rOscDiodeWPDAct[1][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB1_D7WPD | AdL-REF-OPC1.rOscDiodeWPDAct[1][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB2_D1WPD | AdL-REF-OPC1.rOscDiodeWPDAct[2][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB2_D2WPD | AdL-REF-OPC1.rOscDiodeWPDAct[2][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB2_D3WPD | AdL-REF-OPC1.rOscDiodeWPDAct[2][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB2_D4WPD | AdL-REF-OPC1.rOscDiodeWPDAct[2][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB2_D5WPD | AdL-REF-OPC1.rOscDiodeWPDAct[2][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB2_D6WPD | AdL-REF-OPC1.rOscDiodeWPDAct[2][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB2_D7WPD | AdL-REF-OPC1.rOscDiodeWPDAct[2][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB3_D1WPD | AdL-REF-OPC1.rOscDiodeWPDAct[3][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB3_D2WPD | AdL-REF-OPC1.rOscDiodeWPDAct[3][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB3_D3WPD | AdL-REF-OPC1.rOscDiodeWPDAct[3][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB3_D4WPD | AdL-REF-OPC1.rOscDiodeWPDAct[3][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB3_D5WPD | AdL-REF-OPC1.rOscDiodeWPDAct[3][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB3_D6WPD | AdL-REF-OPC1.rOscDiodeWPDAct[3][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB3_D7WPD | AdL-REF-OPC1.rOscDiodeWPDAct[3][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB4_D1WPD | AdL-REF-OPC1.rOscDiodeWPDAct[4][1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB4_D2WPD | AdL-REF-OPC1.rOscDiodeWPDAct[4][2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB4_D3WPD | AdL-REF-OPC1.rOscDiodeWPDAct[4][3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB4_D4WPD | AdL-REF-OPC1.rOscDiodeWPDAct[4][4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB4_D5WPD | AdL-REF-OPC1.rOscDiodeWPDAct[4][5] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB4_D6WPD | AdL-REF-OPC1.rOscDiodeWPDAct[4][6] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_DB4_D7WPD | AdL-REF-OPC1.rOscDiodeWPDAct[4][7] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | was ment to give the relative diode power of diode # in DB # (percentage), not used
{ifo}:PSL-OSC_BOXHUM | AdL-REF-OPC1.rOscHumidityAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | % |  |  |  |  | HPO humidity in HPO laser box
{ifo}:PSL-OSC_FLOWDOGWINDOW | AdL-REF-OPC1.rOscFlowWatchdogWindow | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | l/min |  |  |  |  | OSC flow window for the watchdog (0,2l/min)
{ifo}:PSL-OSC_PRESS1MAX | AdL-REF-OPC1.rOscPressure1MaxSet | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | bar |  |  |  |  | OSC Water pressure max ()
{ifo}:PSL-OSC_HEAD1TEMP | AdL-REF-OPC1.rOscLaserHeadTempAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO Laser Head Temperature, head #
{ifo}:PSL-OSC_HEAD2TEMP | AdL-REF-OPC1.rOscLaserHeadTempAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO Laser Head Temperature, head #
{ifo}:PSL-OSC_HEAD3TEMP | AdL-REF-OPC1.rOscLaserHeadTempAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO Laser Head Temperature, head #
{ifo}:PSL-OSC_HEAD4TEMP | AdL-REF-OPC1.rOscLaserHeadTempAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO Laser Head Temperature, head #
{ifo}:PSL-OSC_HEAD1FLOW | AdL-REF-OPC1.rOscLaserHeadWaterFlowAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | l/min |  |  |  |  | HPO Laser Head water flow rate [l/min], head #
{ifo}:PSL-OSC_HEAD2FLOW | AdL-REF-OPC1.rOscLaserHeadWaterFlowAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | l/min |  |  |  |  | HPO Laser Head water flow rate [l/min], head #
{ifo}:PSL-OSC_HEAD3FLOW | AdL-REF-OPC1.rOscLaserHeadWaterFlowAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | l/min |  |  |  |  | HPO Laser Head water flow rate [l/min], head #
{ifo}:PSL-OSC_HEAD4FLOW | AdL-REF-OPC1.rOscLaserHeadWaterFlowAct[4] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | l/min |  |  |  |  | HPO Laser Head water flow rate [l/min], head #
{ifo}:PSL-OSC_PWR1 | AdL-REF-OPC1.rOscLaserPowerAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | HPO Laser Power, backward direction
{ifo}:PSL-OSC_PWR2 | AdL-REF-OPC1.rOscLaserPowerAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | HPO Laser Power, forward direction (only valid if external shutter is closed)
{ifo}:PSL-OSC_PWREXT | AdL-REF-OPC1.rOscLaserPowerAct[3] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | HPO reading of power meter externally connected
{ifo}:PSL-OSC_PWR_SUM | AdL-REF-OPC1.rOscLaserPowerActSum | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | HPO sum of all three power meter
{ifo}:PSL-OSC_PDLOCKDC_PWR | AdL-REF-OPC1.rOscLaserPowerFromPdLockingDcAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | W |  |  |  |  | HPO reading of injection locking photo diode DC, calibrated in Watt
{ifo}:PSL-OSC_LOCKGAIN | AdL-REF-OPC1.rOscLockGainSet | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO loop gain of injection locking servo, set point
{ifo}:PSL-OSC_LOCKOFFSET | AdL-REF-OPC1.rOscLockOffsetSet | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO error point of injection locking servo, set point
{ifo}:PSL-OSC_LOCKREF | AdL-REF-OPC1.rOscLockRefSet | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO injection locking auto lock reference level, set point
{ifo}:PSL-OSC_PDLOCKDC | AdL-REF-OPC1.rOscPdLockingDcAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO reading of injection locking photo diode DC [Volt]
{ifo}:PSL-OSC_PDLOCKDC_MEAN | AdL-REF-OPC1.rOscPdLockingDcActMean | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO reading of injection locking photo diode DC, 2s average
{ifo}:PSL-OSC_PWRMETERFLOW | AdL-REF-OPC1.rOscPowermeterWaterFlowAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | l/min |  |  |  |  | HPO water flow through laser internal aux. beam dumps and power meter
{ifo}:PSL-OSC_PZTVOLT | AdL-REF-OPC1.rOscPztVoltageAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | V |  |  |  |  | HPO injection locking control signal
{ifo}:PSL-OSC_BOXTEMP | AdL-REF-OPC1.rOscTempAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO air temperature in HPO laser box
{ifo}:PSL-OSC_PRESS1 | AdL-REF-OPC1.rOscWaterPressureAct[1] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | bar |  |  |  |  | HPO water pressure inlet at water manifold
{ifo}:PSL-OSC_PRESS2 | AdL-REF-OPC1.rOscWaterPressureAct[2] | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | bar |  |  |  |  | HPO water pressure return at water manifold
{ifo}:PSL-OSC_XCHILCND | AdL-REF-OPC1.rXChilConductAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | uSiemens |  |  |  |  | HPO crystal chiller, water conductivity
{ifo}:PSL-OSC_XCHILFLOW | AdL-REF-OPC1.rXChilFlowAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | l/min |  |  |  |  | HPO crystal chiller, water flow
{ifo}:PSL-OSC_XCHILTEMP | AdL-REF-OPC1.rXChilTempAct | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO crystal chiller, water temperature
{ifo}:PSL-OSC_XCHILTEMPSET | AdL-REF-OPC1.rXChilTempSet | REAL | AdL-REF_Read.db | 125 | 1 |  |  |  | °C |  |  |  |  | HPO crystal chiller, water temperature, set value
{ifo}:PSL-OSC_USER | AdL-REF-OPC1.sUSER | String | AdL-REF_Read.db | ? | ? | [8 empty] | name of laser operator who put laser into service mode

## Pumplight (L)
L1 / laser diode fiber
one front end fiber bundle (4 fibers plus 2 spare fibers)
six high power laser fiber bundels (with 7 fibres each), 4 bundles required for laser operation, two spare bundles
According to the manufacturer the fibers have a minimum bending radius for long term of 132 mm and 22mm for short term.

## Water (W)
W1, W2 / cooling water for laser diodes and laser heads
W1 termotek chiller - laser room water distribution
3/4 '' tube
25 liter / min
T= 18°C
p_in= 4.5 bar
W2 termotek chiller - laser diode rack water distribution
silicon tube
LD: transfer heat via water running through cold plates to water-water chiller (two point control) (connects to LIGO facility water)
Xtal: front end no TEC only cooling water of secondary loop, transfers heat to water-water chiller (analog valve control) (connects to LIGO facility water)
temperature range Xtal loop
temperature stability Xtal loop | 0.1C TBC
temperature range LD loop
temperature stability LD loop | 0.5C TBC
heat load - Xtal | 10kW
heat load - LD | 10kW
AdLIGO Laser Calculations
total electrical heat load (W) | 2194.4898
total thermal load (W) | 4370.5526
Pump Diode Calculations |  |  |  | Personnel Calculations
maximum optical output power Pmax (W) | 45 | maximum input power Pinput (W) | 130 | number of working personnel | 4
maximum operation current Amax (A) | 65 | maximum heat generated Hdiode (W) | 85 | moderate work heat load (W) | 160
maximum operating voltage Vmax (V) | 2
|  |  |  | total personnel heat load Qpeople (W) | 640
maximum TEC cooling power (W) | 346 | maximum TEC input power (W) | 555.96
maximum TEC current (A) | 11.3 | coefficient of performance COP | 0.62235
maximum TEC voltage (V) | 49.2 | maximum TEC heat (W) | 209.96
TEC electrical power input (W) | 136.5798
TEC generated heat (W) | 51.5798
heat generated per diode, Qdiode (W) | 136.5798
MOPA Calculations |  |  |  | High Power Stage Calculations
number of diodes per amplifier | 4 |  |  | number of heads | 4
amplifier heat load Qamp (W) | 546.3191 |  |  | number of diodes per head | 7
|  |  |  | high power stage heat load QHPS (W) | 3824.2335
Power Supply Calculations |  |  |  | Power Supply Calculations
maximum output voltage (V) |  |  |  | maximum output voltage (V) | 110
maximum output current (A) |  |  |  | maximum output current (A) | 16
maximum output power (W) | 600 |  |  | maximum output power (W) | 1505
power factor | 0.98 | maximum input power (W) | 612.2449 | power supply efficiency | 0.85511
heat load per power supply (W) | 12.2449 |  |  | power supply heat QPS (W) | 255
number of supplies per MOPA | 2 |  |  | number of diode boxes | 4
power supply heat load (W) | 24.4898 |  |  | number of power supplies per diode box | 1
|  |  |  | number of TEC power supplies | 2
|  |  |  | total heat from power supplies Qsupplies (W) | 1530
Note on A40: This is the load that has to be removed by the air handling system.
Note on A41: This is the load that has to be removed by the chillers.
Note on A43: Jenoptik fibre-coupled passively cooled cw diode laser JOLD-45-CPXF-1L Design 15415624
Note on H46: This number was obtained from http://www.engineeringtoolbox.com. Listed is the average metabolic rate for an adult male conducting moderate work.
Note on A53: Assuming the coefficient of performance stays the same. Then to remove the 85 W of heat produced by the pump diode, the electrical power input should be 85 W divided by the coefficient of performance. The difference between the electrical power input and the 85 W is assumed to be the heat generated by the TEC.
Note on H65: The power supplies for the high power stage are assumed to be from Deltra Elektronika SM52-AR-60. These are assumed to be the power supplies for the TECs too, since we seem to have a problem with the 3 phase power supplies.

## external (E)
E1, E2 optical interfaces to IO
E1: PSL IO interface on PSl table (beam waist in PMC)
xy position in PSL coordinates (Visio file)
beam height
beam waist raduis
power
polarization
E2: PSL IO interface in HAM2 for outer loop ISS sensing
xy position in PSL coordinates (Visio file)
beam height
beam waist raduis
power
polarization
E3: PSL / ISC interface
wide-band FSS actuator
tidal feed-back actuator
laser power control
IO HAM2 tip/tilt steering mirror for ISS outer loop diode 1
IO HAM2 tip/tilt steering mirror for ISS outer loop diode 2
E4: not assigned
E5 / laser diode enclosure - facility
cooling
into water (cold plates/ Xtals) | 4.4kW
into air (power supplies / personnel) | 2.1 kW
electricity
eLIGO | Vanfier diode box | 2 power supplies | single phase | 600W each
AdvLIGO | 4 diode boxes | 1 power supplie each | single phase | 1505W each
monitoring and services
telephone
ethernet (CDS network)
dust monitor via CDS
temperature monitor via CDS
webcam monitor via CDS
optical fiber to laser enclosure
communication fiber to laser enclosure | Beckhofff network / labtop
chilled water tubes to laser enclosure
E6 LVEA - facility
E6 Laser diode room - facility
Note on B32: bewill: from Peter Document thermal_load_LDR.xls
Note on B33: bewill: from Peter Document thermal_load_LDR.xls

## channel description
aLIGO PSL channel list |  | 2013-06-10
premode cleaner - via CDS
{ifo}:PSL-PMC_TF_IN_ON | enables summation of TEST2 input | DAC_2_1
{ifo}:PSL-PMC_LOCK_ON | turns lock automation on | DAC_2_2
{ifo}:PSL-PMC_RAMP_ON | turns ramp on | DAC_2_3
{ifo}:PSL-PMC_GAIN | set gain of control loop | DAC_2_8 | dB
{ifo}:PSL-PMC_REF | sets threshold which lock acquisition automation assumes PMC is locked | DAC_2_9 | V
{ifo}:PSL-PMC_INOFFSET | correction of electronic offset in control loop | DAC_2_10 | V
{ifo}:PSL-PMC_HEATER | set offset of heater and hence operation point of PMC temp. loop | DAC_2_11
{ifo}:PSL-PMC_BLANKING | opens feedback loop | DAC_2_4
{ifo}:PSL-PMC_BOOST | activates integrator in control loop | DAC_2_5
{ifo}:PSL-PMC_TF_IN | summation point for calibration lines, used for alignment ramp | DAC_2_6
{ifo}:PSL-PMC_RAMP | not used (would be ramp input in case digital lock acquisition is used) | DAC_2_7 | V
{ifo}:PSL-PMC_HV_MON | monitor of HV signal send to PZT actuator of PMC | ADC_2_4 | V
{ifo}:PSL-PMC_MIXER | error signal of control loop | ADC_2_5 | V
{ifo}:PSL-PMC_LO_POWER_MON | power of local oscillator used in demodulation electronic | ADC_2_6 | dBm
{ifo}:PSL-PMC_TRIGGER | DC signal of locking diode, used by autolock to dertermine if PMC is resonant | ADC_2_7 | V
{ifo}:PSL-PMC_TF_B | test point for transfer function measurement | ADC_2_0 | V
{ifo}:PSL-PMC_TEMP | temperature sensor on PMC | ADC_2_1 | K
{ifo}:PSL-PMC_TF_A | test point for transferfunction measurement | ADC_2_2 | V
{ifo}:PSL-PMC_RESONANT_MON | binary signals, 1 if auto lock thinks PMC is resonant | ADC_2_3
power stabilization - via CDS
{ifo}:PSL-ISS_AOM | ISS actuator signal send to AOM | ADC_0_8 | V
{ifo}:PSL-ISS_QPDDX | quadrant photodiode dx signal | ADC_0_9 | mm
{ifo}:PSL-ISS_QPDDY | quadrant photodiode dy signal | ADC_0_10 | mm
{ifo}:PSL-ISS_REFSIGNAL | reference to which the ISS stabilizes the power | DAC_0_0 | V
{ifo}:PSL-ISS_TRANSFER1_INJ | summation into error point | DAC_0_1 | V
{ifo}:PSL-ISS_TRANSFER2_INJ | summation into outer-loop signal | DAC_0_2 | V
{ifo}:PSL-ISS_GAIN | ISS control loop gain | DAC_0_3 | dB
{ifo}:PSL-ISS_CTRL_OFFSET | offest added to control signal to enable AOM to increase power | DAC_0_4 | V
{ifo}:PSL-ISS_INLOOP_PD_SELECT | selects if photodiode A of B is used as in-loop sensor | DAC_0_5
{ifo}:PSL-ISS_SECONDLOOP_CLOSED | switch to open/close second ISS feedback loop | DAC_0_6
{ifo}:PSL-ISS_LOOP_STATE | switch to open/close feedback loop | DAC_0_7
{ifo}:PSL-ISS_PDAMON | monitor signal of photodiode A, (at output of signal conditioning filter) | ADC_0_0 | V
{ifo}:PSL-ISS_PDBMON | monitor signal of photodiode B, (at output of signal conditioning filter) | ADC_0_1 | V
{ifo}:PSL-ISS_OUTERPDMON | monitor signal of second loop photodiode | ADC_0_2 | V
{ifo}:PSL-ISS_REFSIGNALOUT | low pass filtered readback of ref. signal | ADC_0_3 | V
{ifo}:PSL-ISS_TRANSFER1A | signal infront of summation point | ADC_0_4 | V
{ifo}:PSL-ISS_TRANSFER1B | signal behind summation point | ADC_0_5 | V
{ifo}:PSL-ISS_TRANSFER2A | signal infront of summation point in second loop | ADC_0_6 | V
{ifo}:PSL-ISS_TRANSFER2B | signal behind summation point in second loop | ADC_0_7 | V
frequency stabilization - via CDS
{ifo}:PSL-FSS_MIXER | FSS loop errorpoint | ADC_1_0 | V
{ifo}:PSL-FSS_IN1 | LEMO input IN1 at TTFSS fieldbox | ADC_1_1
{ifo}:PSL-FSS_FAST_MON | monitor of NPRO PZT signal | ADC_1_2 | V
{ifo}:PSL-FSS_PC_MON | monitor of signal to phase correcting pockels cell | ADC_1_3 | V
{ifo}:PSL-FSS_LO_POWER_MON | strength of local oscilator signal | ADC_1_4 | dBm
{ifo}:PSL-FSS_TEST2_MON | LEMO input TEST2 at TTFSS fieldbox | ADC_1_5 | V
{ifo}:PSL-FSS_RFPD_DC | DC signal of photodiode in reflection of reference cavity | ADC_1_6 | V
{ifo}:PSL-FSS_TPD_DC | DC signal of photodiode in transmission of reference cavity | ADC_1_7 | V
{ifo}:PSL-FSS_COMMON_GAIN | common gain | DAC_1_0 | dB
{ifo}:PSL-FSS_MIXER_OFS | signal to be added into error point | DAC_1_1 | V
{ifo}:PSL-FSS_FAST_GAIN | fast path gain (NPRO PZT) | DAC_1_2 | dB
{ifo}:PSL-FSS_FAST_RAMP_INJECTION | signal given to the SLOW BNC output of TTFSS (connected to RAMP IN) | DAC_1_3 | V
{ifo}:PSL-FSS_TEST1_ON | enables BNC TEST1 IN (summation at output of MIXER (at input of LP filter) | DAC_1_4
{ifo}:PSL-FSS_TEST2_ON | enables BNC TEST2 IN (summation at error point, at output of MIXER LP filter) | DAC_1_5
{ifo}:PSL-FSS_LOOP_CLOSED | open/close control loop after error point and infront of ramp injection in fast path | DAC_1_6
{ifo}:PSL-FSS_NPRO_TEMP | signal to LEMO TEMP at TTFSS FB | DAC_1_7 | °C (offset)
{ifo}:PSL-FSS_VCO_POWER_MON | VCO power monitor | ADC_1_8 | dBm
{ifo}:PSL-FSS_VCO_MODLEVEL | sets the level of the rf-signal send from the VCO to the FSS-AOM | DAC_1_12 | ????
{ifo}:PSL-FSS_VCO_TEST_ON | enables the TEST input (BNC at VCO box) | DAC_1_13
{ifo}:PSL-FSS_WIDE_ON | enables the WIDE input (BNC at VCO box, wideband actuator signal from MC loop) | DAC_1_14
diagnostic breadboard - via CDS
{ifo}:PSL-DBB_QPD1_QX | quadrant phtotodiode 1 dx -DC | ADC_3_0 | V
{ifo}:PSL-DBB_QPD1_QY | quadrant phtotodiode 1 dy -DC | ADC_3_1 | V
{ifo}:PSL-DBB_QPD1_QS | quadrant phtotodiode 1 sum -DC | ADC_3_2 | V
{ifo}:PSL-DBB_QPD2_QX | quadrant phtotodiode 1 dx -DC | ADC_0_3 | V
{ifo}:PSL-DBB_QPD2_QY | quadrant phtotodiode 1 dy -DC | ADC_3_4 | V
{ifo}:PSL-DBB_QPD2_QS | quadrant phtotodiode 1 sum -DC | ADC_3_5 | V
{ifo}:PSL-DBB_RPD_DC | power noise sensing photodiode DC | ADC_3_6 | V
{ifo}:PSL-DBB_TPD_0DB | photodiode in transmission of cavity (not amplified) | ADC_3_7 | V
{ifo}:PSL-DBB_QPD1_DX | quadrant phtotodiode 1 dx - demodulated | ADC_3_8 | V
{ifo}:PSL-DBB_QPD1_DY | quadrant phtotodiode 1 dy - demodulated | ADC_3_9 | V
{ifo}:PSL-DBB_QPD_DS | quadrant phtotodiode 1 sum - demodulated | ADC_3_10 | V
{ifo}:PSL-DBB_QPD2_DX | quadrant phtotodiode 1 dx - demodulated | ADC_0_11 | V
{ifo}:PSL-DBB_QPD2_DY | quadrant phtotodiode 1 dy - demodulated | ADC_3_12 | V
{ifo}:PSL-DBB_PZTMON | actuator signal of locking loop with signal conditioning | ADC_3_13 | V
{ifo}:PSL-DBB_HVMON | actuator signal of locking loop | ADC_3_14 | V
{ifo}:PSL-DBB_RPD_AC | power noise sensing photodiode AC | ADC_3_15 | V
{ifo}:PSL-DBB_TPD_40DB | photodiode in transmission of cavity (40dB gain output) | ADC_3_16 | V
{ifo}:PSL-DBB_TPD_80DB | photodiode in transmission of cavity (80db gain output) | ADC_3_17 | V
{ifo}:PSL-DBB_SHUTTER_REQ | close shutter request | ADC_3_18
{ifo}:PSL-DBB_AUX_INPUT1 | auxilary input 1 | ADC_0_19 | V
{ifo}:PSL-DBB_AUX_INPUT2 | auxilary input 2 | ADC_3_20 | V
{ifo}:PSL-DBB_AUX_INPUT3 | auxilary input 3 | ADC_3_21 | V
{ifo}:PSL-DBB_AUX_INPUT4 | auxilary input 4 | ADC_3_22 | V
{ifo}:PSL-DBB_RMT_READBACK | indicates if remote operation is enabled | ADC_3_23
{ifo}:PSL-DBB_CTRL0 | control signal length control | DAC_3_0 | V
{ifo}:PSL-DBB_CTRL1X | control signal alignment control | DAC_3_1 | V
{ifo}:PSL-DBB_CTRL1Y | control signal alignment control | DAC_3_2 | V
{ifo}:PSL-DBB_CTRL2X | control signal alignment control | DAC_0_3 | V
{ifo}:PSL-DBB_CTRL2Y | control signal alignment control | DAC_3_4 | V
{ifo}:PSL-DBB_PHASE | demodulation phase | DAC_3_5 | deg
{ifo}:PSL-DBB_LENS1 | actuates mode matching lens 1 | DAC_3_6 | mm
{ifo}:PSL-DBB_LENS2 | actuates mode matching lens 2 | DAC_3_7 | mm
{ifo}:PSL-DBB_LENS-LATCH | holds lenses in current position | DAC_3_8
{ifo}:PSL-DBB_SHUTTER_DBB | shutter state | DAC_3_9
{ifo}:PSL-DBB_RMT | remote operation is requested | DAC_3_10
{ifo}:PSL-DBB_MODOFF | indicates if modulation is enabled | DAC_0_11
{ifo}:PSL-DBB_SCANMODE | switch between scan- and lockmode | DAC_3_12
{ifo}:PSL-DBB_SHUTTER_MULTIPLEX | select shutter in FE or HPO path | DAC_3_13
{ifo}:PSL-DBB_FILTER_SWITCH1 |  | DAC_3_14
{ifo}:PSL-DBB_FILTER_SWITCH2 |  | DAC_3_15
injection locking/power/miscellaneous - via CDS
{ifo}:PSL-ILS_HV_MON | monitor of HV signal send to PZT actuator of ILS | ADC_2_12 | V
{ifo}:PSL-ILS_MIXER | error signal of injection locking control loop | ADC_2_13 | V
{ifo}:PSL-ILS_LO_POWER_MON | power of local oscillator used in demodulation electronic | ADC_2_14 | dBm
{ifo}:PSL-ILS_TRIGGER | DC signal of locking diode, used by autolock to dertermine if HPL is resonant | ADC_2_15 | V
{ifo}:PSL-ILS_TF_B | test point for transfer function measurement | ADC_2_8 | V
{ifo}:PSL-ILS_TF_A | test point for transferfunction measurement | ADC_2_10 | V
{ifo}:PSL-ILS_RESONANT_ON | binary signals, 1 if auto lock thinks PMC is resonant | ADC_2_11
{ifo}:PSL-OSC_PDAMP_DC | monitoring photodiode in HPL: signal from front-end at output of Faraday Isolator | ADC_2_16 | V
{ifo}:PSL-OSC_PDISO_DC | monitoring photodiode in HPL: power in reverse direction split of at FI | ADC_2_17 | V
{ifo}:PSL-OSC_PDBP_DC | monitoring photodiode in HPL: reflection at Brewster Plate in HPL oscillator | ADC_2_18 | V
{ifo}:PSL-OSC_PDINT_DC | monitoring photodiode in HPL: circulating light power in HPL oscillator | ADC_2_19 | V
{ifo}:PSL-OSC_PDAMP_AC | monitoring photodiode in HPL: signal from front-end at output of Faraday Isolator | ADC_2_20 | V
{ifo}:PSL-OSC_PDISO_AC | monitoring photodiode in HPL: power in reverse direction split of at FI | ADC_2_21 | V
{ifo}:PSL-OSC_PDBP_AC | monitoring photodiode in HPL: reflection at Brewster Plate in HPL oscillator | ADC_2_22 | V
{ifo}:PSL-OSC_PDINT_AC | monitoring photodiode in HPL: circulating light power in HPL oscillator | ADC_2_23 | V
{ifo}:PSL-PWR_PMC_REFL | power in reflection of PMC (power meter used for installation and maintenance) | ADC_2_24 | W
{ifo}:PSL-PWR_PMC_TRANS | power in transmission of PMC (power meter used for installation and maintenance) | ADC_2_25 | W
{ifo}:PSL-PWR_HPL_DC | monitor photodiode (PD001) at output of HPL, main HPL power monitor | ADC_2_26 | W
{ifo}:PSL-PWR_HPL_AC | monitor photodiode (PD001) at output of HPL, main HPL power monitor | ADC_2_27 | V
{ifo}:PSL-PWR_NPRO | monitor photodiode at input of 35W amplifier: senses NPRO power | ADC_2_29 | W
{ifo}:PSL-MIS_FLOW | used to detect too low water flow through beam dumps and power meters | ADC_2_28
{ifo}:PSL-MIS_NPRO_RRO | binary signal for NPRO rf- power (in band around relaxation oscillation frequency) | ADC_2_30
{ifo}:PSL-MIS_SPARE |  | 0
35W front end - via EPICS-OPC connection from PSL-BECKHOFF system
{ifo}:PSL-AMP_DBHTSNKOVRTEMP | Amp. Heatsink Temperature Error | Beckhoff
{ifo}:PSL-AMP_DTEMPERR | Amp. Diode Temperature Error | Beckhoff
{ifo}:PSL-AMP_ENABLEPWRDOG | Amp. Enable Power Watchdog (>15% deviation) | Beckhoff
{ifo}:PSL-AMP_ERR | Amp. Error state positive | Beckhoff
{ifo}:PSL-AMP_LIDCLOSED | Amp. Lid Closed | Beckhoff
{ifo}:PSL-AMP_LIDERR | Amp. Lid opened without Lid override active | Beckhoff
{ifo}:PSL-AMP_MANUALMODE | Amp. Manual Mode (enables individual control of NPRO and amplifier currents) | Beckhoff
{ifo}:PSL-AMP_NPROOK | Amp. NPRO Running (from NPRO internal diagnostic) | Beckhoff
{ifo}:PSL-AMP_LIDERR_OVR | Amp. Lid Interlock Override acitve | Beckhoff
{ifo}:PSL-AMP_WATCHDOG | Amp. Power Watchdog triggered | Beckhoff
{ifo}:PSL-AMP_SHUTTERCLSD | Amp.Shutter Closed INPUT signal | Beckhoff
{ifo}:PSL-AMP_SHUTTEROPEN | Amp. Shutter Open INPUT Signal | Beckhoff
{ifo}:PSL-AMP_SYSTEMOK | Amp. System in automatic mode - Running | Beckhoff
{ifo}:PSL-AMP_WARN | Amp. Warning NPRO or water flow | Beckhoff
{ifo}:PSL-AMP_LOWFLOW | Amp. Water Flow Warning (< x l/min) | Beckhoff
{ifo}:PSL-AMP_FLOWERR | Amp. Water Flow Error (< x l/min) | Beckhoff
{ifo}:PSL-AMP_XTALOVRTEMP | Amp. crystal temperature error | Beckhoff
{ifo}:PSL-AMP_LIDCOUNT | Amp. Lid # of opening | Beckhoff
{ifo}:PSL-AMP_HRS | Amp. # of operation hours | Beckhoff | hours
{ifo}:PSL-AMP_DBHTSNKTEMP | Amp. Heatsink Temperature | Beckhoff | deg C
{ifo}:PSL-AMP_DCUR1 | Amp. Diode current (Diode 1+2) | Beckhoff | A
{ifo}:PSL-AMP_DCUR2 | Amp. Diode current (Diode 3+4) | Beckhoff | A
{ifo}:PSL-AMP_DCURSET1 | Amp. Diode current set point (Diode 1+2) | Beckhoff | A
{ifo}:PSL-AMP_DCURSET2 | Amp. Diode current set point (Diode 3+4) | Beckhoff | A
{ifo}:PSL-AMP_D1PWR | Amp. Diode power (internal monitor diode) | Beckhoff | W
{ifo}:PSL-AMP_D2PWR | Amp. Diode power (internal monitor diode) | Beckhoff | W
{ifo}:PSL-AMP_D3PWR | Amp. Diode power (internal monitor diode) | Beckhoff | W
{ifo}:PSL-AMP_D4PWR | Amp. Diode power (internal monitor diode) | Beckhoff | W
{ifo}:PSL-AMP_D1TEMP | Amp. Diode temperature | Beckhoff | deg C
{ifo}:PSL-AMP_D2TEMP | Amp. Diode temperature | Beckhoff | deg C
{ifo}:PSL-AMP_D3TEMP | Amp. Diode temperature | Beckhoff | deg C
{ifo}:PSL-AMP_D4TEMP | Amp. Diode temperature | Beckhoff | deg C
{ifo}:PSL-AMP_D1TEC | Amp. Diode temp. control signal | Beckhoff
{ifo}:PSL-AMP_D2TEC | Amp. Diode temp. control signal | Beckhoff
{ifo}:PSL-AMP_D3TEC | Amp. Diode temp. control signal | Beckhoff
{ifo}:PSL-AMP_D4TEC | Amp. Diode temp. control signal | Beckhoff
{ifo}:PSL-AMP_D1TEMPSET | Amp. Diode temp. set point | Beckhoff | deg C
{ifo}:PSL-AMP_D2TEMPSET | Amp. Diode temp. set point | Beckhoff | deg C
{ifo}:PSL-AMP_D3TEMPSET | Amp. Diode temp. set point | Beckhoff | deg C
{ifo}:PSL-AMP_D4TEMPSET | Amp. Diode temp. set point | Beckhoff | deg C
{ifo}:PSL-AMP_DVOLT1 | Amp. Diode voltage (Diode 1+2) | Beckhoff | V
{ifo}:PSL-AMP_DVOLT2 | Amp. Diode voltage (Diode 3+4) | Beckhoff | V
{ifo}:PSL-AMP_PWR1 | Amp. power monitor (behind stage 1) | Beckhoff | W
{ifo}:PSL-AMP_PWR2 | Amp. power monitor (behind stage 2) | Beckhoff | W
{ifo}:PSL-AMP_PWR3 | Amp. power monitor (behind stage 3) | Beckhoff | W
{ifo}:PSL-AMP_FLOW | Amp. cooling water flow rate | Beckhoff | lpm
{ifo}:PSL-AMP_XTALTEMP | Amp. crystal temperature | Beckhoff | deg C
{ifo}:PSL-NPRO_TEMPGUARD | NPRO Temperature Error | Beckhoff | K
{ifo}:PSL-NPRO_HRS | NPRO operation hours | Beckhoff | hours
{ifo}:PSL-NPRO_D1PWR | NPRO LD1 power monitor, sensed | Beckhoff | W
{ifo}:PSL-NPRO_D2PWR | NPRO LD2 power monitor, sensed | Beckhoff | W
{ifo}:PSL-NPRO_D1TEMPERR | NPRO LD1 Temperatures , error signal | Beckhoff | K
{ifo}:PSL-NPRO_D2TEMPERR | NPRO LD2 Temperatures , error signal | Beckhoff | K
{ifo}:PSL-NPRO_NEMON | NPRO Noise Eater monitor (NE DC signal) | Beckhoff | mA
{ifo}:PSL-NPRO_XTALTEMPERR | NPRO Xtal Temperature, error signal | Beckhoff | K
{ifo}:PSL-NPRO_D1TEMPSET | NPRO LD1 Temperatures, set | Beckhoff | deg C
{ifo}:PSL-NPRO_D2TEMPSET | NPRO LD2 Temperatures, set | Beckhoff | deg C
{ifo}:PSL-NPRO_D1TEMP | NPRO LD Temperatures, sensed | Beckhoff | deg C
{ifo}:PSL-NPRO_D2TEMP | NPRO LD Temperatures, sensed | Beckhoff | deg C
{ifo}:PSL-NPRO_XTALTEMPSET | NPRO Xtal Temperature, set | Beckhoff | deg C
{ifo}:PSL-NPRO_XTALTEMP | NPRO Xtal Temperature, sensed | Beckhoff | deg C
{ifo}:PSL-NPRO_CURSET | NPRO LD current, set | Beckhoff | A
{ifo}:PSL-NPRO_CUR | NPRO LD current, sensed | Beckhoff | A
Beckhoff
high power oscillator - via EPICS-OPC connection from PSL-BECKHOFF system |  | Beckhoff
Beckhoff
{ifo}:PSL-EPICSALARM | EPICS signal to close the high power laser shutter | Beckhoff
{ifo}:PSL-IL_DBTEMPGUARDS | hardware HPO laser diode temperature interlock | Beckhoff
{ifo}:PSL-IL_DCHILFLOW | hardware diode chiller flow interlock | Beckhoff
{ifo}:PSL-IL_DIODERMEMSTOP | hardware laser diode room interlock box safety button | Beckhoff
{ifo}:PSL-IL_DIODERMFACILITY | hardware laser diode room facility interlock | Beckhoff
{ifo}:PSL-IL_DIODERMKEYLOCK | hardware laser diode room, interlock-box key interlock | Beckhoff
{ifo}:PSL-IL_AMPTEMPGUARDS | hardware amplifier laser diode temperature interlock | Beckhoff
{ifo}:PSL-IL_LVEAEMSTOP | hardware laser area enclosure safety buttom interlock | Beckhoff
{ifo}:PSL-IL_LVEAFACILITY | hardware LVEA facility interlock | Beckhoff
{ifo}:PSL-IL_LVEAKEYLOCK | hardware LVEA control box key interlock | Beckhoff
{ifo}:PSL-IL_OK | no interlock | Beckhoff
{ifo}:PSL-IL_TECSTATUS | HPO laser diode temperature contol power supplies on | Beckhoff
{ifo}:PSL-IL_TWINSAFEFBERR | ??? Not clear ??? | Beckhoff
{ifo}:PSL-IL_XCHILFLOW | hardware crystal chiller flow interlock | Beckhoff
{ifo}:PSL-OSC_COUPLERSTATEERR | TwinCad interface error | Beckhoff
{ifo}:PSL-OSC_DCHILALARM | Alarm Diode Chiller | Beckhoff
{ifo}:PSL-OSC_LRARANGEWARN | LRA out of range | Beckhoff
{ifo}:PSL-OSC_DBHTSNKOVRTEMP | OscDBHeatsinkOvertemp | Beckhoff
{ifo}:PSL-OSC_DTEMPERR | OscDiodeOvertemp | Beckhoff
{ifo}:PSL-OSC_LRAON | OscEnableLRA | Beckhoff
{ifo}:PSL-OSC_LOCKON | OscEnableLocking | Beckhoff
{ifo}:PSL-OSC_PWRDOGON | OscEnablePowerWatchdog | Beckhoff
{ifo}:PSL-OSC_FLOWDOGON | OscEnableFlowWatchdog | Beckhoff
{ifo}:PSL-OSC_RAMPON | OscEnableRamp (injection locking) | Beckhoff
{ifo}:PSL-OSC_SW2ON | ??? Not longer used…? | Beckhoff
{ifo}:PSL-OSC_ERROR | OscError | Beckhoff
{ifo}:PSL-OSC_VBNOTALIVE | OscFiberSwitchVB not Alive | Beckhoff
{ifo}:PSL-OSC_HPSCLSD | OscHighPowerShutterClosed | Beckhoff
{ifo}:PSL-OSC_LIDCLSD | OscLidClosed | Beckhoff
{ifo}:PSL-OSC_LIDERR | OscLidError (were opened) | Beckhoff
{ifo}:PSL-OSC_LOCKED | OscLocked | Beckhoff
{ifo}:PSL-OSC_LPSCLSD | OscLowPowerShutterClosed | Beckhoff
{ifo}:PSL-OSC_LIDERR_OVR | OscOverrideLidInterlock | Beckhoff
{ifo}:PSL-OSC_PWRMETERLOWERR | OscPowermeterWaterFlowError | Beckhoff
{ifo}:PSL-OSC_PWRMETERLOWFLOW | OscPowermeterWaterFlowWarning | Beckhoff
{ifo}:PSL-OSC_WATCHDOG | OscPowerwatchdogError | Beckhoff
{ifo}:PSL-OSC_WARNING | OscWarning one other warnings active | Beckhoff
{ifo}:PSL-OSC_FLOWERR | OscWaterFlowError | Beckhoff
{ifo}:PSL-OSC_PRESS1ERR | OscPressure1MaxError | Beckhoff
{ifo}:PSL-OSC_LOWFLOW | OscWaterFlowWarning | Beckhoff
{ifo}:PSL-OSC_SERVICEMODE | OscServiceMode | Beckhoff
{ifo}:PSL-OSC_XCHILALARM | XChilAlarm | Beckhoff
{ifo}:PSL-OSC_DCHILHRS | DChilOpHrs | Beckhoff
{ifo}:PSL-OSC_DB1HRS | AdL-REF-OPC1.iOscDB1OpHrs | Beckhoff | hours
{ifo}:PSL-OSC_DB2HRS | AdL-REF-OPC1.iOscDB2OpHrs | Beckhoff | hours
{ifo}:PSL-OSC_DB3HRS | AdL-REF-OPC1.iOscDB3OpHrs | Beckhoff | hours
{ifo}:PSL-OSC_DB4HRS | AdL-REF-OPC1.iOscDB4OpHrs | Beckhoff | hours
{ifo}:PSL-OSC_LIDCOUNT | AdL-REF-OPC1.iOscLidCounter | Beckhoff
{ifo}:PSL-OSC_RELOCKCOUNT | AdL-REF-OPC1.iOscRelockCounter | Beckhoff
{ifo}:PSL-OSC_XCHILHRS | HPO crystal chiller operating hours | Beckhoff | hours
{ifo}:PSL-OSC_DCHILCND | HPO diode chiller, water conductivity | Beckhoff | uS
{ifo}:PSL-OSC_DCHILFLOW | HPO diode chiller, water flow | Beckhoff | lpm
{ifo}:PSL-OSC_DCHILTEMP | HPO diode chiller, water temperature | Beckhoff | deg C
{ifo}:PSL-OSC_DCHILTEMPSET | HPO diode chiller, water temperature, set value | Beckhoff | deg C
{ifo}:PSL-OSC_LRAPOS | HPO postion of long range actuator | Beckhoff | um
{ifo}:PSL-OSC_LRAPZTVOLTMAX | HPO trigger LRA movement (upper boundary) | Beckhoff | V
{ifo}:PSL-OSC_LRAPZTVOLTMIN | HPO trigger LRA movement (lower boundary) | Beckhoff | V
{ifo}:PSL-OSC_LRARANGE | HPO set maximum range for LRA with respect to reference point | Beckhoff | um
{ifo}:PSL-OSC_LRAREFPOINT | HPO LRA refence point | Beckhoff | um
{ifo}:PSL-OSC_LRASTEPSIZE | HPO LRA step size | Beckhoff | um
{ifo}:PSL-OSC_BASEPL_REARLTEMP | HPO Base Plate Temperature, sensor 1 | Beckhoff | deg C
{ifo}:PSL-OSC_BASEPL_CENTTEMP | HPO Base Plate Temperature, sensor 2 | Beckhoff | deg C
{ifo}:PSL-OSC_BASEPL_REARRTEMP | HPO Base Plate Temperature, sensor 3 | Beckhoff | deg C
{ifo}:PSL-OSC_BASEPL_FRONTLTEMP | HPO Base Plate Temperature, sensor 4 | Beckhoff | deg C
{ifo}:PSL-OSC_BASEPL_FRONTRTEMP | HPO Base Plate Temperature, sensor 5 | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_CNTWAVELENGTH | HPO center wavelength of light from diode box 1 | Beckhoff | nm
{ifo}:PSL-OSC_DB2_CNTWAVELENGTH | HPO center wavelength of light from diode box 2 | Beckhoff | nm
{ifo}:PSL-OSC_DB3_CNTWAVELENGTH | HPO center wavelength of light from diode box 3 | Beckhoff | nm
{ifo}:PSL-OSC_DB4_CNTWAVELENGTH | HPO center wavelength of light from diode box 4 | Beckhoff | nm
{ifo}:PSL-OSC_DB1_FWHM | HPO full width half max. of light from diode box 1 | Beckhoff | nm
{ifo}:PSL-OSC_DB2_FWHM | HPO full width half max. of light from diode box 2 | Beckhoff | nm
{ifo}:PSL-OSC_DB3_FWHM | HPO full width half max. of light from diode box 3 | Beckhoff | nm
{ifo}:PSL-OSC_DB4_FWHM | HPO full width half max. of light from diode box 4 | Beckhoff | nm
{ifo}:PSL-OSC_DB1_SMOMENT | HPO first moment of spectral distribution of light from diode box 1 | Beckhoff
{ifo}:PSL-OSC_DB2_SMOMENT | HPO first moment of spectral distribution of light from diode box 2 | Beckhoff
{ifo}:PSL-OSC_DB3_SMOMENT | HPO first moment of spectral distribution of light from diode box 3 | Beckhoff
{ifo}:PSL-OSC_DB4_SMOMENT | HPO first moment of spectral distribution of light from diode box 4 | Beckhoff
{ifo}:PSL-OSC_DB1_FMOMENT | HPO second moment of spectral distribution of light from diode box 1 | Beckhoff
{ifo}:PSL-OSC_DB2_FMOMENT | HPO second moment of spectral distribution of light from diode box 2 | Beckhoff
{ifo}:PSL-OSC_DB3_FMOMENT | HPO second moment of spectral distribution of light from diode box 3 | Beckhoff
{ifo}:PSL-OSC_DB4_FMOMENT | HPO second moment of spectral distribution of light from diode box 4 | Beckhoff
{ifo}:PSL-OSC_DB1_HTSNK1TEMP | HPO heat sink temperature diode box #, sensor # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_HTSNK2TEMP | HPO heat sink temperature diode box #, sensor # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_HTSNK1TEMP | HPO heat sink temperature diode box #, sensor # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_HTSNK2TEMP | HPO heat sink temperature diode box #, sensor # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_HTSNK1TEMP | HPO heat sink temperature diode box #, sensor # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_HTSNK2TEMP | HPO heat sink temperature diode box #, sensor # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_HTSNK1TEMP | HPO heat sink temperature diode box #, sensor # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_HTSNK2TEMP | HPO heat sink temperature diode box #, sensor # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_PWR | HPO pump light monitor at laser head 1 | Beckhoff | W
{ifo}:PSL-OSC_DB2_PWR | HPO pump light monitor at laser head 2 | Beckhoff | W
{ifo}:PSL-OSC_DB3_PWR | HPO pump light monitor at laser head 3 | Beckhoff | W
{ifo}:PSL-OSC_DB4_PWR | HPO pump light monitor at laser head 4 | Beckhoff | W
{ifo}:PSL-OSC_DB1_SPECPWR |  | Beckhoff
{ifo}:PSL-OSC_DB2_SPECPWR |  | Beckhoff
{ifo}:PSL-OSC_DB3_SPECPWR |  | Beckhoff
{ifo}:PSL-OSC_DB4_SPECPWR |  | Beckhoff
{ifo}:PSL-OSC_DB1_WPD | HPO pump light monitor at laser head 1, percentage of initial power | Beckhoff | %
{ifo}:PSL-OSC_DB2_WPD | HPO pump light monitor at laser head 2, percentage of initial power | Beckhoff | %
{ifo}:PSL-OSC_DB3_WPD | HPO pump light monitor at laser head 3, percentage of initial power | Beckhoff | %
{ifo}:PSL-OSC_DB4_WPD | HPO pump light monitor at laser head 4, percentage of initial power | Beckhoff | %
{ifo}:PSL-OSC_DB1_CUR | HPO Diode current, diode box 1 | Beckhoff | A
{ifo}:PSL-OSC_DB2_CUR | HPO Diode current, diode box 2 | Beckhoff | A
{ifo}:PSL-OSC_DB3_CUR | HPO Diode current, diode box 3 | Beckhoff | A
{ifo}:PSL-OSC_DB4_CUR | HPO Diode current, diode box 4 | Beckhoff | A
{ifo}:PSL-OSC_DB1_CURSET | HPO Diode current set point, diode box 1 | Beckhoff | A
{ifo}:PSL-OSC_DB2_CURSET | HPO Diode current set point, diode box 2 | Beckhoff | A
{ifo}:PSL-OSC_DB3_CURSET | HPO Diode current set point, diode box 3 | Beckhoff | A
{ifo}:PSL-OSC_DB4_CURSET | HPO Diode current set point, diode box 4 | Beckhoff | A
{ifo}:PSL-OSC_DB1_D1PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB1_D2PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB1_D3PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB1_D4PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB1_D5PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB1_D6PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB1_D7PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB2_D1PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB2_D2PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB2_D3PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB2_D4PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB2_D5PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB2_D6PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB2_D7PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB3_D1PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB3_D2PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB3_D3PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB3_D4PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB3_D5PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB3_D6PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB3_D7PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB4_D1PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB4_D2PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB4_D3PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB4_D4PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB4_D5PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB4_D6PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB4_D7PWR | HPO diode power internal monitor, diode box #, diode # | Beckhoff | W
{ifo}:PSL-OSC_DB1_D1TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D2TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D3TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D4TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D5TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D6TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D7TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D1TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D2TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D3TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D4TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D5TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D6TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D7TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D1TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D2TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D3TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D4TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D5TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D6TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D7TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D1TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D2TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D3TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D4TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D5TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D6TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D7TEMP | HPO diode temperature, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D1TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB1_D2TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB1_D3TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB1_D4TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB1_D5TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB1_D6TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB1_D7TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB2_D1TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB2_D2TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB2_D3TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB2_D4TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB2_D5TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB2_D6TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB2_D7TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB3_D1TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB3_D2TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB3_D3TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB3_D4TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB3_D5TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB3_D6TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB3_D7TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB4_D1TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB4_D2TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB4_D3TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB4_D4TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB4_D5TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB4_D6TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB4_D7TEC | HPO diode temperature control signal, diode box #, diode # | Beckhoff
{ifo}:PSL-OSC_DB1_D1TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D2TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D3TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D4TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D5TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D6TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_D7TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D1TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D2TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D3TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D4TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D5TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D6TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB2_D7TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D1TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D2TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D3TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D4TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D5TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D6TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB3_D7TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D1TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D2TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D3TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D4TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D5TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D6TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB4_D7TEMPSET | HPO diode temperature set point, diode box #, diode # | Beckhoff | deg C
{ifo}:PSL-OSC_DB1_VOLT | HPO supply voltage diode box # | Beckhoff
{ifo}:PSL-OSC_DB2_VOLT | HPO supply voltage diode box # | Beckhoff
{ifo}:PSL-OSC_DB3_VOLT | HPO supply voltage diode box # | Beckhoff
{ifo}:PSL-OSC_DB4_VOLT | HPO supply voltage diode box # | Beckhoff
{ifo}:PSL-OSC_DB1_D1WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB1_D2WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB1_D3WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB1_D4WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB1_D5WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB1_D6WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB1_D7WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB2_D1WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB2_D2WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB2_D3WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB2_D4WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB2_D5WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB2_D6WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB2_D7WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB3_D1WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB3_D2WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB3_D3WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB3_D4WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB3_D5WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB3_D6WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB3_D7WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB4_D1WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB4_D2WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB4_D3WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB4_D4WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB4_D5WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB4_D6WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_DB4_D7WPD | was ment to give the relative diode power of diode # in DB # (percentage), not used | Beckhoff
{ifo}:PSL-OSC_BOXHUM | HPO humidity in HPO laser box | Beckhoff | %
{ifo}:PSL-OSC_FLOWDOGWINDOW | OSC flow window for the watchdog (0,2l/min) | Beckhoff | lpm
{ifo}:PSL-OSC_PRESS1MAX | OSC Water pressure max () | Beckhoff | bar
{ifo}:PSL-OSC_HEAD1TEMP | HPO Laser Head Temperature, head # | Beckhoff | deg C
{ifo}:PSL-OSC_HEAD2TEMP | HPO Laser Head Temperature, head # | Beckhoff | deg C
{ifo}:PSL-OSC_HEAD3TEMP | HPO Laser Head Temperature, head # | Beckhoff | deg C
{ifo}:PSL-OSC_HEAD4TEMP | HPO Laser Head Temperature, head # | Beckhoff | deg C
{ifo}:PSL-OSC_HEAD1FLOW | HPO Laser Head water flow rate [l/min], head # | Beckhoff | lpm
{ifo}:PSL-OSC_HEAD2FLOW | HPO Laser Head water flow rate [l/min], head # | Beckhoff | lpm
{ifo}:PSL-OSC_HEAD3FLOW | HPO Laser Head water flow rate [l/min], head # | Beckhoff | lpm
{ifo}:PSL-OSC_HEAD4FLOW | HPO Laser Head water flow rate [l/min], head # | Beckhoff | lpm
{ifo}:PSL-OSC_PWR1 | HPO Laser Power, backward direction | Beckhoff | W
{ifo}:PSL-OSC_PWR2 | HPO Laser Power, forward direction (only valid if external shutter is closed) | Beckhoff | W
{ifo}:PSL-OSC_PWREXT | HPO reading of power meter externally connected | Beckhoff | W
{ifo}:PSL-OSC_PWR_SUM | HPO sum of all three power meter | Beckhoff | W
{ifo}:PSL-OSC_PDLOCKDC_PWR | HPO reading of injection locking photo diode DC, calibrated in Watt | Beckhoff | W
{ifo}:PSL-OSC_LOCKGAIN | HPO loop gain of injection locking servo, set point | Beckhoff | V
{ifo}:PSL-OSC_LOCKOFFSET | HPO error point of injection locking servo, set point | Beckhoff | V
{ifo}:PSL-OSC_LOCKREF | HPO injection locking auto lock reference level, set point | Beckhoff | V
{ifo}:PSL-OSC_PDLOCKDC | HPO reading of injection locking photo diode DC [Volt] | Beckhoff | V
{ifo}:PSL-OSC_PDLOCKDC_MEAN | HPO reading of injection locking photo diode DC, 2s average | Beckhoff | V
{ifo}:PSL-OSC_PWRMETERFLOW | HPO water flow through laser internal aux. beam dumps and power meter | Beckhoff | lpm
{ifo}:PSL-OSC_PZTVOLT | HPO injection locking control signal | Beckhoff | V
{ifo}:PSL-OSC_BOXTEMP | HPO air temperature in HPO laser box | Beckhoff | deg C
{ifo}:PSL-OSC_PRESS1 | HPO water pressure inlet at water manifold | Beckhoff | bar
{ifo}:PSL-OSC_PRESS2 | HPO water pressure return at water manifold | Beckhoff | bar
{ifo}:PSL-OSC_XCHILCND | HPO crystal chiller, water conductivity | Beckhoff | uS
{ifo}:PSL-OSC_XCHILFLOW | HPO crystal chiller, water flow | Beckhoff | lpm
{ifo}:PSL-OSC_XCHILTEMP | HPO crystal chiller, water temperature | Beckhoff | deg C
{ifo}:PSL-OSC_XCHILTEMPSET | HPO crystal chiller, water temperature, set value | Beckhoff | deg C
{ifo}:PSL-OSC_USER | name of laser operator who put laser into service mode | Beckhoff
derived channels (calculated and used in RT control modules and EPICS user screens)
{ifo}:PSL-FSS
{ifo}:PSL-FSS_AUTOLOCK_DELAY | time in between different autolock states | software
{ifo}:PSL-FSS_AUTOLOCK_ON | automatic lock request | software
{ifo}:PSL-FSS_AUTOLOCK_STATE | autolock state | software
{ifo}:PSL-FSS_COMMON_GAIN_CALI | common mode gain calibration | software | counts/dB
{ifo}:PSL-FSS_FAST_GAIN_CALI | fast gain calibration | software | counts/dB
{ifo}:PSL-FSS_GOOD_TEMP | NPRO temperature offset in the middle PZT ramp | software | deg C
{ifo}:PSL-FSS_LOOP_MODE | switch loop status (if autolock off) | software
{ifo}:PSL-FSS_MIXER_OFS_CALI | modulation-demodulation scheme to optimize the mixer offset | software
{ifo}:PSL-FSS_MIXER_OFS_MOD_AMP | modulation-demodulation scheme to optimize the mixer offset | software
{ifo}:PSL-FSS_MIXER_OFS_MOD_CLKGAIN | modulation-demodulation scheme to optimize the mixer offset | software
{ifo}:PSL-FSS_MIXER_OFS_MOD_COSGAIN | modulation-demodulation scheme to optimize the mixer offset | software
{ifo}:PSL-FSS_MIXER_OFS_MOD_FREQ | modulation-demodulation scheme to optimize the mixer offset | software
{ifo}:PSL-FSS_MIXER_OFS_MOD_LOWPASS | modulation-demodulation scheme to optimize the mixer offset | software
{ifo}:PSL-FSS_MIXER_OFS_MOD_PHASE | modulation-demodulation scheme to optimize the mixer offset | software
{ifo}:PSL-FSS_MIXER_OFS_MOD_SINGAIN | modulation-demodulation scheme to optimize the mixer offset | software
{ifo}:PSL-FSS_MIXER_PP | Mixer peak-to-peak value | software | V
{ifo}:PSL-FSS_OSCILLATION | detects oscillation of loop | software | V
{ifo}:PSL-FSS_OSCILLATION_THRES | threshold for oscillation detection | software | V
{ifo}:PSL-FSS_PC_PP | Pockels Cell peak-to-peak value | software | V
{ifo}:PSL-FSS_PZT_RAMP_ACTIVE | status of PZT ramp | software
{ifo}:PSL-FSS_PZT_RAMP_FRQ | frequency of PZT ramp | software | Hz
{ifo}:PSL-FSS_PZT_RAMP_MAX | maximum voltage of PZT ramp | software | V
{ifo}:PSL-FSS_PZT_RAMP_MIN | minimum voltage of PZT ramp | software | V
{ifo}:PSL-FSS_PZT_RAMP_MODEPOS | FSS lock acquisition variable | software
{ifo}:PSL-FSS_PZT_RAMP_PHASE | phase of PZT ramp | software | deg
{ifo}:PSL-FSS_PZT_RAMP_PHASEPOS | FSS lock acquisition variable | software | deg
{ifo}:PSL-FSS_RELOCK_COUNT | number of relocks | software | #
{ifo}:PSL-FSS_RELOCK_DAY | days since last relock | software | days
{ifo}:PSL-FSS_RELOCK_DUR | duration of last relock | software | months
{ifo}:PSL-FSS_RELOCK_HOUR | hours since last relock | software | hours
{ifo}:PSL-FSS_RELOCK_MIN | minutes since last relck | software | minutes
{ifo}:PSL-FSS_RELOCK_RESET | resets relock counter | software
{ifo}:PSL-FSS_RESONANT | indicates if reference cavity is resonant | software
{ifo}:PSL-FSS_RESONANT_THRES | threshold for resonant detection | software | V
{ifo}:PSL-FSS_TEMP | NPRO temperature offset | software | deg C
{ifo}:PSL-FSS_TEMP_FIXEDRAMP_DUR | duration of NPRO temperature ramp | software | s
{ifo}:PSL-FSS_TEMP_FIXEDRAMP_MAX | maximum value for NPRO temperature ramp | software | V
{ifo}:PSL-FSS_TEMP_FIXEDRAMP_MIN | minimum value for NPRO temperature ramp | software | V
{ifo}:PSL-FSS_TEMP_LOOP | filter module for temperature loop | software
{ifo}:PSL-FSS_TEMP_LOOP_ON | state of temperature loop | software
{ifo}:PSL-FSS_TEMP_LOOP_ON_REQUEST | requested state of temperature loop | software
{ifo}:PSL-FSS_TEMP_MODE | switch between different temperature ramps (if autolock off) | software
{ifo}:PSL-FSS_TEMP_MODE_REQUEST | requested temperature ramp | software
{ifo}:PSL-FSS_TEMP_SEARCH_LOOP | FSS lock acquisition variable | software
{ifo}:PSL-FSS_TEMP_SEARCH_MODEFOUND | indicates if reference cavity is resonant | software
{ifo}:PSL-FSS_TEMP_SEARCH_RAMP_DUR | duration of NPRO temperature ramp | software | s
{ifo}:PSL-FSS_TEMP_SEARCH_RAMP_MAX | maximum value for NPRO temperature ramp | software | V
{ifo}:PSL-FSS_TEMP_SEARCH_RAMP_MIN | minimum value for NPRO temperature ramp | software | V
{ifo}:PSL-FSS_TEST1_ON_CALI | calibration of Test1 | software
{ifo}:PSL-FSS_TEST2_ON_CALI | calibration of Test2 | software
{ifo}:PSL-FSS_TPD_PP | peak-to-peak value of reference cavity transmission photo diode | software | V
{ifo}:PSL-FSS_VCO_MODLEVEL_CALI | calibration of modulation level | software
{ifo}:PSL-FSS_VCO_TEST_ON_CALI | calibration for test inout of VCO | software
{ifo}:PSL-FSS_VCO_WIDE_ON | switch for wide-band-input of VCO | software
{ifo}:PSL-FSS_VCO_WIDE_ON_CALI | calibration of wide-band-input of VCO | software
{ifo}:PSL-ISS
{ifo}:PSL-ISS_AOM_DRIVER_MON | voltage inserted in AOM driver | software | V
{ifo}:PSL-ISS_AUTOLOCK_DELAY | time in between different autolock states | software | s
{ifo}:PSL-ISS_AUTOLOCK_ON | automatic lock request | software
{ifo}:PSL-ISS_AUTOLOCK_SIGN | no longer used | software
{ifo}:PSL-ISS_CTRL_OFFSET_CALI | calibration of Offset % to V to counts | software
{ifo}:PSL-ISS_DIFFRACTION | AOM diffracted power | software | %
{ifo}:PSL-ISS_DIFFRACTION_AVG | average AOM diffracted power (1 sec average) | software | %
{ifo}:PSL-ISS_DIFFRACTION_MAX | maximum AOM diffracted power (1 sec average) | software | %
{ifo}:PSL-ISS_DIFFRACTION_MIN | minimum AOM diffracted power (1 sec average) | software | %
{ifo}:PSL-ISS_DIFF_POLY_0 | 0th order of calibration polynominal for AOM diffracted power | software
{ifo}:PSL-ISS_DIFF_POLY_1 | 1st order of calibration polynominal for AOM diffracted power | software
{ifo}:PSL-ISS_DIFF_POLY_2 | 2nd order of calibration polynominal for AOM diffracted power | software
{ifo}:PSL-ISS_DIGITAL_LOOP | output of digital support loop | software
{ifo}:PSL-ISS_DIGITAL_SCALING | compensation for band-pass in the analog loop | software
{ifo}:PSL-ISS_GAIN_CALI | calibration of loop gain | software
{ifo}:PSL-ISS_INLOOP_PD_SELECT_CALI | calibration for {ifo}:PSL-ISS_INLOOP_PD_SELECT | software
{ifo}:PSL-ISS_INTERLOCK_LP | no longer used | software
{ifo}:PSL-ISS_INTERLOCK_RESET | no longer used | software
{ifo}:PSL-ISS_INTERLOCK_THRES | no longer used | software
{ifo}:PSL-ISS_INTERLOCK_TRIPPED | no longer used | software
{ifo}:PSL-ISS_LOOP_STATE_REQUEST | requested loop state | software
{ifo}:PSL-ISS_OSCILLATION_MON | oscillation monitor at error point | software
{ifo}:PSL-ISS_OSCILLATION_MON_BP | band pass for oscillation monitor | software
{ifo}:PSL-ISS_OSCILLATION_MON_LP | low pass for oscillation monitor | software
{ifo}:PSL-ISS_OSCILLATION_THRESHOLD | threshold for oscillation monitor | software
{ifo}:PSL-ISS_PDA | PD A filter output signal | software | V
{ifo}:PSL-ISS_PDA_AVG | averaged PD A filter output signal | software | V
{ifo}:PSL-ISS_PDA_CALI_AC | calibration of PD A filter output to AC value used for in-model noise-calculations | software
{ifo}:PSL-ISS_PDA_CALI_DC | calibration of PD A filter output to DC value | software
{ifo}:PSL-ISS_PDA_DUR | averaging time of PD A filter output signal | software | s
{ifo}:PSL-ISS_PDA_LSD | calculated linear spectral density | software | /rt(Hz)
{ifo}:PSL-ISS_PDA_LSD_BANDPASS | bandpass filter for calculating linear spectral density at different frequencies | software
{ifo}:PSL-ISS_PDA_LSD_INTEGRATION | integration time for calculating linear spectral density | software
{ifo}:PSL-ISS_PDA_MAX | maximum of averaged PD A filter output signal | software | V
{ifo}:PSL-ISS_PDA_MIN | minimum of averaged PD A filter output signal | software | V
{ifo}:PSL-ISS_PDA_REL | calculated relative power noise at PD A | software
{ifo}:PSL-ISS_PDB | PD B filter output signal | software | V
{ifo}:PSL-ISS_PDB_AVG | averaged PD B filter output signal | software | V
{ifo}:PSL-ISS_PDB_CALI_AC | calibration of PD B filter output to AC value used for in-model noise-calculations | software
{ifo}:PSL-ISS_PDB_CALI_DC | calibration of PD B filter output to DC value | software
{ifo}:PSL-ISS_PDB_DUR | averaging time of PD B filter output signal | software | s
{ifo}:PSL-ISS_PDB_LSD | calculated linear spectral density | software | /rt(Hz)
{ifo}:PSL-ISS_PDB_LSD_BANDPASS | bandpass filter for calculating linear spectral density at different frequencies | software
{ifo}:PSL-ISS_PDB_LSD_INTEGRATION | integration time for calculating linear spectral density | software
{ifo}:PSL-ISS_PDB_MAX | maximum of averaged PD B filter output signal | software | V
{ifo}:PSL-ISS_PDB_MIN | minimum of averaged PD B filter output signal | software | V
{ifo}:PSL-ISS_PDB_REL | calculated relative power noise at PD B | software
{ifo}:PSL-ISS_QPD_DX | difference of right half and left half of QPD | software
{ifo}:PSL-ISS_QPD_DY | difference of upper and lower half of QPD | software
{ifo}:PSL-ISS_REFSIGNAL_CALI | calibration of refence signal | software
{ifo}:PSL-ISS_REFSIGNAL_MON | reference signal monitor | software
{ifo}:PSL-ISS_SAT_DAY | days since last saturation event | software
{ifo}:PSL-ISS_SAT_DELAY | delay between the saturation event and the next attempt to lock | software
{ifo}:PSL-ISS_SAT_DUR | last lock aquisition duration | software | s
{ifo}:PSL-ISS_SAT_EVENTS | number of saturation events | software
{ifo}:PSL-ISS_SAT_HOUR | hours since last saturation event | software | hours
{ifo}:PSL-ISS_SAT_MIN | minutes since last saturation event | software | min
{ifo}:PSL-ISS_SAT_RESET | resets the saturation event counter | software
{ifo}:PSL-ISS_SAT_THRES | threshold for saturation monitor | software | V
{ifo}:PSL-ISS_SECONDLOOP_CLOSED_CALI | calibration of {ifo}:PSL-ISS_SECONDLOOP_CLOSED | software
{ifo}:PSL-ISS_SECONDLOOP_SIGNAL | input of second loop detector | software
{ifo}:PSL-ISS_SLOW_POWER_CONTROL | possibility to change reference signal | software
{ifo}:PSL-ISS_SLOW_POWER_CONTROL_MON | monitor of {ifo}:PSL-ISS_SLOW_POWER_CONTROL | software
{ifo}:PSL-ISS_TRANSFER1_A | signal infront of summation point in first loop | software
{ifo}:PSL-ISS_TRANSFER1_B | signal behind summation point in first loop | software
{ifo}:PSL-ISS_TRANSFER2_A | signal infront of summation point in second loop | software
{ifo}:PSL-ISS_TRANSFER2_B | signal behind summation point in second loop | software
{ifo}:PSL-ISS_TRANS_PWR | no longer used | software
{ifo}:PSL-MIS_CLOSE_SHUTTER_CTRL_ROOM | external shutter close request from control room | software
{ifo}:PSL-MIS_CLOSE_SHUTTER_IO | external shutter close request from IO | software
{ifo}:PSL-MIS_CLOSE_SHUTTER_ISC | external shutter close request from ISC | software
{ifo}:PSL-MIS_EPICSALARM_MON | status of epics alarm | software
{ifo}:PSL-MIS_FLOW_LOWER | lower boundary of flow watch dog | software
{ifo}:PSL-MIS_FLOW_OK | indicates if flow is in between the upper and lower flow boundary | software
{ifo}:PSL-MIS_FLOW_UPPER | upper boundary of flow watch dog | software
{ifo}:PSL-OSC_PD_AMP_AC | no longer used | software
{ifo}:PSL-OSC_PD_AMP_DC | DC value of Photo Diode to monitor amplifier power | software | W
{ifo}:PSL-OSC_PD_BP_AC | no longer used | software
{ifo}:PSL-OSC_PD_BP_DC | DC value of Photo Diode to monitor power at brewster plate | software | V
{ifo}:PSL-OSC_PD_INT_AC | no longer used | software
{ifo}:PSL-OSC_PD_INT_DC | DC value of Photo Diode to monitor resonator internal power | software | V
{ifo}:PSL-OSC_PD_ISO_AC | no longer used | software
{ifo}:PSL-OSC_PD_ISO_DC | DC value of Photo Diode to monitor power at faraday isolator | software | V
{ifo}:PSL-PMC
{ifo}:PSL-PMC_ALIGNRAMP_FREQ | frequency of alignment ramp | software | Hz
{ifo}:PSL-PMC_ALIGNRAMP_MAX | maximum value of alignment ramp amplitude | software | V
{ifo}:PSL-PMC_ALIGNRAMP_MIN | minimum value of alignment ramp amplitude | software | V
{ifo}:PSL-PMC_ALIGNRAMP_ON | alignment ramp state | software
{ifo}:PSL-PMC_BLANKING_CALI | on/off switch for servo (not used during autolock mode) | software
{ifo}:PSL-PMC_BOOST_CALI | on/off switch for integrators (not used during autolock mode | software
{ifo}:PSL-PMC_HEATER_CALI | calibrated signal to Heater | software
{ifo}:PSL-PMC_HEATER_MAX | maximum voltage to Heater | software | V
{ifo}:PSL-PMC_HEATER_OFFSET | Offset for temperature control loop | software
{ifo}:PSL-PMC_HEATER_POWER | percentage of current heating power with respect to maximum value | software | %
{ifo}:PSL-PMC_HEATER_POWER_LP | low pass for Heater power | software
{ifo}:PSL-PMC_HV_MON_AVG | averaged of HV monitor | software
{ifo}:PSL-PMC_HV_MON_FREQ_CALI | designated output for frequncy noise measurement | software
{ifo}:PSL-PMC_HV_MON_MAX | maximum averaged of HV monitor | software | V
{ifo}:PSL-PMC_HV_MON_MIN | minimum averaged of HV monitor | software | V
{ifo}:PSL-PMC_HV_REF | reference signal of temperature loop | software | V
{ifo}:PSL-PMC_GAIN_CALI | calibration of gain | software
{ifo}:PSL-PMC_INOFFSET_CALI | calibration of input offset for PMC locking loop | software
{ifo}:PSL-PMC_INOFFSET_MOD_AMP | modulation-demodulation scheme of input offset calibration | software
{ifo}:PSL-PMC_INOFFSET_MOD_CLKGAIN | modulation-demodulation scheme of input offset calibration | software
{ifo}:PSL-PMC_INOFFSET_MOD_COSGAIN | modulation-demodulation scheme of input offset calibration | software
{ifo}:PSL-PMC_INOFFSET_MOD_FREQ | modulation-demodulation scheme of input offset calibration | software
{ifo}:PSL-PMC_INOFFSET_MOD_LOWPASS | modulation-demodulation scheme of input offset calibration | software
{ifo}:PSL-PMC_INOFFSET_MOD_PHASE | modulation-demodulation scheme of input offset calibration | software
{ifo}:PSL-PMC_INOFFSET_MOD_SINGAIN | modulation-demodulation scheme of input offset calibration | software
{ifo}:PSL-PMC_LOCKED | loop state | software
{ifo}:PSL-PMC_LOCK_ON_CALI | calibration of requested loop state | software
{ifo}:PSL-PMC_MIN_TRIGGER | lower boundary for resonant condition | software | V
{ifo}:PSL-PMC_OSCILLATION_MON | oscillation monitor at error point | software
{ifo}:PSL-PMC_OSCILLATION_MON_BP | band pass for oscillation monitor | software
{ifo}:PSL-PMC_OSCILLATION_MON_LP | low pass for oscillation monitor | software
{ifo}:PSL-PMC_OSCILLATION_THRESHOLD | threshold for oscillation monitor | software
{ifo}:PSL-PMC_RAMP_ON_CALI | calibration of ramp on | software
{ifo}:PSL-PMC_REFL_PWR | PMC reflected power | software | W
{ifo}:PSL-PMC_REF_CALI | calibration of reference signal | software
{ifo}:PSL-PMC_RELOCK_COUNT | number of relocks | software
{ifo}:PSL-PMC_RELOCK_DAY | days since last relock | software | days
{ifo}:PSL-PMC_RELOCK_DUR | duration of last relock | software | s
{ifo}:PSL-PMC_RELOCK_HOUR | hours since last relock | software | hours
{ifo}:PSL-PMC_RELOCK_MIN | minutes since last relck | software | min
{ifo}:PSL-PMC_RELOCK_RESET | resets relock counter | software
{ifo}:PSL-PMC_RESONANT | indicates if PMC is resonant | software
{ifo}:PSL-PMC_RESONANT_THRES | threshold for resonant detection | software
{ifo}:PSL-PMC_TEMP_CTRL_SIGNAL | error point of temperature loop | software
{ifo}:PSL-PMC_TEMP_LOOP | requested loop state of temperature loop | software
{ifo}:PSL-PMC_TF_IN_ON_CALI | calibration of TF_IN | software
{ifo}:PSL-PWR_HPL_DC_LP | output power of high power oscillator | software
{ifo}:PSL-PWR_PMC_SUM | total power at PMC | software
{ifo}:PSL-DBB_AO_LIMITEDSLEW | see RT control model or EPICS screen | software
{ifo}:PSL-DBB_AO_NEXTCYCLING | see RT control model or EPICS screen | software
{ifo}:PSL-DBB_AO_PZTCYCLE | see RT control model or EPICS screen | software
{ifo}:PSL-DBB_AO_PZTCYCLEMAX | see RT control model or EPICS screen | software
{ifo}:PSL-DBB_AO_PZTCYCLEMIN | see RT control model or EPICS screen | software
{ifo}:PSL-DBB_AO_PZTCYCLING | see RT control model or EPICS screen | software
{ifo}:PSL-DBB_CTRL0_ADD | PZT HV signal for DBB modecleaner | software | cts
{ifo}:PSL-DBB_CTRL_ALIGN_HOLD | holds alignment PZT voltage constant when enabled | software
{ifo}:PSL-DBB_CTRL_CTRL0 | output of MC alignment control loop | software
{ifo}:PSL-DBB_CTRL_CTRL1X_CAL | calibration for alignment PZT 1x | software
{ifo}:PSL-DBB_CTRL_CTRL1X_EPS | calibration for alignment PZT 1x | software
{ifo}:PSL-DBB_CTRL_CTRL1X_MON | HV signal for alignment PZT 1x | software | cts
{ifo}:PSL-DBB_CTRL_CTRL1Y_CAL | calibration for alignment PZT 1y | software
{ifo}:PSL-DBB_CTRL_CTRL1Y_EPS | calibration for alignment PZT 1y | software
{ifo}:PSL-DBB_CTRL_CTRL1Y_MON | HV signal for alignment PZT 1y | software | cts
{ifo}:PSL-DBB_CTRL_CTRL2X_CAL | calibration for alignment PZT 2x | software
{ifo}:PSL-DBB_CTRL_CTRL2X_EPS | calibration for alignment PZT 2x | software
{ifo}:PSL-DBB_CTRL_CTRL2X_MON | HV signal for alignment PZT 2x | software | cts
{ifo}:PSL-DBB_CTRL_CTRL2Y_CAL | calibration for alignment PZT 2y | software
{ifo}:PSL-DBB_CTRL_CTRL2Y_EPS | calibration for alignment PZT 2y | software
{ifo}:PSL-DBB_CTRL_CTRL2Y_MON | HV signal for alignment PZT 2y | software | cts
{ifo}:PSL-DBB_CTRL_DCTRL1X | integrator for demodulated signal of autoalignment loop 1x | software
{ifo}:PSL-DBB_CTRL_DCTRL1Y | integrator for demodulated signal of autoalignment loop 1y | software
{ifo}:PSL-DBB_CTRL_DCTRL2X | integrator for demodulated signal of autoalignment loop 2x | software
{ifo}:PSL-DBB_CTRL_DCTRL2Y | integrator for demodulated signal of autoalignment loop 2y | software
{ifo}:PSL-DBB_CTRL_DMATRIX_1_1 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_1_2 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_1_3 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_1_4 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_2_1 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_2_2 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_2_3 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_2_4 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_3_1 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_3_2 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_3_3 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_3_4 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_4_1 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_4_2 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_4_3 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DMATRIX_4_4 | matrix element for assining QPD demodulated signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_DSWITCH | switch to enable/disable autoalignment loop | software
{ifo}:PSL-DBB_CTRL_DSWITCH_ON | state of autoalignment loop | software
{ifo}:PSL-DBB_CTRL_LOCKED | state of MC lock acquisition | software
{ifo}:PSL-DBB_CTRL_LOWER_THRES | lower threshold for MC locking range | software
{ifo}:PSL-DBB_CTRL_MOD1X_AMP | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1X_CLKGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1X_COSGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1X_FREQ | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1X_PHASE | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1X_SINGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1Y_AMP | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1Y_CLKGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1Y_COSGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1Y_FREQ | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1Y_PHASE | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD1Y_SINGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2X_AMP | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2X_CLKGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2X_COSGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2X_FREQ | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2X_PHASE | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2X_SINGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2Y_AMP | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2Y_CLKGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2Y_COSGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2Y_FREQ | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2Y_PHASE | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_MOD2Y_SINGAIN | modulation/demodulation scheme to compensate offsets in alignment loop 1x | software
{ifo}:PSL-DBB_CTRL_PRE_THRES | threshold for power on RPD | software
{ifo}:PSL-DBB_CTRL_QCTRL1X | integrator for prealignment loop 1x | software
{ifo}:PSL-DBB_CTRL_QCTRL1Y | integrator for prealignment loop 1y | software
{ifo}:PSL-DBB_CTRL_QCTRL2X | integrator for prealignment loop 2x | software
{ifo}:PSL-DBB_CTRL_QCTRL2Y | integrator for prealignment loop 2y | software
{ifo}:PSL-DBB_CTRL_QMATRIX_1_1 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_1_2 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_1_3 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_1_4 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_2_1 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_2_2 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_2_3 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_2_4 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_3_1 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_3_2 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_3_3 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_3_4 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_4_1 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_4_2 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_4_3 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QMATRIX_4_4 | matrix element for assining QPD DC signal to alignment PZT | software
{ifo}:PSL-DBB_CTRL_QSWITCH | switch to enable/disable prealignment loop | software
{ifo}:PSL-DBB_CTRL_QSWITCH_ON | state of prealignment loop | software
{ifo}:PSL-DBB_CTRL_RESONANT | indicates if MC is resonant | software
{ifo}:PSL-DBB_CTRL_THRESHOLD | threshold for resonant condition of MC | software
{ifo}:PSL-DBB_CTRL_UPPER_THRES | upper threshold for MC locking range | software
{ifo}:PSL-DBB_DBID | DBB serial number | software
{ifo}:PSL-DBB_FILTER_SWITCH1 | see RT control model or EPICS screen | software
{ifo}:PSL-DBB_FILTER_SWITCH2 | see RT control model or EPICS screen | software
{ifo}:PSL-DBB_INPUT_1 | auxilary input 1 | software
{ifo}:PSL-DBB_INPUT_2 | auxilary input 2 | software
{ifo}:PSL-DBB_INPUT_3 | auxilary input 3 | software
{ifo}:PSL-DBB_INPUT_4 | auxilary input 4 | software
{ifo}:PSL-DBB_INTERLOCK_BP | bandpass for interlock | software
{ifo}:PSL-DBB_INTERLOCK_HARDWARE | hardware interlock | software
{ifo}:PSL-DBB_INTERLOCK_LP | lowpass for interlock | software
{ifo}:PSL-DBB_INTERLOCK_SOFTWARE | software interlock | software
{ifo}:PSL-DBB_INTERLOCK_THRESHOLD | threshold for interlock | software
{ifo}:PSL-DBB_LENS1_CAL | calibration for lens position | software | mm/cts
{ifo}:PSL-DBB_LENS2_CAL | calibration for lens position | software | mm/cts
{ifo}:PSL-DBB_LENS_LATCH | switch to latch lenses | software
{ifo}:PSL-DBB_MAN_RAMP | slider to manually set HV voltage for MC PZT | software
{ifo}:PSL-DBB_MAN_RAMP_CAL | calibration for manually set HV voltage | software | cts/V
{ifo}:PSL-DBB_MODE_COUNTDOWN | time left until DBB will return to Standby | software
{ifo}:PSL-DBB_MODE_NUM | mode state | software
{ifo}:PSL-DBB_MODE_REQUEST | mode state request | software
{ifo}:PSL-DBB_MODE_RESET_TIMEOUT | time until DBB will return to Standby after switching mode state | software
{ifo}:PSL-DBB_MODE_RMT | remote control of DBB enabled/disabled | software
{ifo}:PSL-DBB_MOD_ON | modulation enabled/disabled | software
{ifo}:PSL-DBB_MON_HV | MC HV monitor | software | V
{ifo}:PSL-DBB_MON_HV_FSR | calibration of HV monitor in FSR | software | FSR/V
{ifo}:PSL-DBB_MON_PZT | sensitive AC coupled HV monitor for frequency noise measurement | software
{ifo}:PSL-DBB_MON_PZT_CAL | calibration for PZT monitor | software
{ifo}:PSL-DBB_MON_SHUTTER | shutter calibration | software
{ifo}:PSL-DBB_MON_SHUTTER_CLOSED | shutter state | software
{ifo}:PSL-DBB_PH_CAL | calibration for phase | software
{ifo}:PSL-DBB_QPD_1DX | demodulated signal of QPD 1dx | software
{ifo}:PSL-DBB_QPD_1DX_CAL | calibration for QPD 1dx | software
{ifo}:PSL-DBB_QPD_1DX_OFS | offset for QPD 1dx | software
{ifo}:PSL-DBB_QPD_1DY | demodulated signal of QPD 1dy | software
{ifo}:PSL-DBB_QPD_1DY_CAL | calibration for QPD 1dy | software
{ifo}:PSL-DBB_QPD_1DY_OFS | offset for QPD 1dy | software
{ifo}:PSL-DBB_QPD_1QS | sum of DC signals of QPD1 | software
{ifo}:PSL-DBB_QPD_1QX | DC signal of QPD 1dx | software
{ifo}:PSL-DBB_QPD_1QY | DC signal of QPD 1dy | software
{ifo}:PSL-DBB_QPD_2DX | demodulated signal of QPD 2dx | software
{ifo}:PSL-DBB_QPD_2DX_CAL | calibration for QPD 2dx | software
{ifo}:PSL-DBB_QPD_2DX_OFS | offset for QPD 2dx | software
{ifo}:PSL-DBB_QPD_2DY | demodulated signal of QPD 2dy | software
{ifo}:PSL-DBB_QPD_2DY_CAL | calibration for QPD 2dy | software
{ifo}:PSL-DBB_QPD_2DY_OFS | offset for QPD 2dy | software
{ifo}:PSL-DBB_QPD_2QS | sum of DC signals of QPD2 | software
{ifo}:PSL-DBB_QPD_2QX | DC signal of QPD 2dx | software
{ifo}:PSL-DBB_QPD_2QY | DC signal of QPD 2dy | software
{ifo}:PSL-DBB_QPD_CALIOSC_CLKGAIN | calibration for error signals | software
{ifo}:PSL-DBB_QPD_CALIOSC_COSGAIN | calibration for error signals | software
{ifo}:PSL-DBB_QPD_CALIOSC_FREQ | calibration for error signals | software
{ifo}:PSL-DBB_QPD_CALIOSC_SINGAIN | calibration for error signals | software
{ifo}:PSL-DBB_QPD_CALI_AGE | calibration for error signals | software
{ifo}:PSL-DBB_QPD_CALI_AMP | calibration for error signals | software
{ifo}:PSL-DBB_QPD_CALI_COUNTDOWN | calibration for error signals | software
{ifo}:PSL-DBB_QPD_CALI_DURATION | calibration for error signals | software
{ifo}:PSL-DBB_QPD_CALI_FACTOR | calibration for error signals | software
{ifo}:PSL-DBB_QPD_CALI_TRIGGER | calibration for error signals | software
{ifo}:PSL-DBB_QPD_DS_CAL | calibration for demodulated sum signals of QPD | software
{ifo}:PSL-DBB_QPD_MCALI_AGE | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_AGE_1X | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_AGE_1Y | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_AGE_2X | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_AGE_2Y | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_AMP | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_COUNTDOWN | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_DURATION | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_FACTOR_1X | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_FACTOR_1Y | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_FACTOR_2X | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_FACTOR_2Y | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MCALI_TRIGGER | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MOD1X_LP | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MOD1Y_LP | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MOD2X_LP | calibration for error signals | software
{ifo}:PSL-DBB_QPD_MOD2Y_LP | calibration for error signals | software
{ifo}:PSL-DBB_RAMP_LOCK_AMP | amplitude of ramp in lockmode | software
{ifo}:PSL-DBB_RAMP_LOCK_FREQ | frequency of ramp in lockmode | software
{ifo}:PSL-DBB_RAMP_SCAN_AMP | amplitude of ramp in scanmode | software
{ifo}:PSL-DBB_RAMP_SCAN_FREQ | frequency of ramp in scanmode | software
{ifo}:PSL-DBB_RPD_AC_CAL | calibration for RPD AC | software
{ifo}:PSL-DBB_RPD_CURRENT | photo current of RPD | software
{ifo}:PSL-DBB_RPD_DC_LP | lowpassed DC output of RPD | software
{ifo}:PSL-DBB_RPD_REL_PWR | RPD AC voltage divided by RPD DC voltage | software
{ifo}:PSL-DBB_RPD_SHOTNOISE | shotnoise level of RPD | software
{ifo}:PSL-DBB_SHUTTER | request for opening of shutter 1 or 2 | software
{ifo}:PSL-DBB_SHUTTER_DELAY | time delay between closing shutter 1 and opening shutter 2 and vice versa | software
{ifo}:PSL-DBB_TEM00_AVG | no longer used | software
{ifo}:PSL-DBB_TEM00_RPD | no longer used | software
{ifo}:PSL-DBB_TEM00_TPD | no longer used | software
{ifo}:PSL-DBB_TPD_40DB_INT | offset compensation for 40dB output | software
{ifo}:PSL-DBB_TPD_80DB_INT | offset compensation for 80dB output | software
{ifo}:PSL-DBB_TPD_VALUE | combined value from TPD 0dB, 40dB and 80dB output | software
{ifo}:PSL-DBB_VERSION | software version | software