LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1000325 	        LIGO                          August 16, 2010

aLIGO HAM-ISI, Pre integration Test report, Phase 1,

LHO Unit #2

Eric Allwine, Hugh Radkins, Corey Gray, Jeffrey Garcia, Fabrice Matichard, Vincent Lhuillier

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

.

Table of contents:

Introduction	2

I.	Pre-Assembly Testing	3

	Step 1: Position Sensors	3

	Step 2: GS13	4

	Step 3: Actuators	4

II.	Tests to be performed during assembly	5

	Step 1: Check torques on all bolts	5

	Step 2: Check gaps under Support Posts	5

	Step 3: Pitchfork/Boxwork flatness before Optical Table install	5

	Step 4: Blade spring profile	5

	Step 5: Gap checks on actuators	6

	Step 6: Check level of Stage 0	6

	Step 7: Mass budget	6

	Step 8: Lockers adjustment	8

III.	Tests to be performed after assembly	9

	Step 1 - Actuators Inventory	9

	Step 2 - Sensors Inventory	9

	Step 3 - Electronics Inventory	10

	Step 4 - Set up sensors gap	10

	Step 5 - Measure the Sensor gap	11

	Step 6 - Check Sensor gaps after the platform release	11

	Step 7 - Check range of motion (hand pushing)	12

o	Step 7.1 – Test Nº1	12

o	Step 7.2 – Test Nº2	12

	Step 8 - Capacitive position sensor Power Spectrum	13

	Step 9 - GS13 power spectrum	14

	Step 10 - Coil Driver, cabling and resistance check	16

	Step 11 - Actuators Sign and range of motion (Local drive)	16

o	Step 11.1 - Actuators sign	16

o	Step 11.2 - Range of motion - Local drive	17

	Step 12 - Vertical Sensor Calibration	17

	Step 13 - Vertical Spring Constant	18

	Step 14 - Static Testing (Tests in the local basis)	18

	Step 15 - Linearity test (might not be necessary)	19

	Step 16 – Static tests in the general coordinate basis	20

	Step 17 - Frequency response	21

o	Step 17.1 - Local to local measurements	21

o	Step 17.2 - Cartesian to Cartesian measurements	22

	Step 18 - Transfer function comparison	24

o	Step 18.1 - Local to local - Comparison with HAM6	24

o	Step 18.2 - Cartesian toCartesian - Comparison with HAM6	25

	Step 19 - Lower Zero Moment Plan	26

	Step 20 - Damping loops	27

o	Step 20.1 - Transfer functions - Simulation	27

o	Step 20.2 - Powerspectrum – Experimental	28

IV.	Testing Summary	30

Introduction

This document presents the tests performed to characterize and validate the “HAM-ISI LLO Unit #1”. This unit is the first unit assembled for aLIGO at LLO.

The procedure document used to perform this test is:

-   E1000309 - aLIGO HAM-ISI, Pre-Integration Testing Procedure, Phase I (post assembly, before storage)

Other useful information can be found in:

-  E1000300 - HAM-ISI LLO test stand: software and electronic check

-

# Pre-Assembly Testing

## Step 1: Position Sensors

Note: The back panel reads 0.508V/0.001"

| S/N sensor | S/N board | ADE Gap Standoff(mm) | Location on the Jig | Gap Standoff on Jig(mm/in) | Voltage before zeroing | Voltage after zeroing. Prebake | Voltage after zeroing. Post bake |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 11999 | NR | NR | NR | NR | NR | NR | NR |
| 11987 | NR | NR | NR | NR | NR | NR | NR |
| 12041 | NR | NR | NR | NR | NR | NR | NR |
| 11998 | NR | NR | NR | NR | NR | NR | NR |
| 11990 | NR | NR | NR | NR | NR | NR | NR |
| 12048 | NR | NR | NR | NR | NR | NR | NR |

NR: not recorded

Will be measured for the next units.

Sensors noise spectra measured before baking:

Figure - H1 and V1 sensor noise                                                 Figure - H2 and V2 sensor noise

Figure  - H3 and V3 sensor noise

Acceptance Criteria: To be defined.

Issues/difficulties/comments regarding this test: Values of sensor gaps and zeroing were nor recorded. Waived for this unit.

Test result:	Passed:       	Failed:         .

## Step 2: GS13

All the data related to GS-13 post podding testing can be found in the SVN at : SeismicSVN\seismic\Common\Data\aLIGO_GS13_TestData\PostMod_TestResults_PDFs.

E1000058 spreadsheet provides the status of each individual GS-13 at LLO site during aLIGO HAM assembly

Acceptance Criteria: To be defined.

Test result:	Passed:       	Failed:         .

## Step 3: Actuators

Actuator data can be found at: T0900564. Actuator inventory is made at Section II – Step 1.

Acceptance Criteria: To be defined.

Test result:	Passed:       	Failed:         .

# Tests to be performed during assembly

## Step 1: Check torques on all bolts

Acceptance Criteria:

- All bolts should trip the wrench, and start moving immediately after. If any bolts in a pattern move before torque is reached, recheck after all bolts are brought to spec.

Test result:	Passed:   X    	Failed:         .

## Step 2: Check gaps under Support Posts

Acceptance Criteria:

- The test is passed if: a 0.001 inch shim cannot be passed freely through any connection to Stage 0 or between post and gussets. If shim can pass through, loosen all constraining bolts, and then retighten iteratively from the center of the part to the edges. Retest.

Test result:	Passed:   X   	Failed:         .

## Step 3: Pitchfork/Boxwork flatness before Optical Table install

Acceptance Criteria:

- Shim inserted won’t pass between parts.

Test result:	Passed:   X   	Failed:         .

## Step 4: Blade spring profile

| Blade # | Base (") | Middle (") | Tip(") | Flatness (mils) |
| --- | --- | --- | --- | --- |
| 1 | 0.376 |  | 0.371 | 0.005 |
| 2 | 0.391 |  | 0.372 | 0.019 |
| 3 | 0.38 |  | 0.373 | 0.007 |

Table 1 - Blade profile

Acceptance Criteria

- Blades must be flat within 0.020" inches. To be approved

Test result:	Passed:   X   	Failed:         .

## Step 5: Gap checks on actuators

Acceptance Criteria

- Gaps must be within 0.010” of design (i.e. 0.090” and .070” pass, but 0.095” and 0.065” doesn’t).

Test result:	Passed:   X   	Failed:         .

## Step 6: Check level of Stage 0

Acceptance Criteria To be defined.

Test result:	Passed:   X   	Failed:         .

## Step 7: Mass budget

| Optical Mass (Kg) | Wall Mass (Kg) | Keel Mass (Kg) | Total Mass (Kg) |
| --- | --- | --- | --- |
| 313.40 | 173.93 | 89.65 | 576.97 |

Table - Masses distribution

The total weight is 577Kg.

Figure - Masses distribution

Figure – Keel Masses and Wall masses location

|  |  | Mass |  |
| --- | --- | --- | --- |
|  | Type | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |
|  | Weight | 0.246 | 0.512 | 0.992 | 2.052 | 3.558 | 7.028 | 12.268 | Kg |
| Location | w1 |  |  |  |  |  | 1 | 1 | 19.3 |
|  | w2 |  |  |  |  |  | 1 | 1 | 19.3 |
|  | w3 |  |  |  |  | 2 |  | 1 | 19.4 |
|  | w4 |  |  |  |  |  | 1 | 1 | 19.3 |
|  | w5 |  |  |  |  |  | 1 | 1 | 19.3 |
|  | w6 |  |  |  |  | 2 |  | 1 | 19.4 |
|  | w7 |  |  |  |  |  | 1 | 1 | 19.3 |
|  | w8 |  |  |  |  |  | 1 | 1 | 19.3 |
|  | w9 |  |  |  |  | 2 |  | 1 | 19.4 |
|  |  | 0 | 0 | 0 | 0 | 4 | 6 | 9 | 173.9 |
| Location | k1 |  |  |  |  | 1 |  | 1 | 15.8 |
|  | k2 |  |  |  |  |  | 2 |  | 14.1 |
|  | k3 |  |  |  |  | 1 |  | 1 | 15.8 |
|  | k4 |  |  |  |  |  | 2 |  | 14.1 |
|  | k5 |  |  |  |  | 1 |  | 1 | 15.8 |
|  | k6 |  |  |  |  |  | 2 |  | 14.1 |
|  |  | 0 | 0 | 0 | 0 | 3 | 6 | 3 | 89.6 |

Table – Wall masses and Keel masses distribution

Locker shim thickness indicated in mils.

| Lockers | Shim thickness (mil) |
| --- | --- |
| A | 128 |
| B | 127 |
| C | 120 |
| D | 121 |

Table – Shims Thickness

Acceptance Criteria

- The total weight must be Kg +/- Kg

Test result:	Passed:       	Failed:         .

## Step 8: Lockers adjustment

| D.I at Lockers | Dial indicators V | Dial indicators H |
| --- | --- | --- |
| A | 0 | -0.5 |
| B | 0 | 0.5 |
| C | -1 | -0.5 |
| D | 0 | 0 |

Table – Dial indicators read-out

Issues/difficulties encountered during this test :

Acceptance Criteria

- Vertical and horizontal displacement near the lockers must be lower than 2 mils

Test result:	Passed:   X    	Failed:         .

# Tests to be performed after assembly

## Step 1 - Actuators Inventory

| Actuator | Actuator S/N |
| --- | --- |
| H1 | 26 |
| H2 | 12 |
| H3 | 1 |
| V1 | 16-L038 |
| V2 | 20-L040 |
| V3 | 14-L058 |

Table - Actuators' inventory

Issues/difficulties/comments regarding this test: Horizontal actuators S/N missing

Acceptance Criteria: all S/N must be recorded.

Test result:	Passed:   X   	Failed:        .

## Step 2 - Sensors Inventory

| Sensor | CPS S/N | Mount S/N | ADE board serial # |
| --- | --- | --- | --- |
| H1 | 11999 |  |  |
| H2 | 11987 |  |  |
| H3 | 12041 |  |  |
| V1 | 11998 |  |  |
| V2 | 11990 |  |  |
| V3 | 12048 |  |  |

Table - Capacitive position sensors' inventory

| Geophones GS13 | Serial Number | POD | Adaptor |
| --- | --- | --- | --- |
| H1 | 799 | 58 | 32 |
| H2 | 803 | 13 | 42 |
| H3 | 805 | 68 | 33 |
| V1 | 701 | 40 | 40 |
| V2 | 695 | 49 | 41 |
| V3 | 705 | 59 | 37 |

Table - Geophones' inventory

Issues/difficulties/comments regarding this test:

Test result:	Passed:   X    	Failed:         .

## Step 3 - Electronics Inventory

Write down serial number of coil driver, Anti aliasing chassis, Anti image chassis and interface chassis used for this test are listed below:

| Hardware | Ligo reference | S/N |
| --- | --- | --- |
| Coil driver | D0902744 | S1000266 |
|  |  | S1000269 |
| Anti Image filter | D070081 | S1000250 |
| Anti aliasing filter | D1000269 | S1000247 |
|  |  | S1000246 |
| Interface chassis | D1000067 | S1000270 |
|  |  | S1000271 |
|  |  | S1000272 |

Table - Inventory electronics

Issues/difficulties/comments regarding this test:

Test result:	Passed:   X   	Failed:         .

## Step 4 - Set up sensors gap

|  | 10 Kg masses at each corners | No mass | No mass |
| --- | --- | --- | --- |
| Table locked | ADE boxes on | ADE boxes on | ADE boxes on/off |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation |
| H1 | -455 | 1.0 | 100 | 0.8 | 102 | 0.8 |
| H2 | -80 | 0.9 | -107 | 0.7 | -110 | 0.8 |
| H3 | 293 | 1.0 | -36 | 1.0 |  |  |
| V1 | -144 | 0.6 | 28 | 1.1 | 30 | 1.4 |
| V2 | -248 | 1.2 | -27 | 1.6 | -27 | 1.6 |
| V3 | 537 | 1.0 | 449 | 1.8 |  |  |

Table - Set-up sensors gap

Issues/difficulties/comments regarding this test: HAM-ISI – LHO unit#2 uses synchronized satellite boxes

Acceptance criteria:

- All mean values must be lower than 400 cts (a bit less than .0005”).

- All standard deviations below 5 counts.

Test result:	Passed:   X   	Failed:         .

## Step 5 - Measure the Sensor gap

| Sensors | Gap measured on the Jig | Gap measured on the table | % of change | Offset sensors (counts) |
| --- | --- | --- | --- | --- |
| H1 | NR | 0.080" |  | 100 |
| H2 | NR | 0.080" |  | -107 |
| H3 | NR | 0.080" |  | -36 |
| V1 | NR | 0.080" |  | 28 |
| V2 | NR | 0.080" |  | -27 |
| V3 | NR | 0.080" |  | 449 |

Table  - Sensors gap

Issues/difficulties/comments regarding this test: No information of gaps measured on the Jig. Difficult to measure without scratching the target.

Acceptance criteria:

- Change of gaps lower than 2% (reference is the gap measured on the Jig)

Test result:	Passed:     X   	Failed:         .

## Step 6 - Check Sensor gaps after the platform release

|  | Table locked | Table unlocked |
| --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Diff |
| H1 | 100 | 0.8 | -511 | -611 |
| H2 | -107 | 0.7 | -1215 | -1108 |
| H3 | -36 | 1.0 | -1367 | -1331 |
| V1 | 28 | 1.1 | 866 | 838 |
| V2 | -27 | 1.6 | 369 | 396 |
| V3 | 449 | 1.8 | 60 | -389 |

Table – Sensor gaps after platform release

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Absolute values of the difference between the unlocked and the locked table must be below:

  - 1600 cts for horizontal sensors (~0.002”)

  - 1600 cts for vertical sensors (~0.002”)

- Considering the acceptance criteria of step 4, all mean values must be lower than

  - 2000 cts for horizontal sensors (~0.0025”)

  - 2000 cts for vertical sensors (~0.0025”)

Test result:	Passed:   X   	Failed:         .

## Step 7 - Check range of motion (hand pushing)

## Step 7.1 – Test Nº1

|  | CPS read out | Calculated after calibration |
| --- | --- | --- |
| Sensors | UP (Counts) | Down (Counts) | UP (mil) | Down (mil) |
| V1 | 20600 | -19900 | 24.7 | -23.9 |
| V2 | 18900 | -19990 | 22.7 | -24.0 |
| V3 | 20280 | -18500 | 24.3 | -22.2 |
|  | CPS read out | Calculated after calibration |
| Sensors | CW(-RZ) | CCW (+RZ) | CW (mil) | CCW (mil) |
| H1 | 22350 | -23600 | 26.8 | -28.3 |
| H2 | 19200 | -21400 | 23.1 | -25.7 |
| H3 | 23350 | -19600 | 28.0 | -23.5 |

Table - Optic table range of motion

Issues/difficulties/comments regarding this test: the results indicate that all displacement sensor signals have the same sign all the way through the read chain.

## Step 7.2 – Test Nº2

Displacement sensors and actuator gap check:

|  | Push in positive direction | Push in negative direction | Railing | Actuator Gap Check |
| --- | --- | --- | --- | --- |
| H1 | 25177 | -24906 |  | X |
| H2 | 24179 | -26046 |  | X |
| H3 | 26335 | -28076 |  | X |
| V1 | 20883 | -19813 |  | X |
| V2 | 32767 | -32768 | X | X |
| V3 | 25209 | -25426 |  | X |

Table - Sensor and Actuator gap check

Acceptance criteria:

- The vertical sensor readout be positive when the optic table is pushed in the +Z direction

- The horizontal sensor readout be negative when the optic table is pushed in the +RZ direction

- Step 7.1

  - Absolutes value of all estimated motions must be higher than 16000counts (~0.020”)

- Step 7.2

  - No contact point on sensors

  - Absolute value of sensor read out must be higher than 16000counts (~0.020”)

  - No contact point on actuators

Test result:	Passed:   X   	Failed:         .

## Step 8 - Capacitive position sensor Power Spectrum

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/DTT_Powerspectrum_Undamped

- Power_spectrum_table_locked_ADE_Sync_082410.xml

- Power_spectrum_table_unlocked_ADE_Sync_082410.xml

- Power_spectrum_table_locked_ADE_Sync_082410.txt

- Power_spectrum_table_unlocked_ADE_Sync_082410.txt

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/DTT_Powerspectrum_Undamped

- PSD_CPS_GS13_lock_unlock_HAM_ISI_LHO_unit_2_2010_08_24.m

Figures in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Figures/Powerspectrum_GS13_CPS

- Calibrated_CPS_Powerspectrum_Locked_Unlocked.fig

CPS calibration:

The CPS power spectrums are calibrated by using a sensitivity of 30.2nm/count.

Figure - Calibrated CPS power spectrum

Acceptance criteria:

- No cross talk (peaks at low frequencies + harmonics on measurements)

- Magnitude of power spectrums must be

  - Less than 5e-10m/√Hz at 0.1Hz

  - Less than 5e-10m/√Hz at 1Hz

  - Less than 5e-10m/√Hz at 10Hz

Test result:	Passed:       	Failed:          .

## Step 9 - GS13 power spectrums

## Step 9.1 –Table locked and unlocked (free)

Figure – Power spectrum Calibrated GS13

## Step 9.2 – Table tilted

The figure below presents the GS13 powerspectrum when the table is unlocked and loaded with a 20Kg mass at one of its corner.

Figure – Power spectrum Calibrated GS13 with mass at corner

Data files in SVN at:

opt//svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_1/DTT_Powerspectrum_Undamped

- Power_spectrum_table_locked_ADE_Sync_082410.xml

- Power_spectrum_table_unlocked_ADE_Sync_082410.xml

- Power_spectrum_table_locked_ADE_Sync_082410.txt

- Power_spectrum_table_unlocked_ADE_Sync_082410.txt

- Power_spectrum_table_unlocked_Mass_corner_A_ADE_Sync_082410.xml

- Power_spectrum_table_unlocked_Mass_corner_B_ADE_Sync_082410.xml

- Power_spectrum_table_unlocked_Mass_corner_C_ADE_Sync_082410.xml

- Power_spectrum_table_unlocked_Mass_corner_D_ADE_Sync_082410.xml

- Power_spectrum_table_unlocked_Mass_corner_E_ADE_Sync_082410.xml

- Power_spectrum_table_unlocked_Mass_corner_F_ADE_Sync_082410.xml

- Power_spectrum_table_unlocked_Mass_corner_A_ADE_Sync_082410.txt

- Power_spectrum_table_unlocked_Mass_corner_B_ADE_Sync_082410.txt

- Power_spectrum_table_unlocked_Mass_corner_C_ADE_Sync_082410.txt

- Power_spectrum_table_unlocked_Mass_corner_D_ADE_Sync_082410.txt

- Power_spectrum_table_unlocked_Mass_corner_E_ADE_Sync_082410.txt

- Power_spectrum_table_unlocked_Mass_corner_F_ADE_Sync_082410.txt

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_1/DTT_Powerspectrum_Undamped

- PSD_CPS_GS13_Lock_Unlock_HAM_ISI_LHO_Unit_2_2010_08_24.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Figures/Powerspectrum_GS13_CPS

- Calibrated_Geophones_Powerspectrum_Locked_Unlocked.fig

- Calibrated_Geophones_Powerspectrum_with_20Kg_Mass_at_one corner.fig

Note: Powerspr

Acceptance criteria:

- Table locked

  - Less than xx at 0.1Hz

  - Less than xx at 1Hz

  - Less than xx at 10Hz

- Table unlocked (no tilt)

  - Less than xx at 0.1Hz

  - Less than xx at 1Hz

  - Less than xx at 10Hz

- Table unlocked (tilt Rx & Ry)

  - Less than xx at 0.1Hz

  - Less than xx at 1Hz

  - Less than xx at 10Hz

Test result:	Passed:       	Failed:       .

## Step 10 - Coil Driver, cabling and resistance check

| Actuator | V1 | H1 | V2 |
| --- | --- | --- | --- |
| Coil driver | S1000266 - Coarse 1 OUT | S1000266 - Fine 1 OUT | S1000266 - Fine 2 OUT |
| Cable # | 25 | 26 | 27 |
| Resistance (Ohm) | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 |
|  | 6.5 | O.L (infinity) | 6.5 | O.L (infinity) | 6.5 | O.L (infinity) |
| MEDM offset (1000 counts) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) |
|  | 0.303 V | 0.306 V | 0.305 V |
| Actuator | H2 | V3 | H3 |
| Coil driver | S1000266 - Coarse 2 OUT | S1000267 - Coarse 1 OUT | S1000267 - Fine 1 OUT |
| Cable # | 28 | 29 | 30 |
| Resistance (Ohm) | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 |
|  | 6.5 | O.L (infinity) | 6.5 | O.L (infinity) | 6.5 | O.L (infinity) |
| MEDM offset (1000 counts) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) |
|  | 0.307 V | 0.298 V | 0.301 V |

Table - Actuators resistance check

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- The measured resistance between the middle pin and one side pin must be 6.5 +/-1 ohms

- Actuator neutral pins must be connected on pin #1 (left side pin of the plug)

- Actuator drive pins must be connected on pin #2 (middle pin of the plug)

- Actuator ground shield pins must be connected on pin #3 (right pin of the plug)

- All LEDs on the coil driver front panel must be green

Test result:	Passed:   X   	Failed:         .

## Step 11 - Actuators Sign and range of motion (Local drive)

## Step 11.1 - Actuators sign

Acceptance criteria: A positive offset drive on one actuator must give positive sensor readout on the collocated sensor. Signs will also be tested when measuring local to local transfer functions.

Test result:	Passed:   X    	Failed:         .

## Step 11.2 - Range of motion - Local drive

|  | Positive and negative offset Drive(+/-30000 counts) |
| --- | --- |
|  | V1 | V2 | V3 | H1 | H2 | H3 |
| V1 readout (counts) | 20088 | -1989 | -8540 | X | X | X |
|  | -19439 | -353 | 8685 |  |  |  |
| V2 readout (counts) | -6256 | 26420 | -8118 | X | X | X |
|  | 4818 | -24717 | 7165 |  |  |  |
| V3 readout (counts) | 6022 | -15782 | 21595 | X | X | X |
|  | -6601 | 15201 | -22703 |  |  |  |
| H1 readout (counts) | X | X | X | 23863 | 15560 | 16362 |
|  |  |  |  | -24419 | -17650 | -17558 |
| H2 readout (counts) | X | X | X | 17716 | 23311 | 17221 |
|  |  |  |  | -16627 | -24489 | -16864 |
| H3 readout (counts) | X | X | X | 13668 | 13626 | 25018 |
|  |  |  |  | -13163 | -13296 | -25206 |

Table - Range of motion - Local drive

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Main couplings read out must be at least 16000 counts (~0.002”)

- Signs of actuators drive and sensors read out have to be the same

Test result:	Passed:   X   	Failed:         .

## Step 12 - Vertical Sensor Calibration

| Lockers | D.I readout with for a negative drive | D.I readout without any drive | D.I readout with for a positive drive |  |
| --- | --- | --- | --- | --- |
| A | -18 | 0 | 21 |  |
| B | -18.5 | 0 | 20 |  |
| C | -18 | 0 | 19.5 |  |
| D | -18 | 0 | 19.5 |  |
| Average | -18.125 | 0 | 20 | 38.1 |
| Sensors | Counts | Counts | Counts | Difference (Counts) |
| V1 | -15697 | 16 | 15807 | 31504 |
| V2 | -15577 | -6 | 15782 | 31359 |
| V3 | -16502 | -328 | 15895 | 32397 |
|  |  |  | Average | 31753 |

Table - Calibration of capacitive position sensors

Vertical sensitivity:	31753/38.1 = 832.9 count/mil

or	832.9 count/mil * 1/1638 V/count = 0.508V/mil

or	25400nm/mil * 1/832.9 mil/count = 30.5 nm/count

Difference with Nominal sensitivity = (832.9-840)/840=-0.85%

Issues/difficulties/comments regarding this test:

Acceptance criteria: Deviation from nominal value < 2%. Nominal value is 840 count/mil.

Test result:	Passed:   X   	Failed:          .

## Step 13 - Vertical Spring Constant

| Sensors | Mean diff counts | Mean diff m | K (N/m) |
| --- | --- | --- | --- |
| V1 | 7981 | 2.41E-04 | 8.30E+04 |
| V2 | 8017 | 2.42E-04 | 8.26E+04 |
| V3 | 7955 | 2.40E-04 | 8.33E+04 |
|  |  | Average (N/m) | 2.49E+05 |

Table - Vertical spring constant

The measured vertical stiffness is 2.489e5N/m. The nominal vertical spring constant is 2.428e5N/m. The measured error on the vertical stiffness is 2.53%.

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Spring constant is within +/- 10/-1% of 2.428e5 N/m (HPD FEA Results).

Test result:	Passed:   X    	Failed:         .

## Step 14 - Static Testing (Tests in the local basis)

|  |  | Sensors (counts) |
| --- | --- | --- |
|  |  | H1 | H2 | H3 | V1 | V2 | V3 |
| Actuators (1000 counts) | H1 | 2057 | 1297 | 1295 | 2 | 15 | -3 |
|  | H2 | 1285 | 2071 | 1294 | -4 | 13 | -11 |
|  | H3 | 1296 | 1304 | 2082 | -5 | 11 | 5 |
|  | V1 | 184 | 193 | -372 | 1408 | -34 | -620 |
|  | V2 | -372 | 198 | 190 | -630 | 1470 | -47 |
|  | V3 | 179 | -391 | 183 | -44 | -629 | 1477 |
|  | Main coupling |  |  |  |  |  |  |
|  | Important cross coupling |  |  |  |  |  |

Table - Main and cross coupling

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Vertical

For a +1000 count offset drive on vertical actuators

  - Collocated sensors must be 1400 counts +/- 10%

- Horizontal

For a +1000 count offset drive on horizontal actuators

  - Collocated sensors must be 2000 counts +/- 10%

  - Non-collocated horizontal sensors must be 1250 counts +/-10%

Test result:	Passed:   X    	Failed:         .

## Step 15 - Linearity test

|  | Sensors (counts) |
| --- | --- |
|  | V1/V1 | V2/V2 | V3/V3 | H1/H1 | H2/H2 | H3/H3 |
| -7000 | -14747 | -15424 | -15560 | -10245 | -10622 | -11190 |
| -3000 | -6536 | -7160 | -7265 | -4628 | -4832 | -5280 |
| -1000 | -2474 | -3031 | -3125 | -1821 | -1914 | -2330 |
| -300 | -1002 | -1594 | -1670 | -835 | -907 | -1280 |
| 0 | -395 | -985 | -1050 | -416 | -475 | -830 |
| 300 | 223 | -364 | -436 | 4 | -35 | -382 |
| 1000 | 1662 | 1090 | 1018 | 995 | 974 | 632 |
| 3000 | 5771 | 5217 | 5173 | 3794 | 3864 | 3588 |
| 7000 | 13977 | 13463 | 13481 | 9431 | 9674 | 9470 |

Table - Linearity test of the triplet Actuators – HAM-ISI – Sensors

|  | Slope | Offset | Average slope | Variation from average (%) |
| --- | --- | --- | --- | --- |
| H1 | 1.4052 | -413.44 | 1.4437 | -2.66 |
| H2 | 1.4496 | -474.78 |  | 0.41 |
| H3 | 1.4762 | -844.67 |  | 2.25 |
| V1 | 2.0519 | -391.22 | 2.0631 | -0.54 |
| V2 | 2.0632 | -976.44 |  | 0.01 |
| V3 | 2.0741 | -1048.2 |  | 0.53 |

Table - Slopes and offset of the triplet Actuators - HAM-ISI - Sensors

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Linearity_test

- Linearity_test_HAM-ISI_LHO_Unit_2_2010_08_26.mat

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Linearity_test

- Linearity_test_2010_08_26.m

Figures in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Figures/Linearity_test

- Linearity_test.fig

Figure - Horizontal and vertical actuators x HAM-ISI x sensors

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Horizontal and vertical slopes of the triplet actuators x HAM-ISI x sensors:  Average slope +/- 3%.

Test result:	Passed:   X   	Failed:         .

## Step 16 – Static tests in the general coordinate basis

|  | 1000 counts Drive | X Drive | Y Drive | Z Drive | Rx Drive | Ry Drive | Rz Drive |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sensors readout (count) | H1 | 257 | -434 | 11 | -433 | -236 | -1970 |
|  | H2 | 260 | 469 | 15 | 497 | -246 | -1971 |
|  | H3 | -513 | 13 | 17 | 22 | 543 | -1980 |
|  | V1 | -41 | -53 | 191 | -565 | -1710 | -98 |
|  | V2 | -39 | -25 | 230 | 1634 | 334 | -71 |
|  | V3 | -49 | -53 | 222 | -1315 | 1178 | -89 |
|  | Direction read out | 514 | 512 | 236 | 2631 | 2518 | 2547 |

Table - Tests in the general coordinate basis

Issues/difficulties/comments regarding this test:

Acceptance criteria:

For a positive drive in the Cartesian basis:

- Local sensor readout must have the same sign that the reference table (CONT2ACT check)

- Cartesian sensors read out must be positive (DISP2CEN check) in the drive direction

Test result:	Passed:   X   	Failed:         .

## Step 17 - Frequency response

Compensation filters of the new GS13 interface chassis are located in the geophone pre-filters bank. Powerspectra were measured with masses on the optic table not bolted.

## Step 17.1 - Local to local measurements

Local to local transfer functions have been measured with 90 repetitions.

Data files in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_1/Transfer_functions/Local_to_local

- Data_L2L_50mHz_500mHz_20100828-025758.mat

- Data_L2L_500mHz_5Hz_20100827-235225.mat

- Data_L2L_5_200Hz_20100827-221852.mat

- Data_L2L_200_800Hz_20100827-204519.mat

Data collection script files:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Scripts/Data_Collection

- Run_L2L_50mHz_500mHz.m

- Run_L2L_500mHz_5Hz.m

- Run_L2L_5mHz_200Hz.m

- Run_L2L_200Hz_800Hz.m

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Transfer_functions/Local_to_local

- Plot_TF_Local_to_Local_2010_08_27.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAMISI/X2/Data/Unit_1/Figures/Transfer_functions/Local_to_Local

- TF_Horizontal_CPS.fig

- TF_Vertical_CPS.fig

- TF_Horizontal_Geophone.fig

- TF_Vertical_Geophone.fig

Storage of measured transfer functions in the SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_1/Transfer_functions/Local_to_local

- Local_2_Local_LHO_HAM-ISI-Unit_2_2010_08_27.mat

The local to local transfer functions are presented below.

Figure - Local to Local Measurements – Horizontal sensors

Figure - Local to Local Measurements – Vertical sensors

Issues/difficulties/comments regarding this test:  Local to local transfer functions have been measured from the actuator to pre-filter inputs instead of pre-filter outputs. Consequently, we see the GS13 interface chassis response.

## Step 17.2 - Cartesian to Cartesian measurements

Cartesian to Cartesian transfer functions have been measured with 90 repetitions.

Data files in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Transfer_functions/Cartesian_to_cartesian

- Data_Cart2Cart_50mHz_500mHz_20100828-121600.mat

- Data_Cart2Cart_500mHz_5Hz_20100828-091027.mat

- Data_Cart2Cart_5_200Hz_20100828-073654.mat

- Data_Cart2Cart_200_800Hz_20100828-060321.mat

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Transfer_functions/Cartesian_to_cartesian

- Plot_TF_Cartesian_to_Cartesian_2010_08_28.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAMISI/X1/Data/Unit_2/Figures/Transfer_functions/ Cartesian_to_cartesian

- TF_X_Y_RZ_CPS.fig

- TF_X_Y_RZ_Geophone.fig

- TF_Z_RX_RY_CPS.fig

- TF_Z_RX_RY_Geophone.fig

Storage of measured transfer functions in the SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Transfer_functions/Cartesian_to_Cartesian

- Cartesian_2_Cartesian_LHO_HAM-ISI-Unit_2_2010_08_28.mat

Figure - Cartesian to Cartesian measurements – X, Y, RZ directions

Figure - Cartesian to Cartesian measurements – Z, RX, RY directions

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Local to local measurement

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

  - Modal to modal measurement

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

Test result:	Passed:   X   	Failed:         .

## Step 18 - Transfer function comparison

## Step 18.1 - Local to local - Comparison with HAM6

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Transfer_functions/Local_to_local

- Plot_TF_Local_to_Local_2010_08_27.m

Local to local figures in SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Figures/Transfer_functions/Local_to_Local

- TF_Horizontal_CPS_vs_HAM6.fig

- TF_Vertical_CPS_vs_HAM6.fig

- TF_Horizontal_Geophone_vs_HAM6.fig

- TF_Vertical_Geophone_vs_HAM6.fig

GS13, Local to local measurement

Figure - Local to local measurements comparison with HAM6 – GS13

CPS, Local to local measurement, Undamped

Figure - Local to local measurements comparison – Position sensors

Difference with HAM6:

- Sign difference on CPS and Geophones.

## Step 18.2 - Cartesian to Cartesian - Comparison with HAM6

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Transfer_functions/Cartesian_to_cartesian

- Plot_TF_Cartesian_to_Cartesian_2010_08_28.m

Cartesian to Cartesian figures in SVN at :

/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Transfer_functions/Cartesian_to_Cartesian

- TF_X_Y_RZ_CPS_vs_HAM6.fig

- TF_Z_RX_RY_CPS_vs_HAM6.fig

- TF_X_Y_RZ_Geophone_vs_HAM6.fig

- TF_Z_RX_RY_Geophone_vs_HAM6.fig

GS13, Cartesian to Cartesian measurement, Undamped

Figure - Cartesian to Cartesian measurements comparison with HAM6 – GS13

CPS, Cartesian to Cartesian measurement, Undamped

Figure - Cartesian to Cartesian measurements comparison – Position sensors

Difference with HAM6:

- Resonance at 1 Hz is slightly higher than HAM6

- Sign difference on CPS and Geophones

Acceptance criteria:

- No difference with the reference transfer functions (SVN)

  - Phase – less than 10º - In Phase – Out of Phase

  - Damping (fit by eye with the reference transfer functions)

  - DC gain

  - Eigen frequencies shift less than 5%

Test result:	Passed:    X   	Failed:         .

## Step 19 - Lower Zero Moment Plan

Data collection script files:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Scripts/Data_Collection

- Run_Cart2Cart_10mHz_100mHz.m

Data files in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Transfer_functions/LZMP

- Data_LZMP_Cart2Cart_10mHz_100mHz_20100829-095746.mat

Scripts files for processing and plotting in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Transfer_functions/LZMP

- LZMP_2010_08_29.m

Figures in SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Figures/Transfer_functions/LZMP

- LZMP.fig

X & Y offsets:

| X offset (mm) | 0.38 |
| --- | --- |
| Y offset (mm) | 0.11 |

Table – Offset of the Lower Zero Moment Plane

The results from two measurements are presented on the figure below:

Figure - Lower Zero moment plane – Main and cross coupling at low frequency

Acceptance criteria:

- Vertical offset must be less than mm

- Horizontal offset must be less than mm

Test result:	Passed:       	Failed:         .

## Step 20 - Damping loops

Filters used by Damping loops in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/FilterDesign/Unit_2/txt_file

- G1ISIHAM_Back_up_2010_09_01.txt (digitalized filters copied and rename to G1ISIHAM.txt in /opt/rtcds/geo/g1/chans)

## Step 20.1 - Transfer functions - Simulation

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Damping_Loops/Transfer_functions_simulation

- HAM_ISI_LHO_unit_2_Damping_TF_2010_08_26.m

Figures in SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_1/Figures/Damping_Loops

- Damping_LOOP_H1_H2_H3.fig

- Damping_LOOP_V1_V2_V3.fig

Results are saved in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Damping_Loops/Transfer_functions_simulation

- Damping_loop_transfer_functions_HAM-ISI-LHO-unit2_2010_08_26.mat

The following figures present the plant, controller, open loop, closed loop and sensitivity of vertical and horizontal damping loops. H1/V1 are plotted in solid line, H2/V2 are plotted in dash line, H3/V3 are plotted in dash-dot line.

Figure - Vertical damping loops - Simulation

Figure – Horizontal damping loops - Simulation

## Step 20.2 - Powerspectrum – Experimental

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Damping_Loops/
DTT_Powerspectrum_experimental

- Power_spectrum_table_unlocked_Damping_OFF_082510.xml

- Power_spectrum_table_unlocked_Damping_ON_082510.xml

- Power_spectrum_table_unlocked_Damping_OFF_082510.txt

- Power_spectrum_table_unlocked_Damping_ON_082510.txt

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Damping_Loops/
DTT_Powerspectrum_experimental

- HAM_ISI_LHO_unit_2_Damping_Powerspectrum_2010_08_26.m

Figures in SVN at:

/svncommon/seisvn/seismic/HAM-ISI/X1/Data/Unit_2/Figures/Damping_Loops

- Damping_LOOP_GEO_Powerspectrum.fig

- Simulation_vs_experimental_Suppression.fig

Powerspectrum:

The figure below shows power spectrum of Geophones when there when there is no damping loop and when all damping loops are engaged.

Figure – Horizontal (left) and vertical (right) damping loops - Experimental

Sensitivity:

The figure below compare the sensitivity ('Undamped/'Damped') of LLO HAM (Aug 2008) and LLO Unit 1. Performances are very similar, which confirms that we can use the damping loop as they are (modulo electronics change compensation). The plot also shows that the measured performance matches with the prediction.

Figure - Damping loop sensitivity

Acceptance criteria:

- Ham 6 damping loop must implemented and stable with

  - Phase margin must be at least 45º

  - Gain margin must be at least 20dB

Test result:	Passed:   X   	Failed:         .

# Testing Summary