LIGO Laboratory / LIGO Scientific Collaboration

LIGO- T1000468-v2		8/5/2010

Q’s of LASTI Violin Modes

Gregg Harry, Alastair Heptonstall, Matt Evans, Norna Robertson

Distribution of this document:

LIGO Science Collaboration

This is an internal working note

of the LIGO Project.

| California Institute of Technology LIGO Project - MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW 22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

A monolithic quadruple suspension is in the LASTI vacuum chamber and the quality factors of the fused silica suspension fibers have been measured. We compare these measurements to expectations based on theory to help determine if there is any excess loss affecting the suspension fibers.

The theoretical expectation for violin mode Q’s is based on two factors; the theoretical mechanical loss and the dissipation dilution factor.  The theoretical mechanical loss of fused silica can be calculated from Penn et al, Physics Letters A 352 (2006) 3-6.   This paper gives this formula for mechanical loss in silica

= C1 (S/V) + C2 (f / (1 Hz))C3 + C4thermo

where  is the overall mechanical loss, S is the surface area of the fiber, V is the volume of the fiber, f is the frequency of the oscillation, thermo is the thermoelastic loss in the fiber, and C1, C2, C3, and C4 are coefficients that come from fitting mechanical loss data.  The Penn paper give values of C1 = 12.1 10-12 m, C2 = 11.8 10-12, C3 = 0.77, and C4 = 0.61 for Suprasil 2 fused silica.  The distinctions between the different types of fused silica are often lost when the silica is drawn into fibers, so the exact type of silica is likely not important.  The only other type of silica where these parameters are available, Suprasil 312, has C1 and C2 within a factor of 2 of these numbers, C3 identical, and no value available for C4.

Using the Suprasil 2 numbers, the geometry factors of the LASTI fibers allow for an estimate of S and V depend on where the bending in the violin mode takes place.  FEA models (see Cumming et al, Classical and Quantum Gravity 26 (2009) 215012 and Mark Barton) indicate that this bending is in the 800 m diameter section of the fibers. S/V = 4/d for fibers, and thus 5000 1/m.  The frequency of the lowest violin mode is about 500 Hz, well above the 3 Hz thermoelastic peak for an 800 m silica fiber, so thermoelastic damping is ignored from here on.  These parameters all together predict

f0= 7.33X 10-8

f1= 6.83 X 10-8

for the silica fibers at the first two violin modes.

Dissipation dilution for violin modes has been calculated by Gonzalez and Saulson in  Journal of the Acoustical Society of America 96 (1994) 207 as

rn = 2 /l ( 1 + (n)2 /(2 l) )

where rn is the dissipation dilution factor,  is √Y I/T, where Y is the fiber’s Young’s modulus, I is its moment of area, and T is the tension, n is the mode number, and l is the fiber length.  Using 7.27 109 Pa for Y, ¼  (d/2)4 for I, 98 N for T, and 60 cm for l, the dissipation dilution factor for the first two violin modes are

r0 = 1 / 77.2

r1 = 1 / 74.8

The expect violin mode Qs can then be expressed as

Qn = 1 / ( rn fn)

Using the above values for r and , the Q’s become

Q0 = 1.3 109

Q1 = 1.0 109

This theoretical value can be compared to the experimentally value on the LASTI suspension measured to be

Q0 = 5.9 108

Q0 = 6.1 108

Q1 = 4.6 108

as shown in the LASTI ILOG July 31, 2010.

The use of a single thickness for the fiber in both the silica mechanical loss (where it enters through S/V) and the dissipation dilution factor (where it enters through I) is simplistic, the actual fibers have a very large diameter stock region of d = 3000 m, a thick region of d = 800 m, and a long thin middle region of d = 400 m.  A true estimate of the mechanical loss and dissipation dilution, and thus the violin mode Q’s, requires finite element modeling, which is in progress by Alan Cummings et al.