LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T1000504-v1	LIGO	25/8/10

Test Procedure for RF Frequency Doubler

Paul Schwinberg and Daniel Sigg

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

The following Test Procedure describes the test of proper operation of the RF Frequency Doubler.

# Test Equipment

- Voltmeter

- Oscilloscope

- Stanford Research SR785 analyzer

- Tektronix AFG3101 function generator (or similar)

- RF Power Meter HP E4418A

- Board Schematics--Frequency Doubler

# Tests

The RF Frequency Doubler comes with a number of different power supply boards so I will assume that we are using the latest which is the Low Noise Power Module (D0901846) with the RF Distribution  Amplifier :Interface (D1000064).

- Verify the proper current draw.  Using a bench DC supply apply ±24Volts to P7 and 
±17 Volts to P6 of the low noise power Module (D0901846).  Measure the current draw of the board.

+24 Volt current ______________     0.1 A Nom.

–24 Volt current  ______________     0.0 A Nom.

+17 Volt current ______________    less than 1.0 A

–17  Volt current ______________    less than 0.01 A

- On the low noise power module check the voltage on TP 1-13.

TP1 ( +17V ) ______________		TP2 (-17V ) _______________

TP3 , 4 ( GND )				TP5 (+ 5V)________________

TP6 (-15V ) ________________		TP7 (+24V ) ________________

TP8 ( GND )				TP9 ( -24V ) ____________________

TP10 ( GND )				TP11 (+15V ) ____________________

TP12 (+VREF ) _____________		TP13  (-VREF ) ____________________

- If  TP 1 , 2, 7 , 9 and 8 are correct then pin 5 on U1 and U7,  TP14 (OK) should be

Logic high ~3Volts. The front panel LED should be on.

Confirm._____________________

- The noise on TP 12, 13, 11 and 6 should be measured with a SR785 using an

rms power spectrum.

TP12  noise ____________________less than 200 nVrms/sqrt Hz at 140 Hz

TP13 noise _____________________less than 900 nVrms/sqrt Hz at 140 Hz

TP11 noise _____________________ less than 1 uVrms/sqrt Hz at 140 Hz

TP6 noise ______________________ less than 1 uVrms/sqrt Hz at 140 Hz.

- Test the RF monitor by applying a 160 MHz RF signal to J1.  Measure the output voltage at mon1.

Mon1

| Nom output pwr | Input pwr dBm | Mon volt (M) | Measured volt. | Measured Pwr |
| --- | --- | --- | --- | --- |
| 13 dBm |  | 2.9V (0.725) |  |  |
| 10 dBm |  | 3.2V (0.800) |  |  |
| 7 dBm |  | 3.5V (0.875) |  |  |
| 0 dBm |  | 4.2V (1.05) |  |  |
| -10 dBm |  | 5.2V (1.30) |  |  |
| -20 dBm |  | 6.2V (1.55) |  |  |

6)  Test the RF output powers by applying a 80 MHz/10dBm RF signal to J1. With a RF power meter measure the power at the output (13 dBm nominal). If the output power is consistently too high an attenuator A1 has to be adjusted accordingly. Nominal output power is 13 dBm.

Output: ____________________ (13 dBm nominal)

7)  Measure  the Phase noise of the RF Oscillator Source driving the RF Frequency Doubler , with   1PPS locking using the Wenzel single channel phase noise measurement technique (3.5.3), Figure 3.5.2-1, which can be found at

http://www.wenzel.com/pdffiles1/BP1000Manual/BP_1000_v101_2_.pdf .

A reasonable FFT analyzer is the SR785, which can be set to measure power units if you start in Display Setup.  A Reference Source must be provided which can be just a Wenzel crystal oscillator of frequency close enough to lock, properly powered and connected to the Wenzel phase noise measurement system. The output of the RF Frequency Doubler will need to be attenuated to the amplitude needed by the wenzel phase noise measurement system (about 10 dBm ). Compare to the Phase noise of the RF Oscillator Source alone, it should be within 3dB.

J2

| Offset freq. Hz | Phase noise spec. | RF osc. phase noise | RF osc + doubler |
| --- | --- | --- | --- |
| 10 Hz | -110 dBc/Hz |  |  |
| 100 Hz | -140 dBc/Hz |  |  |
| 1 kHz | -160 dBc/Hz |  |  |