*LIGO Laboratory / LIGO Scientific Collaboration*

LIGO-T1000504-v3 *LIGO* 2/8/2018

Test Procedure for RF Frequency Doubler

Paul Schwinberg, Daniel Sigg, Marc Pirello

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

|  |  |
| --- | --- |
| **California Institute of Technology****LIGO Project – MS 18-34****1200 E. California Blvd.****Pasadena, CA 91125**Phone (626) 395-2129Fax (626) 304-9834E-mail: info@ligo.caltech.edu | **Massachusetts Institute of Technology****LIGO Project – NW22-295****185 Albany St****Cambridge, MA 02139**Phone (617) 253-4824Fax (617) 253-7014E-mail: info@ligo.mit.edu |
| **LIGO Hanford Observatory****P.O. Box 159****Richland WA 99352**Phone 509-372-8106Fax 509-372-8137 | **LIGO Livingston Observatory****P.O. Box 940****Livingston, LA 70754**Phone 225-686-3100Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

The following Test Procedure describes the test of proper operation of the RF Frequency Doubler.

# Test Equipment

* Voltmeter
* Oscilloscope
* Stanford Research SR785 analyzer
* Tektronix AFG3101 function generator (or similar)
* RF Power Meter HP E4418A
* Board Schematics—[Frequency Doubler](https://awiki.ligo-wa.caltech.edu/aLIGO/FrequencyDoubler)

# Tests

*The RF Frequency Doubler comes with a number of different power supply boards so I will assume that we are using the latest which is the Low Noise Power Module (D0901846) with the RF Distribution Amplifier :Interface (D1000064).*

*Determine doubler base frequency and record here: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_*

1. **Verify the proper current draw.**  Using a bench DC supply apply ±24Volts to P7 and
±17 Volts to P6 of the low noise power Module (D0901846). Measure the current draw of the board.

+24 Volt current \_\_\_\_\_\_\_\_\_\_\_\_\_\_ 30mA Nom.

–24 Volt current \_\_\_\_\_\_\_\_\_\_\_\_\_\_ 30mA Nom.

+17 Volt current \_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 500mA

–17 Volt current \_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 50mA

1. **On the low noise power module check the voltage on TP 1-13.**

TP1 ( +17V ) \_\_\_\_\_\_\_\_\_\_\_\_\_\_ TP2 (–17V ) \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

TP3 , 4 ( GND ) TP5 (+ 5V)\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

TP6 (–15V ) \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ TP7 (+24V ) \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

TP8 ( GND ) TP9 (–24V ) \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

TP10 ( GND ) TP11 (+15V ) \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

TP12 (+VREF ) \_\_\_\_\_\_\_\_\_\_\_\_\_ TP13 (–VREF ) \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

1. **If TP 1 , 2, 7 , 9 and 8 are correct then pin 5 on U1 and U7, TP14 (OK) should be**

**Logic high ~3Volts. The front panel LED should be on.**

**Confirm.\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

1. **The noise on TP 12, 13, 11 and 6 should be measured with a SR785 using an**

**rms power spectrum.**

TP12 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_less than 20 nVrms/√Hz at 140 Hz

TP13 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_less than 30 nVrms/√Hz at 140 Hz

TP11 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 40 nVrms/√Hz at 140 Hz

TP6 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 60 nVrms/√Hz at 140 Hz.

1. **Test the RF monitor by applying the base frequency directly to the power meter at M1.**  Monitor the nominal output power at the DB25 and M1 on the back of the unit.

|  |  |  |  |
| --- | --- | --- | --- |
| **Nom output pwr** |  **Input PWR**  | **M1 Volts** | **DB25 Volts** |
| **10 dBm** |  | **(0.55V)** | **(2.2V)** |
| **0 dBm** |  | **(0.8V)** | **(3.2V)** |
| **-10 dBm** |  | **(1.05V)** | **(4.2V)** |
| **-20 dBm** |  | **(1.3V)** | **(5.2V)** |
| **-30 dBm** |  | **(1.55V)** | **(6.2V)** |
| **none** |  | **(2.15V)** | **(8.6V)** |

**6) Test the RF output powers by applying the base frequency to the input port.** With a RF power meter measure the power at the output (13+/-2 dBm nominal). If the output power is consistently too high, an attenuator A1 has to be adjusted accordingly. Nominal output power is 13+/-2 dBm.

Output: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ (13+/-2 dBm nominal)

**7) Measure the phase noise of an OCXO driving the RF Frequency Doubler at close to the base frequency.** Use a second OCXO which is 2x the first OCXO frequency as the second oscillator to compare the output signal of the doubler, using the Wenzel single channel phase noise measurement technique (3.5.3), Figure 3.5.2-1, which can be found at:

<http://www.wenzel.com/pdffiles1/BP1000Manual/BP_1000_v101_2_.pdf> .

A reasonable FFT analyzer is the SR785, which can be set to measure power units if you start in Display Setup. A Reference Source must be provided which can be just a Wenzel crystal oscillator of frequency close enough to lock, properly powered and connected to the Wenzel phase noise measurement system. The output of the RF Frequency Doubler will need to be attenuated to the amplitude needed by the Wenzel phase noise measurement system (about 10 dBm). Compare to the phase noise of the OCXO datasheet, add 6dB to the noise of the base unit and add it in quadrature to the noise of the 160MHz unit. The noise of the doubler should be within 3dB.

|  |  |  |
| --- | --- | --- |
| **Offset (Hz)** |  **Phase noise spec (dB/Hz)** | **Measured (dB/Hz)** |
| **Base MHz** | **2xBase MHz** | **total** |
|  **10** |  **–90**  | **–84** | **–81** |  |
|  **100** | **–110** | **–104** | **–101** |  |
|  **1000** | **–140** | **–134** | **–131** |  |
|  **10000** | **–160** | **–154** | **–151** |  |