## Outgassing Rates
A-OSEM Head with Wound Coil (404 units)
Species | AMU | Outgassing Rate (Torr-L/s per AOSEM body)
H | 2 | 3.8675e-13
He | 4 | 3.2925e-14
C+ | 12 | 1.0702e-13
N+, N2+ | 14 | 1.9111e-13
water (O+) | 16 | 3.6101e-13
water (HO+) | 17 | 1.103e-12
water (H2O+) | 18 | 1.4519e-12
CO2++ | 22 | 2.3526e-14
N2 | 28 | 1.799e-11
O2 | 32 | 1.6121e-12
Ar | 40 | 1.5164e-13
hydrocarbon | 41 | 2.0198e-13
hydrocarbon | 43 | 1.728e-13
CO2 | 44 | 3.6772e-12
hydrocarbon | 53 | 1.1027e-14
hydrocarbon | 55 | 9.7112e-15
hydrocarbon | 57 | 1.6404e-14
Sum HC |  | 4.1192e-13
| AMU |  | Outgassing Rate (Torr-L/s) per AOSEM Wound Head 404 units
hydrogen (H2) | 2 |  | 3.8675e-13
water (H2O) at 1000 hr | max of 16, 17, 18 |  | 1.4519e-12
air (N2) | 28 |  | 1.799e-11
CO2 | 44 |  | 3.6772e-12
hydrocarbons | sum of 41, 43, 53, 55, 57 |  | 4.1192e-13

## C83010 aLIGO AOSEM Flex Circuit
2010-08-30 | 2010 07:52:41 AM
Residual Gas Analyzer Software
Software Version | 3.205.003
Analog Scan Setup:
Data Points in Scan | 991
Units | Amps
Noise Floor | 3
CEM Status | OFF
Points Per AMU | 10
Scan Start Mass | 1 | amu
Scan Stop Mass | 100 | amu
Focus Voltage | 90 | Volts
Ion Energy | HIGH
Electron Energy | 70 | eV
CEM Voltage | 1000 | Volts
CEM Gain | 1000000
Sensitivity Factor | 0.0001
Filament Current | 1 | mAmps
-4.8076e-16
1 | 3.32e-15 | 3.8008e-15
1.1 | 3.23e-15 | 3.7108e-15
1.2 | 1.23e-14 | 1.2781e-14
1.3 | 3.6e-13 | 3.6048e-13
1.4 | 4.77e-12 | 4.7705e-12
1.5 | 6.38e-12 | 6.3805e-12
1.6 | 4.67e-12 | 4.6705e-12
1.7 | 3.05e-12 | 3.0505e-12
1.8 | 7.92e-13 | 7.9248e-13
1.9 | 3.93e-14 | 3.9781e-14
2 | 6.28e-15 | 6.7608e-15
2.1 | 2.68e-15 | 3.1608e-15
2.2 | 1.37e-15 | 1.8508e-15
2.3 | 7.4e-16 | 1.2208e-15
2.4 | 5.76e-16 | 1.0568e-15
2.5 | 1.37e-15 | 1.8508e-15
2.6 | 4.83e-17 | 5.2906e-16
2.7 | -1.76e-16 | 3.0476e-16
2.8 | -2.23e-16 | 2.5776e-16
2.9 | -2.72e-16 | 2.0876e-16
3 | -3.57e-16 | 1.2376e-16
3.1 | 3.83e-16 | 8.6376e-16
3.2 | 3.93e-15 | 4.4108e-15
3.3 | 4.43e-15 | 4.9108e-15
3.4 | 4.16e-15 | 4.6408e-15
3.5 | 2.43e-15 | 2.9108e-15
3.6 | 3.11e-15 | 3.5908e-15
3.7 | 2.28e-15 | 2.7608e-15
3.8 | 2.93e-15 | 3.4108e-15
3.9 | 5.72e-16 | 1.0528e-15
4 | 9.48e-17 | 5.7556e-16
4.1 | -2.27e-16 | 2.5376e-16
4.2 | -3.85e-16 | 9.576e-17
4.3 | -3.78e-16 | 1.0276e-16
4.4 | -2.68e-16 | 2.1276e-16
4.5 | -3.76e-16 | 1.0476e-16
4.6 | -4.27e-16 | 5.376e-17
4.7 | -3.82e-16 | 9.876e-17
4.8 | -3.5e-16 | 1.3076e-16
4.9 | -4.22e-16 | 5.876e-17
5 | -3.06e-16 | 1.7476e-16
5.1 | 1.18e-15 | 1.6608e-15
5.2 | 4.05e-16 | 8.8576e-16
5.3 | -7.26e-17 | 4.0816e-16
5.4 | 2.34e-17 | 5.0416e-16
5.5 | -3.01e-16 | 1.7976e-16
5.6 | -3.3e-16 | 1.5076e-16
5.7 | -3.38e-16 | 1.4276e-16
5.8 | -2.97e-16 | 1.8376e-16
5.9 | -3.23e-16 | 1.5776e-16
6 | -2.86e-16 | 1.9476e-16
6.1 | 1.51e-15 | 1.9908e-15
6.2 | 5.39e-17 | 5.3466e-16
6.3 | -2.94e-16 | 1.8676e-16
6.4 | -3.19e-16 | 1.6176e-16
6.5 | 1.24e-16 | 6.0476e-16
6.6 | 3.93e-17 | 5.2006e-16
6.7 | -2.54e-16 | 2.2676e-16
6.8 | -2.83e-16 | 1.9776e-16
6.9 | -2.88e-16 | 1.9276e-16
7 | -3.25e-16 | 1.5576e-16
7.1 | 3.03e-16 | 7.8376e-16
7.2 | -3.02e-17 | 4.5056e-16
7.3 | -2.39e-16 | 2.4176e-16
7.4 | -1.11e-16 | 3.6976e-16
7.5 | -3.04e-16 | 1.7676e-16
7.6 | -3.05e-16 | 1.7576e-16
7.7 | -2.45e-16 | 2.3576e-16
7.8 | -2.46e-16 | 2.3476e-16
7.9 | -3.31e-16 | 1.4976e-16
8 | -3.79e-16 | 1.0176e-16
8.1 | 2.29e-15 | 2.7708e-15
8.2 | 5.75e-16 | 1.0558e-15
8.3 | 3.03e-16 | 7.8376e-16
8.4 | -6.38e-17 | 4.1696e-16
8.5 | 3.31e-15 | 3.7908e-15
8.6 | 1.91e-16 | 6.7176e-16
8.7 | -1.14e-16 | 3.6676e-16
8.8 | -2.48e-16 | 2.3276e-16
8.9 | -1.08e-16 | 3.7276e-16
9 | -3.38e-16 | 1.4276e-16
9.1 | -2.8e-16 | 2.0076e-16
9.2 | 1.56e-15 | 2.0408e-15
9.3 | 7e-16 | 1.1808e-15
9.4 | -6.22e-17 | 4.1856e-16
9.5 | -2.8e-16 | 2.0076e-16
9.6 | -3.36e-16 | 1.4476e-16
9.7 | -2.83e-16 | 1.9776e-16
9.8 | -3.28e-16 | 1.5276e-16
9.9 | -2.15e-16 | 2.6576e-16
10 | -3.4e-16 | 1.4076e-16
10.1 | -3.19e-16 | 1.6176e-16
10.2 | -3.26e-16 | 1.5476e-16
10.3 | -2.92e-16 | 1.8876e-16
10.4 | -2.65e-16 | 2.1576e-16
10.5 | -2.78e-16 | 2.0276e-16
10.6 | -3.39e-16 | 1.4176e-16
10.7 | -3.26e-16 | 1.5476e-16
10.8 | -1.22e-17 | 4.6856e-16
10.9 | 4.59e-16 | 9.3976e-16
11 | -2.59e-17 | 4.5486e-16
11.1 | 3.86e-16 | 8.6676e-16
11.2 | 5.7e-15 | 6.1808e-15
11.3 | 1.21e-14 | 1.2581e-14
11.4 | 1.79e-14 | 1.8381e-14
11.5 | 2.41e-14 | 2.4581e-14
11.6 | 1.99e-14 | 2.0381e-14
11.7 | 1.5e-14 | 1.5481e-14
11.8 | 8.56e-15 | 9.0408e-15
11.9 | 3.99e-15 | 4.4708e-15
12 | 1.39e-15 | 1.8708e-15
12.1 | 4.79e-16 | 9.5976e-16
12.2 | 5.14e-16 | 9.9476e-16
12.3 | 1.45e-15 | 1.9308e-15
12.4 | 1.04e-15 | 1.5208e-15
12.5 | 3.22e-15 | 3.7008e-15
12.6 | 3.05e-15 | 3.5308e-15
12.7 | 5.23e-15 | 5.7108e-15
12.8 | 2.81e-15 | 3.2908e-15
12.9 | 3.34e-15 | 3.8208e-15
13 | 1.16e-15 | 1.6408e-15
13.1 | 1.34e-15 | 1.8208e-15
13.2 | 5.88e-15 | 6.3608e-15
13.3 | 1.76e-14 | 1.8081e-14
13.4 | 3.95e-14 | 3.9981e-14
13.5 | 3.64e-14 | 3.6881e-14
13.6 | 3.31e-14 | 3.3581e-14
13.7 | 3.45e-14 | 3.4981e-14
13.8 | 3.02e-14 | 3.0681e-14
13.9 | 1.05e-14 | 1.0981e-14
14 | 2.86e-15 | 3.3408e-15
14.1 | 1.24e-15 | 1.7208e-15
14.2 | 8.69e-16 | 1.3498e-15
14.3 | 1.43e-15 | 1.9108e-15
14.4 | 8.43e-15 | 8.9108e-15
14.5 | 1e-14 | 1.0481e-14
14.6 | 1.34e-14 | 1.3881e-14
14.7 | 1.2e-14 | 1.2481e-14
14.8 | 1.05e-14 | 1.0981e-14
14.9 | 5.89e-15 | 6.3708e-15
15 | 2.83e-15 | 3.3108e-15
15.1 | 2.76e-15 | 3.2408e-15
15.2 | 3.91e-15 | 4.3908e-15
15.3 | 1.57e-14 | 1.6181e-14
15.4 | 2.45e-14 | 2.4981e-14
15.5 | 4.25e-14 | 4.2981e-14
15.6 | 3.25e-14 | 3.2981e-14
15.7 | 3.28e-14 | 3.3281e-14
15.8 | 2.56e-14 | 2.6081e-14
15.9 | 1.86e-14 | 1.9081e-14
16 | 5.83e-15 | 6.3108e-15
16.1 | 1.23e-14 | 1.2781e-14
16.2 | 2.23e-14 | 2.2781e-14
16.3 | 7.55e-14 | 7.5981e-14
16.4 | 1.76e-13 | 1.7648e-13
16.5 | 1.66e-13 | 1.6648e-13
16.6 | 1.58e-13 | 1.5848e-13
16.7 | 1.5e-13 | 1.5048e-13
16.8 | 1.06e-13 | 1.0648e-13
16.9 | 6.05e-14 | 6.0981e-14
17 | 1.88e-14 | 1.9281e-14
17.1 | 2.09e-14 | 2.1381e-14
17.2 | 8.25e-14 | 8.2981e-14
17.3 | 2.53e-13 | 2.5348e-13
17.4 | 5.7e-13 | 5.7048e-13
17.5 | 6.52e-13 | 6.5248e-13
17.6 | 7.08e-13 | 7.0848e-13
17.7 | 5.78e-13 | 5.7848e-13
17.8 | 4.63e-13 | 4.6348e-13
17.9 | 2.69e-13 | 2.6948e-13
18 | 2.49e-14 | 2.5381e-14
18.1 | 4.84e-15 | 5.3208e-15
18.2 | 1.66e-15 | 2.1408e-15
18.3 | 6.16e-16 | 1.0968e-15
18.4 | 9.34e-16 | 1.4148e-15
18.5 | 3.23e-15 | 3.7108e-15
18.6 | 3.22e-15 | 3.7008e-15
18.7 | 2.6e-15 | 3.0808e-15
18.8 | 3.39e-15 | 3.8708e-15
18.9 | 4.41e-15 | 4.8908e-15
19 | 1.94e-15 | 2.4208e-15
19.1 | 5.5e-16 | 1.0308e-15
19.2 | 2.89e-16 | 7.6976e-16
19.3 | 1.28e-16 | 6.0876e-16
19.4 | 1.18e-15 | 1.6608e-15
19.5 | 2.66e-15 | 3.1408e-15
19.6 | 4.52e-15 | 5.0008e-15
19.7 | 7.33e-15 | 7.8108e-15
19.8 | 6.76e-15 | 7.2408e-15
19.9 | 3.61e-15 | 4.0908e-15
20 | 1.51e-15 | 1.9908e-15
20.1 | 7.34e-16 | 1.2148e-15
20.2 | 1.1e-16 | 5.9076e-16
20.3 | -2.51e-16 | 2.2976e-16
20.4 | -3.16e-16 | 1.6476e-16
20.5 | -3.12e-16 | 1.6876e-16
20.6 | -2.26e-16 | 2.5476e-16
20.7 | -2.68e-16 | 2.1276e-16
20.8 | -3.34e-16 | 1.4676e-16
20.9 | -3.61e-16 | 1.1976e-16
21 | -4.03e-16 | 7.776e-17
21.1 | -6.67e-17 | 4.1406e-16
21.2 | -1.68e-16 | 3.1276e-16
21.3 | -2.33e-16 | 2.4776e-16
21.4 | -3.64e-17 | 4.4436e-16
21.5 | 1.26e-16 | 6.0676e-16
21.6 | 1.16e-16 | 5.9676e-16
21.7 | 1.76e-16 | 6.5676e-16
21.8 | 8.97e-17 | 5.7046e-16
21.9 | 9.24e-17 | 5.7316e-16
22 | -6.95e-17 | 4.1126e-16
22.1 | -2.42e-16 | 2.3876e-16
22.2 | -3.7e-16 | 1.1076e-16
22.3 | -2e-16 | 2.8076e-16
22.4 | 2.09e-15 | 2.5708e-15
22.5 | 1.41e-15 | 1.8908e-15
22.6 | 1.49e-16 | 6.2976e-16
22.7 | -8.61e-17 | 3.9466e-16
22.8 | -2.89e-16 | 1.9176e-16
22.9 | -4.27e-16 | 5.376e-17
23 | -3.36e-16 | 1.4476e-16
23.1 | -2.61e-16 | 2.1976e-16
23.2 | -2.99e-16 | 1.8176e-16
23.3 | -3.65e-16 | 1.1576e-16
23.4 | 1.22e-15 | 1.7008e-15
23.5 | 6.08e-16 | 1.0888e-15
23.6 | 3.21e-16 | 8.0176e-16
23.7 | 1.13e-16 | 5.9376e-16
23.8 | -1.93e-17 | 4.6146e-16
23.9 | 1.22e-16 | 6.0276e-16
24 | 1.41e-16 | 6.2176e-16
24.1 | -1.47e-16 | 3.3376e-16
24.2 | -1.67e-16 | 3.1376e-16
24.3 | -2.02e-16 | 2.7876e-16
24.4 | 2.71e-15 | 3.1908e-15
24.5 | 1.94e-15 | 2.4208e-15
24.6 | 1.51e-15 | 1.9908e-15
24.7 | 1.11e-15 | 1.5908e-15
24.8 | 9.43e-16 | 1.4238e-15
24.9 | 8.82e-16 | 1.3628e-15
25 | 1.86e-15 | 2.3408e-15
25.1 | 8.97e-16 | 1.3778e-15
25.2 | 3.51e-15 | 3.9908e-15
25.3 | 1.39e-15 | 1.8708e-15
25.4 | 3.12e-15 | 3.6008e-15
25.5 | 3.22e-15 | 3.7008e-15
25.6 | 9.61e-15 | 1.0091e-14
25.7 | 9.82e-15 | 1.0301e-14
25.8 | 1.2e-14 | 1.2481e-14
25.9 | 9.26e-15 | 9.7408e-15
26 | 6.39e-15 | 6.8708e-15
26.1 | 3.21e-15 | 3.6908e-15
26.2 | 1.41e-15 | 1.8908e-15
26.3 | 2.18e-15 | 2.6608e-15
26.4 | 1.5e-15 | 1.9808e-15
26.5 | 4.39e-15 | 4.8708e-15
26.6 | 1.22e-14 | 1.2681e-14
26.7 | 1.68e-14 | 1.7281e-14
26.8 | 1.66e-14 | 1.7081e-14
26.9 | 2.29e-14 | 2.3381e-14
27 | 1.62e-14 | 1.6681e-14
27.1 | 1.6e-14 | 1.6481e-14
27.2 | 1.86e-14 | 1.9081e-14
27.3 | 3.81e-14 | 3.8581e-14
27.4 | 6.25e-14 | 6.2981e-14
27.5 | 1.99e-13 | 1.9948e-13
27.6 | 4.27e-13 | 4.2748e-13
27.7 | 5.26e-13 | 5.2648e-13
27.8 | 4.87e-13 | 4.8748e-13
27.9 | 4.47e-13 | 4.4748e-13
28 | 3.14e-13 | 3.1448e-13
28.1 | 5.8e-14 | 5.8481e-14
28.2 | 8.34e-15 | 8.8208e-15
28.3 | 2.12e-15 | 2.6008e-15
28.4 | 7.8e-16 | 1.2608e-15
28.5 | 1.51e-15 | 1.9908e-15
28.6 | 7.27e-15 | 7.7508e-15
28.7 | 1.92e-14 | 1.9681e-14
28.8 | 1.83e-14 | 1.8781e-14
28.9 | 2.26e-14 | 2.3081e-14
29 | 1.48e-14 | 1.5281e-14
29.1 | 9.86e-15 | 1.0341e-14
29.2 | 2.68e-15 | 3.1608e-15
29.3 | 4.2e-15 | 4.6808e-15
29.4 | 1.08e-15 | 1.5608e-15
29.5 | 1.05e-15 | 1.5308e-15
29.6 | 5.74e-15 | 6.2208e-15
29.7 | 7.67e-15 | 8.1508e-15
29.8 | 9.83e-15 | 1.0311e-14
29.9 | 9.77e-15 | 1.0251e-14
30 | 1.25e-14 | 1.2981e-14
30.1 | 1.01e-14 | 1.0581e-14
30.2 | 2.4e-15 | 2.8808e-15
30.3 | 7.05e-16 | 1.1858e-15
30.4 | 7.55e-17 | 5.5626e-16
30.5 | 8.25e-17 | 5.6326e-16
30.6 | 2.02e-16 | 6.8276e-16
30.7 | 4.3e-16 | 9.1076e-16
30.8 | 7.46e-16 | 1.2268e-15
30.9 | 1.04e-15 | 1.5208e-15
31 | 2.11e-15 | 2.5908e-15
31.1 | 1.38e-15 | 1.8608e-15
31.2 | 8.29e-16 | 1.3098e-15
31.3 | 9.66e-16 | 1.4468e-15
31.4 | 8e-16 | 1.2808e-15
31.5 | 3.3e-15 | 3.7808e-15
31.6 | 9.67e-15 | 1.0151e-14
31.7 | 2.35e-14 | 2.3981e-14
31.8 | 3.92e-14 | 3.9681e-14
31.9 | 3.73e-14 | 3.7781e-14
32 | 2.77e-14 | 2.8181e-14
32.1 | 1.56e-14 | 1.6081e-14
32.2 | 3.93e-15 | 4.4108e-15
32.3 | 1.5e-15 | 1.9808e-15
32.4 | 4.37e-16 | 9.1776e-16
32.5 | -4.14e-17 | 4.3936e-16
32.6 | 2.49e-16 | 7.2976e-16
32.7 | 5.78e-16 | 1.0588e-15
32.8 | 4.83e-15 | 5.3108e-15
32.9 | 9.4e-16 | 1.4208e-15
33 | 8.48e-16 | 1.3288e-15
33.1 | 8.01e-16 | 1.2818e-15
33.2 | 2.22e-16 | 7.0276e-16
33.3 | -1.19e-16 | 3.6176e-16
33.4 | 2.4e-16 | 7.2076e-16
33.5 | 3.39e-16 | 8.1976e-16
33.6 | 3.05e-16 | 7.8576e-16
33.7 | 9.21e-16 | 1.4018e-15
33.8 | 2.4e-15 | 2.8808e-15
33.9 | 4.81e-15 | 5.2908e-15
34 | 3.69e-15 | 4.1708e-15
34.1 | 2.18e-15 | 2.6608e-15
34.2 | 8.53e-16 | 1.3338e-15
34.3 | 2.47e-16 | 7.2776e-16
34.4 | 2.03e-16 | 6.8376e-16
34.5 | -1.35e-16 | 3.4576e-16
34.6 | -5.3e-17 | 4.2776e-16
34.7 | 3.22e-16 | 8.0276e-16
34.8 | 5.98e-16 | 1.0788e-15
34.9 | 8.51e-16 | 1.3318e-15
35 | 1.49e-15 | 1.9708e-15
35.1 | 7.97e-16 | 1.2778e-15
35.2 | 4.73e-16 | 9.5376e-16
35.3 | 4.56e-16 | 9.3676e-16
35.4 | 5.1e-17 | 5.3176e-16
35.5 | -2.68e-17 | 4.5396e-16
35.6 | 1.67e-15 | 2.1508e-15
35.7 | 4.39e-15 | 4.8708e-15
35.8 | 7.05e-15 | 7.5308e-15
35.9 | 8.39e-15 | 8.8708e-15
36 | 9.75e-15 | 1.0231e-14
36.1 | 9.64e-15 | 1.0121e-14
36.2 | 3.83e-15 | 4.3108e-15
36.3 | 1.21e-15 | 1.6908e-15
36.4 | 3.49e-16 | 8.2976e-16
36.5 | 2.06e-17 | 5.0136e-16
36.6 | 1.52e-17 | 4.9596e-16
36.7 | 1.13e-16 | 5.9376e-16
36.8 | 2.71e-16 | 7.5176e-16
36.9 | 4.87e-16 | 9.6776e-16
37 | 5.4e-16 | 1.0208e-15
37.1 | 2.6e-16 | 7.4076e-16
37.2 | 5.35e-17 | 5.3426e-16
37.3 | -7.73e-17 | 4.0346e-16
37.4 | -8.9e-17 | 3.9176e-16
37.5 | 6.3e-17 | 5.4376e-16
37.6 | 3.61e-16 | 8.4176e-16
37.7 | 1.08e-15 | 1.5608e-15
37.8 | 1.25e-15 | 1.7308e-15
37.9 | 2.31e-15 | 2.7908e-15
38 | 2.85e-15 | 3.3308e-15
38.1 | 6.17e-15 | 6.6508e-15
38.2 | 3.56e-15 | 4.0408e-15
38.3 | 1.18e-15 | 1.6608e-15
38.4 | 5.84e-16 | 1.0648e-15
38.5 | 2.23e-16 | 7.0376e-16
38.6 | 4.15e-16 | 8.9576e-16
38.7 | 4.06e-16 | 8.8676e-16
38.8 | 5.26e-16 | 1.0068e-15
38.9 | 1.04e-15 | 1.5208e-15
39 | 1.65e-15 | 2.1308e-15
39.1 | 2.7e-15 | 3.1808e-15
39.2 | 3.61e-15 | 4.0908e-15
39.3 | 1.63e-15 | 2.1108e-15
39.4 | 2.56e-15 | 3.0408e-15
39.5 | 1.02e-15 | 1.5008e-15
39.6 | 2.99e-16 | 7.7976e-16
39.7 | 2.28e-15 | 2.7608e-15
39.8 | 1.98e-15 | 2.4608e-15
39.9 | 2.77e-15 | 3.2508e-15
40 | 2.17e-15 | 2.6508e-15
40.1 | 2.06e-15 | 2.5408e-15
40.2 | 1.49e-15 | 1.9708e-15
40.3 | 8.31e-16 | 1.3118e-15
40.4 | 3.04e-16 | 7.8476e-16
40.5 | 1.87e-16 | 6.6776e-16
40.6 | 5.05e-16 | 9.8576e-16
40.7 | 5.94e-16 | 1.0748e-15
40.8 | 4.61e-15 | 5.0908e-15
40.9 | 1.56e-15 | 2.0408e-15
41 | 3.05e-15 | 3.5308e-15
41.1 | 9.8e-15 | 1.0281e-14
41.2 | 6.87e-15 | 7.3508e-15
41.3 | 3.56e-15 | 4.0408e-15
41.4 | 1.24e-15 | 1.7208e-15
41.5 | 4.92e-16 | 9.7276e-16
41.6 | 1.52e-16 | 6.3276e-16
41.7 | 4.63e-16 | 9.4376e-16
41.8 | 9e-16 | 1.3808e-15
41.9 | 3.69e-15 | 4.1708e-15
42 | 6.63e-15 | 7.1108e-15
42.1 | 6.43e-15 | 6.9108e-15
42.2 | 4.05e-15 | 4.5308e-15
42.3 | 2.87e-15 | 3.3508e-15
42.4 | 1.18e-15 | 1.6608e-15
42.5 | 3.55e-16 | 8.3576e-16
42.6 | 1.57e-15 | 2.0508e-15
42.7 | 1.45e-15 | 1.9308e-15
42.8 | 5.41e-16 | 1.0218e-15
42.9 | 5.79e-15 | 6.2708e-15
43 | 2.54e-15 | 3.0208e-15
43.1 | 5.39e-15 | 5.8708e-15
43.2 | 3.59e-15 | 4.0708e-15
43.3 | 4.02e-15 | 4.5008e-15
43.4 | 6.05e-15 | 6.5308e-15
43.5 | 3.98e-15 | 4.4608e-15
43.6 | 4.72e-15 | 5.2008e-15
43.7 | 9.8e-15 | 1.0281e-14
43.8 | 1.3e-14 | 1.3481e-14
43.9 | 4.65e-14 | 4.6981e-14
44 | 6.38e-14 | 6.4281e-14
44.1 | 8.35e-14 | 8.3981e-14
44.2 | 5.9e-14 | 5.9481e-14
44.3 | 3.11e-14 | 3.1581e-14
44.4 | 6.6e-15 | 7.0808e-15
44.5 | 1.83e-15 | 2.3108e-15
44.6 | 6.07e-16 | 1.0878e-15
44.7 | 1.26e-16 | 6.0676e-16
44.8 | -3.89e-19 | 4.8037e-16
44.9 | 1.46e-16 | 6.2676e-16
45 | 9.46e-16 | 1.4268e-15
45.1 | 7.64e-16 | 1.2448e-15
45.2 | 9.33e-16 | 1.4138e-15
45.3 | 1.01e-15 | 1.4908e-15
45.4 | 3.26e-16 | 8.0676e-16
45.5 | 1.71e-15 | 2.1908e-15
45.6 | 6.19e-16 | 1.0998e-15
45.7 | 4.94e-16 | 9.7476e-16
45.8 | 2.52e-15 | 3.0008e-15
45.9 | 7.72e-16 | 1.2528e-15
46 | 3.09e-16 | 7.8976e-16
46.1 | 4.89e-17 | 5.2966e-16
46.2 | 1.77e-16 | 6.5776e-16
46.3 | 2.85e-15 | 3.3308e-15
46.4 | 6.06e-16 | 1.0868e-15
46.5 | -6.27e-17 | 4.1806e-16
46.6 | -2.55e-16 | 2.2576e-16
46.7 | -1.52e-16 | 3.2876e-16
46.8 | -2.53e-16 | 2.2776e-16
46.9 | -1.57e-16 | 3.2376e-16
47 | -2.69e-17 | 4.5386e-16
47.1 | -3.64e-17 | 4.4436e-16
47.2 | 7.92e-17 | 5.5996e-16
47.3 | 1.97e-16 | 6.7776e-16
47.4 | -8.42e-17 | 3.9656e-16
47.5 | -2.8e-16 | 2.0076e-16
47.6 | -3.98e-16 | 8.276e-17
47.7 | -3.19e-16 | 1.6176e-16
47.8 | -3.36e-16 | 1.4476e-16
47.9 | -2.77e-16 | 2.0376e-16
48 | 1.77e-16 | 6.5776e-16
48.1 | 3.28e-16 | 8.0876e-16
48.2 | 6.68e-16 | 1.1488e-15
48.3 | 4.87e-16 | 9.6776e-16
48.4 | 7.38e-17 | 5.5456e-16
48.5 | -1.59e-16 | 3.2176e-16
48.6 | 3.45e-15 | 3.9308e-15
48.7 | 3.83e-16 | 8.6376e-16
48.8 | -1.34e-16 | 3.4676e-16
48.9 | -3e-16 | 1.8076e-16
49 | -2.72e-16 | 2.0876e-16
49.1 | -3.24e-16 | 1.5676e-16
49.2 | -3.35e-16 | 1.4576e-16
49.3 | -2.21e-16 | 2.5976e-16
49.4 | -2.66e-16 | 2.1476e-16
49.5 | -3.98e-16 | 8.276e-17
49.6 | -4.05e-16 | 7.576e-17
49.7 | -3.82e-16 | 9.876e-17
49.8 | -3.11e-16 | 1.6976e-16
49.9 | -3.56e-16 | 1.2476e-16
50 | -1.75e-16 | 3.0576e-16
50.1 | 1.02e-16 | 5.8276e-16
50.2 | 2.6e-16 | 7.4076e-16
50.3 | 5.41e-16 | 1.0218e-15
50.4 | 5.46e-16 | 1.0268e-15
50.5 | 2.46e-16 | 7.2676e-16
50.6 | 5.72e-17 | 5.3796e-16
50.7 | -1.92e-16 | 2.8876e-16
50.8 | -1.94e-16 | 2.8676e-16
50.9 | -2.02e-16 | 2.7876e-16
51 | -2.72e-16 | 2.0876e-16
51.1 | -1.85e-16 | 2.9576e-16
51.2 | -2.37e-16 | 2.4376e-16
51.3 | 5.52e-17 | 5.3596e-16
51.4 | 1.38e-15 | 1.8608e-15
51.5 | 5.45e-16 | 1.0258e-15
51.6 | 7.77e-17 | 5.5846e-16
51.7 | -2e-16 | 2.8076e-16
51.8 | -2.54e-16 | 2.2676e-16
51.9 | -1.92e-16 | 2.8876e-16
52 | -2e-16 | 2.8076e-16
52.1 | -1.94e-16 | 2.8676e-16
52.2 | 3.56e-16 | 8.3676e-16
52.3 | 3.59e-16 | 8.3976e-16
52.4 | 3.76e-16 | 8.5676e-16
52.5 | 2.41e-15 | 2.8908e-15
52.6 | 1.33e-15 | 1.8108e-15
52.7 | 1.73e-16 | 6.5376e-16
52.8 | -1.11e-16 | 3.6976e-16
52.9 | -2.76e-16 | 2.0476e-16
53 | -2.88e-16 | 1.9276e-16
53.1 | -3.16e-16 | 1.6476e-16
53.2 | -1.28e-16 | 3.5276e-16
53.3 | -2.2e-17 | 4.5876e-16
53.4 | 1.59e-17 | 4.9666e-16
53.5 | -1.95e-16 | 2.8576e-16
53.6 | -2.27e-16 | 2.5376e-16
53.7 | -3.04e-16 | 1.7676e-16
53.8 | -3.68e-16 | 1.1276e-16
53.9 | -3.68e-16 | 1.1276e-16
54 | -3.61e-16 | 1.1976e-16
54.1 | -3.78e-16 | 1.0276e-16
54.2 | -2.58e-16 | 2.2276e-16
54.3 | -1.67e-16 | 3.1376e-16
54.4 | -1.22e-16 | 3.5876e-16
54.5 | -2.82e-16 | 1.9876e-16
54.6 | -2.95e-16 | 1.8576e-16
54.7 | -2.1e-16 | 2.7076e-16
54.8 | -1.73e-16 | 3.0776e-16
54.9 | -2.86e-16 | 1.9476e-16
55 | -3.11e-16 | 1.6976e-16
55.1 | -2.04e-16 | 2.7676e-16
55.2 | -3.34e-18 | 4.7742e-16
55.3 | 3.58e-16 | 8.3876e-16
55.4 | 1.08e-15 | 1.5608e-15
55.5 | 3.01e-16 | 7.8176e-16
55.6 | -1.12e-16 | 3.6876e-16
55.7 | -1.48e-16 | 3.3276e-16
55.8 | -2.53e-16 | 2.2776e-16
55.9 | -8.28e-19 | 4.7993e-16
56 | 1.2e-15 | 1.6808e-15
56.1 | 1.51e-16 | 6.3176e-16
56.2 | -5.91e-18 | 4.7485e-16
56.3 | 1.55e-15 | 2.0308e-15
56.4 | 2.08e-15 | 2.5608e-15
56.5 | 2.07e-15 | 2.5508e-15
56.6 | 1.69e-15 | 2.1708e-15
56.7 | 3.07e-16 | 7.8776e-16
56.8 | -1.15e-16 | 3.6576e-16
56.9 | -2.05e-16 | 2.7576e-16
57 | -1.94e-16 | 2.8676e-16
57.1 | -1.76e-16 | 3.0476e-16
57.2 | -2.94e-16 | 1.8676e-16
57.3 | -1.37e-16 | 3.4376e-16
57.4 | -1.9e-16 | 2.9076e-16
57.5 | -4.01e-17 | 4.4066e-16
57.6 | 9.12e-16 | 1.3928e-15
57.7 | 2.25e-16 | 7.0576e-16
57.8 | -1.03e-16 | 3.7776e-16
57.9 | -1.64e-17 | 4.6436e-16
58 | -1.22e-16 | 3.5876e-16
58.1 | -2.61e-16 | 2.1976e-16
58.2 | -3.49e-16 | 1.3176e-16
58.3 | -2.52e-16 | 2.2876e-16
58.4 | -1.81e-16 | 2.9976e-16
58.5 | -2.42e-16 | 2.3876e-16
58.6 | -2.15e-16 | 2.6576e-16
58.7 | -3.11e-16 | 1.6976e-16
58.8 | -2.75e-16 | 2.0576e-16
58.9 | -3.01e-16 | 1.7976e-16
59 | -2.59e-16 | 2.2176e-16
59.1 | -2.42e-16 | 2.3876e-16
59.2 | -3.43e-16 | 1.3776e-16
59.3 | -3.79e-16 | 1.0176e-16
59.4 | -4.13e-16 | 6.776e-17
59.5 | -3.36e-16 | 1.4476e-16
59.6 | -2.44e-16 | 2.3676e-16
59.7 | -2.07e-16 | 2.7376e-16
59.8 | -3.06e-17 | 4.5016e-16
59.9 | 2.38e-16 | 7.1876e-16
60 | -4.38e-17 | 4.3696e-16
60.1 | -4.24e-17 | 4.3836e-16
60.2 | 1.83e-15 | 2.3108e-15
60.3 | 7.48e-16 | 1.2288e-15
60.4 | 1.57e-16 | 6.3776e-16
60.5 | 3.56e-17 | 5.1636e-16
60.6 | 8.94e-18 | 4.897e-16
60.7 | -1.4e-16 | 3.4076e-16
60.8 | -2.71e-16 | 2.0976e-16
60.9 | -1.11e-16 | 3.6976e-16
61 | 4.43e-18 | 4.8519e-16
61.1 | -3.04e-16 | 1.7676e-16
61.2 | -2.69e-16 | 2.1176e-16
61.3 | -2.3e-16 | 2.5076e-16
61.4 | -2.58e-16 | 2.2276e-16
61.5 | -2.99e-16 | 1.8176e-16
61.6 | -3.02e-16 | 1.7876e-16
61.7 | -3.27e-16 | 1.5376e-16
61.8 | -3.06e-16 | 1.7476e-16
61.9 | 2.6e-16 | 7.4076e-16
62 | 5.98e-17 | 5.4056e-16
62.1 | -2.37e-16 | 2.4376e-16
62.2 | -2.5e-16 | 2.3076e-16
62.3 | -1.61e-16 | 3.1976e-16
62.4 | -2.97e-16 | 1.8376e-16
62.5 | -3.51e-16 | 1.2976e-16
62.6 | -3.26e-16 | 1.5476e-16
62.7 | -3.88e-16 | 9.276e-17
62.8 | -3.95e-16 | 8.576e-17
62.9 | -4.27e-16 | 5.376e-17
63 | -3.71e-16 | 1.0976e-16
63.1 | -3.73e-16 | 1.0776e-16
63.2 | -3.71e-16 | 1.0976e-16
63.3 | -3.32e-16 | 1.4876e-16
63.4 | -3.03e-16 | 1.7776e-16
63.5 | -2.64e-16 | 2.1676e-16
63.6 | -1.3e-16 | 3.5076e-16
63.7 | -3.11e-16 | 1.6976e-16
63.8 | -2.85e-16 | 1.9576e-16
63.9 | -2.16e-16 | 2.6476e-16
64 | 1.22e-16 | 6.0276e-16
64.1 | 4.82e-16 | 9.6276e-16
64.2 | 2.24e-15 | 2.7208e-15
64.3 | 7.16e-16 | 1.1968e-15
64.4 | 7e-17 | 5.5076e-16
64.5 | -3.7e-17 | 4.4376e-16
64.6 | 5.21e-17 | 5.3286e-16
64.7 | -1.16e-16 | 3.6476e-16
64.8 | -3.41e-16 | 1.3976e-16
64.9 | -3.15e-16 | 1.6576e-16
65 | -2.93e-16 | 1.8776e-16
65.1 | -2.68e-16 | 2.1276e-16
65.2 | -4.28e-16 | 5.276e-17
65.3 | -3.73e-16 | 1.0776e-16
65.4 | -2.77e-16 | 2.0376e-16
65.5 | -2.89e-16 | 1.9176e-16
65.6 | -2.77e-16 | 2.0376e-16
65.7 | -2.49e-16 | 2.3176e-16
65.8 | -3.57e-16 | 1.2376e-16
65.9 | -3.65e-16 | 1.1576e-16
66 | 2.99e-15 | 3.4708e-15
66.1 | 6.31e-16 | 1.1118e-15
66.2 | -1.29e-16 | 3.5176e-16
66.3 | -2.76e-16 | 2.0476e-16
66.4 | -2.33e-16 | 2.4776e-16
66.5 | -1.87e-16 | 2.9376e-16
66.6 | -3.07e-16 | 1.7376e-16
66.7 | -3.51e-16 | 1.2976e-16
66.8 | -3.1e-16 | 1.7076e-16
66.9 | 1.97e-15 | 2.4508e-15
67 | 6.24e-16 | 1.1048e-15
67.1 | 5.49e-17 | 5.3566e-16
67.2 | -1.77e-16 | 3.0376e-16
67.3 | -3.34e-16 | 1.4676e-16
67.4 | -3.36e-16 | 1.4476e-16
67.5 | -2.57e-16 | 2.2376e-16
67.6 | -2.69e-16 | 2.1176e-16
67.7 | -2.83e-16 | 1.9776e-16
67.8 | -3.55e-16 | 1.2576e-16
67.9 | -3.69e-16 | 1.1176e-16
68 | -3.3e-16 | 1.5076e-16
68.1 | -3.17e-16 | 1.6376e-16
68.2 | -3.5e-16 | 1.3076e-16
68.3 | -3.38e-16 | 1.4276e-16
68.4 | -1.23e-16 | 3.5776e-16
68.5 | -2.63e-16 | 2.1776e-16
68.6 | 1.86e-17 | 4.9936e-16
68.7 | 2.46e-16 | 7.2676e-16
68.8 | -1.21e-16 | 3.5976e-16
68.9 | 9.79e-17 | 5.7866e-16
69 | 9.5e-17 | 5.7576e-16
69.1 | -1.39e-16 | 3.4176e-16
69.2 | -2.11e-16 | 2.6976e-16
69.3 | -3.4e-16 | 1.4076e-16
69.4 | -9.69e-17 | 3.8386e-16
69.5 | -3.57e-17 | 4.4506e-16
69.6 | 9.14e-17 | 5.7216e-16
69.7 | 2.19e-15 | 2.6708e-15
69.8 | 3.48e-16 | 8.2876e-16
69.9 | -8.26e-17 | 3.9816e-16
70 | -3.05e-16 | 1.7576e-16
70.1 | -4.17e-16 | 6.376e-17
70.2 | -3.03e-16 | 1.7776e-16
70.3 | -2.82e-16 | 1.9876e-16
70.4 | -3.03e-16 | 1.7776e-16
70.5 | -3.22e-16 | 1.5876e-16
70.6 | -2.89e-16 | 1.9176e-16
70.7 | -3.28e-16 | 1.5276e-16
70.8 | -3.14e-16 | 1.6676e-16
70.9 | -2.77e-16 | 2.0376e-16
71 | -2.51e-16 | 2.2976e-16
71.1 | -3.37e-16 | 1.4376e-16
71.2 | -2.81e-16 | 1.9976e-16
71.3 | -3.46e-16 | 1.3476e-16
71.4 | -2.96e-16 | 1.8476e-16
71.5 | -3.64e-16 | 1.1676e-16
71.6 | 1.7e-16 | 6.5076e-16
71.7 | 9.04e-17 | 5.7116e-16
71.8 | 8.47e-17 | 5.6546e-16
71.9 | -1e-16 | 3.8076e-16
72 | -3.14e-16 | 1.6676e-16
72.1 | -3.23e-16 | 1.5776e-16
72.2 | -3.41e-16 | 1.3976e-16
72.3 | -3.27e-16 | 1.5376e-16
72.4 | -3.04e-16 | 1.7676e-16
72.5 | -3e-16 | 1.8076e-16
72.6 | -3.16e-16 | 1.6476e-16
72.7 | -3.81e-16 | 9.976e-17
72.8 | -3.04e-16 | 1.7676e-16
72.9 | -1.32e-16 | 3.4876e-16
73 | -3.89e-16 | 9.176e-17
73.1 | 4.45e-16 | 9.2576e-16
73.2 | -5.85e-17 | 4.2226e-16
73.3 | -2.07e-16 | 2.7376e-16
73.4 | -3e-16 | 1.8076e-16
73.5 | -3.04e-16 | 1.7676e-16
73.6 | -2.42e-16 | 2.3876e-16
73.7 | -1.83e-16 | 2.9776e-16
73.8 | 4.77e-16 | 9.5776e-16
73.9 | -8.01e-17 | 4.0066e-16
74 | -2.46e-16 | 2.3476e-16
74.1 | -2.66e-16 | 2.1476e-16
74.2 | -3.49e-16 | 1.3176e-16
74.3 | -2.75e-16 | 2.0576e-16
74.4 | -2.25e-16 | 2.5576e-16
74.5 | -2.9e-16 | 1.9076e-16
74.6 | -4.03e-16 | 7.776e-17
74.7 | -1.12e-16 | 3.6876e-16
74.8 | -2.87e-16 | 1.9376e-16
74.9 | -4.11e-16 | 6.976e-17
75 | 1.33e-15 | 1.8108e-15
75.1 | 2.86e-16 | 7.6676e-16
75.2 | -9.59e-17 | 3.8486e-16
75.3 | -4.85e-17 | 4.3226e-16
75.4 | 6.86e-17 | 5.4936e-16
75.5 | -1.69e-16 | 3.1176e-16
75.6 | -2.51e-16 | 2.2976e-16
75.7 | -2.58e-16 | 2.2276e-16
75.8 | -3.84e-16 | 9.676e-17
75.9 | -3.33e-16 | 1.4776e-16
76 | -3.34e-16 | 1.4676e-16
76.1 | -1.52e-16 | 3.2876e-16
76.2 | -2.32e-16 | 2.4876e-16
76.3 | -2.96e-16 | 1.8476e-16
76.4 | -2.88e-16 | 1.9276e-16
76.5 | -4.23e-16 | 5.776e-17
76.6 | -3.53e-16 | 1.2776e-16
76.7 | 1.34e-15 | 1.8208e-15
76.8 | 1.63e-16 | 6.4376e-16
76.9 | -2.19e-16 | 2.6176e-16
77 | -2.19e-16 | 2.6176e-16
77.1 | -2.83e-16 | 1.9776e-16
77.2 | -3.58e-16 | 1.2276e-16
77.3 | -3.78e-16 | 1.0276e-16
77.4 | -2.37e-16 | 2.4376e-16
77.5 | -3.01e-16 | 1.7976e-16
77.6 | -3.27e-16 | 1.5376e-16
77.7 | -1.62e-16 | 3.1876e-16
77.8 | 3.49e-17 | 5.1566e-16
77.9 | 1.86e-15 | 2.3408e-15
78 | 4.37e-16 | 9.1776e-16
78.1 | 1.66e-16 | 6.4676e-16
78.2 | 1.29e-16 | 6.0976e-16
78.3 | 5.79e-16 | 1.0598e-15
78.4 | 9.83e-17 | 5.7906e-16
78.5 | -1.45e-16 | 3.3576e-16
78.6 | 3.28e-17 | 5.1356e-16
78.7 | 1.94e-16 | 6.7476e-16
78.8 | 4.24e-16 | 9.0476e-16
78.9 | 2.67e-16 | 7.4776e-16
79 | 6.74e-17 | 5.4816e-16
79.1 | -2.22e-16 | 2.5876e-16
79.2 | -2.74e-16 | 2.0676e-16
79.3 | -3.63e-16 | 1.1776e-16
79.4 | -3.22e-16 | 1.5876e-16
79.5 | -3.73e-16 | 1.0776e-16
79.6 | -3.43e-16 | 1.3776e-16
79.7 | -3.59e-16 | 1.2176e-16
79.8 | -3.13e-16 | 1.6776e-16
79.9 | -3.25e-16 | 1.5576e-16
80 | -2.97e-16 | 1.8376e-16
80.1 | 5.17e-16 | 9.9776e-16
80.2 | -1.13e-16 | 3.6776e-16
80.3 | -2.09e-16 | 2.7176e-16
80.4 | -2.57e-16 | 2.2376e-16
80.5 | -3.22e-16 | 1.5876e-16
80.6 | 5.48e-17 | 5.3556e-16
80.7 | 6.73e-17 | 5.4806e-16
80.8 | -2.2e-16 | 2.6076e-16
80.9 | -2.85e-16 | 1.9576e-16
81 | -3.72e-16 | 1.0876e-16
81.1 | -3.33e-16 | 1.4776e-16
81.2 | -2.98e-16 | 1.8276e-16
81.3 | -4.76e-16 | 4.76e-18
81.4 | -4.3e-16 | 5.076e-17
81.5 | -3.09e-16 | 1.7176e-16
81.6 | -2.94e-16 | 1.8676e-16
81.7 | -2.9e-16 | 1.9076e-16
81.8 | -3.34e-16 | 1.4676e-16
81.9 | -3.75e-16 | 1.0576e-16
82 | -3.25e-16 | 1.5576e-16
82.1 | -2.96e-16 | 1.8476e-16
82.2 | 2.79e-15 | 3.2708e-15
82.3 | 7.5e-16 | 1.2308e-15
82.4 | -2.15e-17 | 4.5926e-16
82.5 | -1.4e-16 | 3.4076e-16
82.6 | -2.54e-16 | 2.2676e-16
82.7 | -3e-16 | 1.8076e-16
82.8 | -3.54e-16 | 1.2676e-16
82.9 | -3.68e-16 | 1.1276e-16
83 | -3.41e-16 | 1.3976e-16
83.1 | -3.61e-16 | 1.1976e-16
83.2 | -3.55e-16 | 1.2576e-16
83.3 | -3.03e-16 | 1.7776e-16
83.4 | -5.85e-17 | 4.2226e-16
83.5 | -1.81e-16 | 2.9976e-16
83.6 | -2.41e-16 | 2.3976e-16
83.7 | -3.05e-16 | 1.7576e-16
83.8 | -2.93e-16 | 1.8776e-16
83.9 | -2.83e-16 | 1.9776e-16
84 | -3.35e-16 | 1.4576e-16
84.1 | -3.31e-16 | 1.4976e-16
84.2 | -3.69e-16 | 1.1176e-16
84.3 | -3.57e-16 | 1.2376e-16
84.4 | -2.23e-17 | 4.5846e-16
84.5 | 4.36e-15 | 4.8408e-15
84.6 | 8.68e-16 | 1.3488e-15
84.7 | -4.21e-17 | 4.3866e-16
84.8 | -2.49e-16 | 2.3176e-16
84.9 | -2.83e-16 | 1.9776e-16
85 | -3.15e-16 | 1.6576e-16
85.1 | -3.02e-16 | 1.7876e-16
85.2 | -3.71e-16 | 1.0976e-16
85.3 | -3.37e-16 | 1.4376e-16
85.4 | -3.72e-16 | 1.0876e-16
85.5 | -3.2e-16 | 1.6076e-16
85.6 | -3.6e-16 | 1.2076e-16
85.7 | -3.86e-16 | 9.476e-17
85.8 | -2.59e-16 | 2.2176e-16
85.9 | -3.33e-16 | 1.4776e-16
86 | -3.98e-16 | 8.276e-17
86.1 | -3.18e-16 | 1.6276e-16
86.2 | -3.29e-16 | 1.5176e-16
86.3 | -2.83e-16 | 1.9776e-16
86.4 | -2.17e-16 | 2.6376e-16
86.5 | -2.88e-16 | 1.9276e-16
86.6 | 9.59e-16 | 1.4398e-15
86.7 | 5.18e-16 | 9.9876e-16
86.8 | 2.91e-16 | 7.7176e-16
86.9 | 2.45e-16 | 7.2576e-16
87 | 6e-17 | 5.4076e-16
87.1 | -1.13e-17 | 4.6946e-16
87.2 | -2.47e-16 | 2.3376e-16
87.3 | -3.02e-16 | 1.7876e-16
87.4 | -2.79e-16 | 2.0176e-16
87.5 | -3.56e-16 | 1.2476e-16
87.6 | -2.99e-16 | 1.8176e-16
87.7 | -3.8e-16 | 1.0076e-16
87.8 | -3.48e-16 | 1.3276e-16
87.9 | -3.41e-16 | 1.3976e-16
88 | -2.1e-16 | 2.7076e-16
88.1 | -2.66e-16 | 2.1476e-16
88.2 | -3.06e-16 | 1.7476e-16
88.3 | -3.96e-16 | 8.476e-17
88.4 | -3.33e-16 | 1.4776e-16
88.5 | -3.13e-16 | 1.6776e-16
88.6 | -3.22e-16 | 1.5876e-16
88.7 | -3.74e-16 | 1.0676e-16
88.8 | 8.12e-16 | 1.2928e-15
88.9 | 1.55e-16 | 6.3576e-16
89 | -1.68e-16 | 3.1276e-16
89.1 | 4.53e-15 | 5.0108e-15
89.2 | 4.51e-16 | 9.3176e-16
89.3 | -9.38e-17 | 3.8696e-16
89.4 | 2.6e-15 | 3.0808e-15
89.5 | 6.08e-16 | 1.0888e-15
89.6 | -5.86e-17 | 4.2216e-16
89.7 | -2.76e-16 | 2.0476e-16
89.8 | -2.93e-16 | 1.8776e-16
89.9 | -6.01e-18 | 4.7475e-16
90 | -1.12e-16 | 3.6876e-16
90.1 | -2.15e-16 | 2.6576e-16
90.2 | -3.09e-16 | 1.7176e-16
90.3 | -2.59e-16 | 2.2176e-16
90.4 | -8.21e-17 | 3.9866e-16
90.5 | -1.67e-16 | 3.1376e-16
90.6 | -1.68e-16 | 3.1276e-16
90.7 | -2.64e-16 | 2.1676e-16
90.8 | -1.04e-16 | 3.7676e-16
90.9 | -2.69e-16 | 2.1176e-16
91 | -3.12e-16 | 1.6876e-16
91.1 | -2.58e-16 | 2.2276e-16
91.2 | -2.81e-16 | 1.9976e-16
91.3 | -2.49e-16 | 2.3176e-16
91.4 | -3e-16 | 1.8076e-16
91.5 | -3.3e-16 | 1.5076e-16
91.6 | -3.62e-16 | 1.1876e-16
91.7 | -3.9e-16 | 9.076e-17
91.8 | -3.69e-16 | 1.1176e-16
91.9 | -2.92e-16 | 1.8876e-16
92 | -2.97e-16 | 1.8376e-16
92.1 | 1.38e-15 | 1.8608e-15
92.2 | 1.13e-15 | 1.6108e-15
92.3 | 1.77e-16 | 6.5776e-16
92.4 | -1.15e-16 | 3.6576e-16
92.5 | -2.25e-16 | 2.5576e-16
92.6 | -3.2e-16 | 1.6076e-16
92.7 | -7.56e-17 | 4.0516e-16
92.8 | -2.53e-16 | 2.2776e-16
92.9 | -3.11e-16 | 1.6976e-16
93 | -3.62e-17 | 4.4456e-16
93.1 | -1.69e-16 | 3.1176e-16
93.2 | -3.45e-16 | 1.3576e-16
93.3 | -2.51e-16 | 2.2976e-16
93.4 | -3.03e-16 | 1.7776e-16
93.5 | -3.08e-16 | 1.7276e-16
93.6 | -3.11e-16 | 1.6976e-16
93.7 | 1.37e-15 | 1.8508e-15
93.8 | 1.62e-16 | 6.4276e-16
93.9 | -1.56e-16 | 3.2476e-16
94 | -2.39e-16 | 2.4176e-16
94.1 | -3.21e-16 | 1.5976e-16
94.2 | -3.34e-16 | 1.4676e-16
94.3 | -3.72e-16 | 1.0876e-16
94.4 | -3.67e-16 | 1.1376e-16
94.5 | -3.59e-16 | 1.2176e-16
94.6 | -3.03e-16 | 1.7776e-16
94.7 | -2.74e-16 | 2.0676e-16
94.8 | -2.2e-16 | 2.6076e-16
94.9 | -3.15e-16 | 1.6576e-16
95 | -3.17e-16 | 1.6376e-16
95.1 | -3.42e-16 | 1.3876e-16
95.2 | 9.36e-16 | 1.4168e-15
95.3 | 5.53e-17 | 5.3606e-16
95.4 | -1.79e-16 | 3.0176e-16
95.5 | -3.59e-17 | 4.4486e-16
95.6 | -3.42e-17 | 4.4656e-16
95.7 | -9.79e-17 | 3.8286e-16
95.8 | -2.51e-16 | 2.2976e-16
95.9 | -2.75e-16 | 2.0576e-16
96 | -2.2e-16 | 2.6076e-16
96.1 | -3.41e-16 | 1.3976e-16
96.2 | -3.17e-16 | 1.6376e-16
96.3 | 1.26e-16 | 6.0676e-16
96.4 | -1.32e-16 | 3.4876e-16
96.5 | -2.78e-16 | 2.0276e-16
96.6 | -2.3e-16 | 2.5076e-16
96.7 | -3.24e-16 | 1.5676e-16
96.8 | -2.76e-16 | 2.0476e-16
96.9 | -2.41e-16 | 2.3976e-16
97 | 3.8e-16 | 8.6076e-16
97.1 | 7.34e-16 | 1.2148e-15
97.2 | -7.54e-17 | 4.0536e-16
97.3 | -1.03e-17 | 4.7046e-16
97.4 | -1.63e-16 | 3.1776e-16
97.5 | -2.62e-16 | 2.1876e-16
97.6 | -2.02e-16 | 2.7876e-16
97.7 | 1.35e-15 | 1.8308e-15
97.8 | 1.54e-15 | 2.0208e-15
97.9 | 2.87e-16 | 7.6776e-16
98 | 2.02e-17 | 5.0096e-16
98.1 | -1.19e-16 | 3.6176e-16
98.2 | -1.68e-16 | 3.1276e-16
98.3 | -3.09e-16 | 1.7176e-16
98.4 | -3.11e-16 | 1.6976e-16
98.5 | -3.06e-16 | 1.7476e-16
98.6 | -2.98e-16 | 1.8276e-16
98.7 | -3.43e-16 | 1.3776e-16
98.8 | -3.17e-16 | 1.6376e-16
98.9 | -3.15e-16 | 1.6576e-16
99 | -3.52e-16 | 1.2876e-16
99.1 | 7.43e-17 | 5.5506e-16
99.2 | 1.14e-16 | 5.9476e-16
99.3 | -1.89e-16 | 2.9176e-16
99.4 | -2.55e-16 | 2.2576e-16
99.5 | -2.59e-16 | 2.2176e-16
99.6 | -3.03e-16 | 1.7776e-16
99.7 | -3.15e-16 | 1.6576e-16
99.8 | -3.76e-16 | 1.0476e-16
99.9 | -3.31e-16 | 1.4976e-16
100 | -3.61e-16 | 1.1976e-16

## C83010 aLIGO AOSEM Flex Cir (2)
2010-08-30 | 2010 07:56:01 AM
Residual Gas Analyzer Software
RGA Software Version | 3.205.003
Pressure vs Time Scan Setup:
Active channels in Scan | 1
Units | Amps
Sample Period | 1 | sec
Focus Voltage | 90 | Volts
Ion Energy | HIGH
Electron Energy | 70 | eV
CEM Voltage | 1000 | Volts
CEM Gain | 1000000
Sensitivity Factor | 0.0001
Filament Current | 1 | mAmps
Focus Voltage | 90 | Volts
Ion Energy | HIGH
Electron Energy | 70 | eV
CEM Voltage | 1000 | Volts
CEM Gain | 1000000
Sensitivity Factor | 0.0001
Filament Current | 1 | mAmps
Start time | Aug 30 | 2010 07:56:01 AM
Channel | Mass(amu) | Name | Cal Factor | Noise Floor | CEM Status
1 40.00 Argon 0.18 4 OFF
Note: *.rga file indicates CEM is on
Time(s) Channel#1
0 | 2.0556e-15
1 | 1.3681e-15
2 | 2.7913e-15
3 | 7.3927e-15
4 | 2.5871e-15
5 | 1.0476e-15
6 | 2.1007e-15
7 | 1.6303e-15
8 | 3.9773e-15
9 | 5.8975e-15
10 | 7.5821e-16
11 | 1.2339e-15
12 | 1.4566e-15
13 | 2.2662e-15
14 | 1.519e-15
15 | 1.7544e-15
16 | 6.8122e-15
17 | 1.1776e-15
18 | 1.0788e-15
19 | 2.6175e-15
20 | 1.2848e-15
21 | 2.4982e-15
22 | 3.9854e-11
23 | 1.0401e-13
24 | 2.756e-14
25 | 2.1546e-14
26 | 2.2553e-14
27 | 1.8964e-14
28 | 1.6836e-14
29 | 1.407e-14
30 | 1.2374e-14
31 | 1.4013e-14
32 | 1.6514e-14
33 | 2.1637e-14
34 | 3.0047e-14
35 | 1.5386e-14
36 | 2.2076e-14
37 | 2.2154e-14
38 | 1.7664e-14
39 | 2.2045e-14
40 | 2.0909e-14
41 | 1.7399e-14
42 | 1.5244e-14
43 | 1.3296e-14
44 | 1.795e-14
45 | 1.2582e-14
46 | 1.7393e-14
47 | 1.6888e-14
48 | 1.8326e-14
49 | 1.6406e-14
50 | 1.9508e-14
51 | 1.7327e-14
52 | 2.9948e-14
53 | 1.5784e-14
54 | 2.6258e-14
55 | 1.7275e-14
56 | 1.612e-14
57 | 1.2445e-15
58 | 1.4814e-15
59 | 1.3012e-15
60 | 1.1404e-15
61 | 1.0556e-15
62 | 8.9808e-16
63 | 2.9274e-15
64 | 9.949e-16
65 | 4.5435e-15
66 | 1.4465e-15
67 | 8.2127e-15
68 | 1.5245e-15
69 | 1.3407e-15
70 | 1.0172e-15
71 | 2.295e-15
72 | 1.2992e-15
73 | 8.1161e-16
74 | 1.0804e-15
75 | 1.611e-15
76 | 1.376e-15
77 | 2.5927e-15
78 | 1.3124e-15
79 | 7.066e-15
80 | 1.0835e-15
81 | 9.7304e-16
82 | 1.8972e-15
83 | 2.35e-15
84 | 1.2402e-15
85 | 1.5795e-15
86 | 1.2499e-15
87 | 1.081e-15
88 | 1.5953e-15
89 | 2.0031e-15
90 | 2.1423e-15
91 | 6.7765e-15
92 | 1.4228e-15
93 | 1.3693e-15
94 | 1.0615e-15
95 | 1.4468e-15
96 | 1.3306e-15
97 | 1.4957e-15
98 | 1.2387e-15
99 | 4.565e-15
100 | 4.3762e-15
101 | 2.0758e-15
102 | 1.7118e-15
103 | 1.3624e-15
104 | 1.625e-15
105 | 8.0933e-16 |  | 1.5306e-15
106 | 1.2145e-15
107 | 1.3391e-15
108 | 1.2317e-15
109 | 9.0334e-16
110 | 1.5325e-15
111 | 1.3356e-15
112 | 1.9894e-15
113 | 1.4271e-15
114 | 1.8327e-15
115 | 1.5471e-15
116 | 1.3329e-15
117 | 1.6416e-15
118 | 1.5193e-15
119 | 1.2499e-15
120 | 9.6562e-16
121 | 1.6093e-15
122 | 1.2036e-15
123 | 2.12e-15
124 | 2.4089e-15
125 | 2.0921e-15
126 | 1.1322e-15
127 | 9.3798e-16
128 | 1.8304e-15
129 | 2.3562e-15
130 | 1.3228e-15
131 | 2.4397e-15

## Calibration
Calibration with leak
Calib leak rate | 2.36e-10 | torr l/s | (Argon)
pump speed | ? | l/s | oven F
AMU 40 (w/leak open) | 2e-14 | amps
AMU 40 (background) | 1.5e-15 | amps
Calib leak contributes | 1.85e-14 | amps | = (w/leak open) - (background)
Leak Rate Calibration factor | 12756.7568 | torr l/s/A
Pressure Calibration factor | #VALUE! | torr/A
Leak Rate Cal factor per Flex Assy | 57.2052 | torr l/s/A/AOSEM
The load consists of 223 A-OSEM Flex Circuit Assemblies
| 223 | units | for the test
| 104 | units | in the vertex volume
| 112 | units | in the LHO diagonal volume
| 8 | units | in the LHO end station volume