Material Qualification Queue

as of 16 Nov-2010

in priority order

# IN-TEST:

- Masterbond low outgassing epoxy (EP30-2): repeat RGA and optical contamination cavity testing with epoxy bonded between 1" glass slides (~.001 thick)
As always, take extreme care in proper mixing ratios. Bond 10 pairs of glass disks of 1" diameter together to obtain 10 bonds each 1" diameter and ~.0.001" thick.

Measure the pair thicknesses before and after epoxying to gauge the bond thickness. You could get glass windows from Edmunds:

http://www.edmundoptics.com/onlinecatalog/displayproduct.cfm?productID=1755&PageNum=2&StartRow=21

The NT45-640 is an uncoated 25mm diam disc.

According to Master Bond reps, the best cure schedule (yielding the strongest bonds) is: overnight (say 12 hr) room temperature cure, followed by a heat cure of 3 hours at 150 deg. F. See the data sheet for the EP30-2 in the dcc:

https://dcc.ligo.org/cgi-bin/private/DocDB/ShowDocument?docid=15100

- Perkin Elmer 3mm InGaAs Photodiode, Part Number - C30665GH 
Intended Use - aLIGO DC readout detectors (ISC)
Quantity - 10 
Purchase Order Number - CP748923 
Max Storage Temperature (from data sheet) - 125 C 
1.  Clean exterior of diodes 
2.  Cut can lids off with clean cutting tool (Chub has tool) 
3.  Bake at 120 C 
4.  Perform RGA scan 
5.  Perform Cavity Contamination Scan

- Chamber 3 being re-commissioned after vacuum failure

# QUEUE

- a) RTD device (Omega #1PT100KN1510CL1/3) 
for TCS temperature measurement
platinum wire wound
embedded in alumina?
b) Fiberglass insulated RTD extension cable (Omega # PN EXGG-4CU-26S)
for TCS temperature measurement

- EG&G YAG-444A -- the proposed diodes for the arm cavity baffle
Bob to order ASAP

- Nedox SF-2 coating (by General Magnaplate) Coating Material for TCS in-vacuum relay mirror (will be applied to the I.D. and O.D. surfaces of a .50" tube)
There will be six specimens, each 16" long for testing
Mindy has ordered samples for testing

- RG-316 (Allied)
Rich Abbott ordered a spool. Chub asked to give to Bob 10/1. 
1.  Manufacturer - Allied Wire and Cable 
http://www.awcwire.com/Part.aspx?code=213F27F27J34

2.  Materials:
Part Number RG316 
Conductor Stranding 7/.0067 
Nom. Dia. of Cond. 0.021 
Dielectric (in) 0.06 
Nom. O.D. (in) 0.102 
Nom. Imp. 50 ohms
Approx LBS/MFT 11 (I think this should be kFt)
Nom. Cap. (pF/ft) 29.4
Shield Material: Silver-Coated Copper 
Conductor Material: Silver-Coated Copper Clad Steel 
Dielectric Material: Polytetrafluoroethylene (PTFE) 
Jacket Material: Fluorinated Ethylene Propylene (FEP) 
Shield: Overall Braid Shield 
Min. Temp -55°C 
Max. Temp 80°C

3.  Application: In vacuum LSC and ASC detectors

4.  Quantities:
The anticipated length for each coaxial bundle in vacuum is ~15ft.
Each LSC detector uses a bundle of 5 individual RG316, there are 2 LSC in-vacuum detectors per IFO
Each ASC detector uses two bundles of 5 individual RG316, there are 4 ASC in-vacuum detectors per IFO

Total lengths of cable(#detectors*connectors per detector*length*# of coaxial cables)
2*1*15*5 = 150 feet for LSC per IFO
4*2*15*5 = 600 feet for ASC per IFO

Grand total of 150 + 600 = 750 feet of coaxial cable per IFO

5.  This is an identified need, must have.

- Ferrite Material
No currently defined need – would be nice to have qualified just in case
Rich Abbott needs to provide a significant number of samples (10 - 20 units) from a known source with a known material grade.
Ferrite materials are based on "Nickel-Zinc" and "Manganese-Zinc". Zinc has a high vapor pressure.
One possibility is BN-43-7051 is a Balun (binocular or multi-aperture) core sold by many companies, such as Amidon Corp.:
https://www.amidoncorp.com/specs/2-34.pdf
where "43" refers to the material.
Material 43 is a NiZn ferrite material:
http://www.cwsbytemark.com/CatalogSheets/Ferrite_datasheet_oct06/FR_MATL.pdf
http://www.fair-rite.com/newfair/materials43.htm

- Copper & polyimide clad fiber optic, IVG Fiber CU1300. 
Not yet a baseline element for aLIGO.
Brian Lantz wants to test this material for use with an optics table, optical lever system. Not yet a baseline element for aLIGO. Could be used for ALS etc. Likely "inherently" vacuum compatible, but worth checking since it has a polymer.

- Tungsten carbide/carbon (WC/C) sample (for potential BSC-ISI tooling use (class b) but perhaps useful in vacuum as well) (had small sample -- likely inadequate in size probably inherently vacuum compatible)
No currently defined need – would be nice to have qualified just in case

- Cesic sample (composite ceramic of SiC, Si and C) possibly inherently vacuum compatible application is not clear (high stiffness to weight structure) (sample size adequate?)
No currently defined need – would be nice to have qualified just in case