Material Qualification Queue

as of 28 Feb-2011

in priority order

# IN-TEST:

- a) RTD device (Omega #1PT100KN1510CL1/3) 
for TCS temperature measurement
platinum wire wound
embedded in alumina?
b) Fiberglass insulated RTD extension cable (Omega # PN EXGG-4CU-26S)
for TCS temperature measurement

- EG&G YAG-444A -- the proposed diodes for the arm cavity baffle
Bob to order ASAP

- Nedox SF-2 coating (by General Magnaplate) Coating Material for TCS in-vacuum relay mirror (will be applied to the I.D. and O.D. surfaces of a .50" tube)
There will be six specimens, each 16" long for testing

# QUEUE

- Conner wire, kapton/polyimide wire (28 awg), part number CW6951rev1.
(has a yellow color, rather than the more usual amber color)

LIGO-C0900156: RFQ No. SM-150, entitled “LIGO 28AWG In-Vacuum Wire” submitted to Cooner Wire Inc.
LIGO-T0900298-v1: Specification for 28AWG Kapton Insulated Wire

- RG-316 (Allied)
Rich Abbott ordered a spool. Chub asked to give to Bob 10/1. 
1.  Manufacturer - Allied Wire and Cable 
http://www.awcwire.com/Part.aspx?code=213F27F27J34

2.  Materials:
Part Number RG316 
Conductor Stranding 7/.0067 
Nom. Dia. of Cond. 0.021 
Dielectric (in) 0.06 
Nom. O.D. (in) 0.102 
Nom. Imp. 50 ohms
Approx LBS/MFT 11 (I think this should be kFt)
Nom. Cap. (pF/ft) 29.4
Shield Material: Silver-Coated Copper 
Conductor Material: Silver-Coated Copper Clad Steel 
Dielectric Material: Polytetrafluoroethylene (PTFE) 
Jacket Material: Fluorinated Ethylene Propylene (FEP) 
Shield: Overall Braid Shield 
Min. Temp -55°C 
Max. Temp 80°C

3.  Application: In vacuum LSC and ASC detectors

4.  Quantities:
The anticipated length for each coaxial bundle in vacuum is ~15ft.
Each LSC detector uses a bundle of 5 individual RG316, there are 2 LSC in-vacuum detectors per IFO
Each ASC detector uses two bundles of 5 individual RG316, there are 4 ASC in-vacuum detectors per IFO

Total lengths of cable(#detectors*connectors per detector*length*# of coaxial cables)
2*1*15*5 = 150 feet for LSC per IFO
4*2*15*5 = 600 feet for ASC per IFO

Grand total of 150 + 600 = 750 feet of coaxial cable per IFO

5.  This is an identified need, must have.

- Ferrite Material
No currently defined need – would be nice to have qualified just in case
Rich Abbott needs to provide a significant number of samples (10 - 20 units) from a known source with a known material grade.
Ferrite materials are based on "Nickel-Zinc" and "Manganese-Zinc". Zinc has a high vapor pressure.
One possibility is BN-43-7051 is a Balun (binocular or multi-aperture) core sold by many companies, such as Amidon Corp.:
https://www.amidoncorp.com/specs/2-34.pdf
where "43" refers to the material.
Material 43 is a NiZn ferrite material:
http://www.cwsbytemark.com/CatalogSheets/Ferrite_datasheet_oct06/FR_MATL.pdf
http://www.fair-rite.com/newfair/materials43.htm

- Copper & polyimide clad fiber optic, IVG Fiber CU1300. 
Not yet a baseline element for aLIGO.
Brian Lantz wants to test this material for use with an optics table, optical lever system. Not yet a baseline element for aLIGO. Could be used for ALS etc. Likely "inherently" vacuum compatible, but worth checking since it has a polymer.

- Tungsten carbide/carbon (WC/C) sample (for potential BSC-ISI tooling use (class b) but perhaps useful in vacuum as well) (had small sample -- likely inadequate in size probably inherently vacuum compatible)
No currently defined need – would be nice to have qualified just in case

- Cesic sample (composite ceramic of SiC, Si and C) possibly inherently vacuum compatible application is not clear (high stiffness to weight structure) (sample size adequate?)
No currently defined need – would be nice to have qualified just in case