LIGO Laboratory / LIGO Scientific Collaboration

E1000387-v4	        LIGO                           21 Oct. 2011

Stage 1-2 Vertical Actuator Assembly

S.Biscans, M.Hillard

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Reference D0902531

Note : ! Make sure there is no trash in the coil !

Remove the two shoulder screws and watch carefully inside the coil.

Clean room standards

For a clean assembly all LIGO standards should be followed, as presented in the latest version of the LIGO Contamination Control Plan (E0900047). Clean room garb including UHV gloves should be worn when working with parts.

All tools that come in contact with assembly should be cleaned to class B standards.

Assembly will be done under a portable clean room.  Any time a part of the assembly is not covered by the portable clean room or not being actively worked on it should be covered with appropriate clean covers. (C3 polyester or equivalent).

Parts prep work: alignment pins and helicoils

To facilitate assembly, parts requiring helicoils and/or pins to be inserted before assembly will be listed to allow preliminary preparation of the parts, after cleaning and before starting general assembly. Preparation need to be complete on the following. See E1000402 part pre-preparation matrix

Helicoils

Helicoil 1/2-23, 1XDIA lg.

#1185-8EN.500

1. D0902155

Helicoil 1/4-20, 1.5XDIA lg.

#1185-4EN.500

2. D0902161

Helicoil 1/4-20, 1.5XDIA lg.

#1185-4EN.375

3. D0902137

Helicoil 8-32, 2 X DIA lg.

#1185-2EN.328

HELICOILS NEED TO BE INSERTED FROM THE OTHER SIDE OF THIS PART (side without the pads)

Dowel pins

1. D0902137

Dowel Pin 1/8X7/16

MC MASTER 90145A481

2. D0902435

|  | Dowel Pin 1/8X1-1/8 MC MASTER 90145A882 |
| --- | --- |

Assembly procedure

1. Insert Item (23) (D1001737) Step Pin into Item (1) (D1000584) Small Actuator.

Note : Non pined face

| 1 | 23 |
| --- | --- |

2.Attach Item (5) (D0902155) Actuator Coil Bracket, to Item (23) (D1001737) Step Pin and Item (1) (D1000584) Small Actuator using Item (13) washer and Item (14) screw.

| 5 1 | Note : Look this face, connector on the right #4, 0.115 ID, 0.209 OD, 0.016 THK UCC WFV-04 SHCS 4-40 X 0.75 lg. MSC 05664073 Torque value : 5.2 in-lbs |
| --- | --- |

3.Attach Items (3) (D0902154) Left Thermal Bar and  Item (4) (D0902153) Right Thermal Bar, to Item (1) (D1000584) Small Actuator and Item (5) (D0902155) Actuator Coil Bracket using Item (13) washer and Items (14)/(15) screws.

| 4 3 | SHCS 4-40 X 0.25 lg. MSC 82272881 Torque value : 5.2 in-lbs |
| --- | --- |
| SHCS 4-40 X 0.75 lg. MSC 05664073 Torque value : 5.2 in-lbs | #4, 0.115 ID, 0.209 OD, 0.016 THK UCC WFV-04 |

4. Insert Item Dowel Pin into Item (1) (D0902749) Large Actuator.

SHCS 1/8 X 3/8 lg.

MC MASTER 90145A470

5. Attach Items (2) (D0902161) Small Actuator Mount, to Item (1) (D1000584) Small Actuator using Item (18) vent washer and Items (20) screw.

| Note : connector on the right |  |
| --- | --- |

SHCS 8-32, 1.50 lg. AG Plate

UCC C-824-NA

Torque value : 19.8 in-lbs

0.17 ID, 0.304 OD, 0.032 THK

UC COMPONENTS WFV 08

6. Tooling Bracket Assembly

Attach Part D0902137 Tooling Bracket (only the bottom), to Step Pin and Item (1) (D1000584) Small Actuator using vent washers and screws.

7. Loose the 2 shoulder screws of Item (1) (D1000584) Small Actuator.

8. Slide Pin Carrier (D0902434) into the Tooling Bar and Item (5) (D0902155).

|  | SHCS 8-32, 0.50 lg. HOLOKROME 78028 Torque value : 19.8 in-lbs |
| --- | --- |

9. Remove the 2 shoulder screws of Item (1) (D1000584) Small Actuator.

10. Add the two last screws and washers in the Tooling Bar

11. Remove the Pin Carrier (D0902434).

Gap Test : Go / No Go

- Check the gap between the coil and the magnet with shims on both sides.

Go : 0.085 inches

No Go : 0.115 inches