| AUTHOR: | CHECKED: | DATE | APPROVALS |
| --- | --- | --- | --- |
|  |  |  | DCN NO. | REV | DATE |
| G. Billingsley |  | 02-28-01 | E010032-00 |  | 02-28-01 |

Applicable Documents

LIGO-D000464-00-D	Substrate, Mode Cleaner Curved Mirror, 40M RSE Experiment
LIGO-D970534-00-D	Mirror Blank Material, Mode Cleaner Curved Mirror
LIGO-D970538-00-D	Mode Cleaner Curved Mirror Blank

Requirements

Physical Configuration

According to

LIGO- D000464		Substrate, Mode Cleaner Curved Mirror, 40M RSE Experiment

Fabricate from

LIGO-D970534		Mirror Blank Material, Mode Cleaner Curved Mirror

Part and Serial Number

The Serial number shall be per D000464 and of the format:
MCCM YY Where 
YY 		is incremental for each optic starting at 01

Registration Mark

Registration mark shall be etched, ground or sandblasted coincident with the registration mark drawn on the Blank within 5 mm.  The arrow orientation used on the Blank will be preserved if possible or changes reported in detail. Reference LIGO-D970538, Mode Cleaner Curved Mirror Blank.

Side and Bevel Polish

All sides and Bevels shall be polished from a five micrometer grit finish. These surfaces shall appear transparent with no gray, scuffs or scratches visible to the naked eye when viewed in normal room light against a black background.

Scratches and Point defects

There shall be no scratches, sleeks or point defects within the central 10 mm

The total area of scratches, sleeks and point defects within the central 30 mm diameter shall not exceed 500 square micrometers (width times length.)

The total area of scratches outside the central 30 mm diameter shall not exceed 5,000 square micrometers.

Minimum Point Defect

Point defects which have a maximum dimension of 5 micrometers are disregarded.

Inspection Method

1. The surface is examined visually by two observers independently. The examination is done against a dark background using a three-bundle fiber-optic illumination system of at least 200 W total power. A 100% inspection of the surface is carried out. Pits and scratches down to 2 micrometers in width can be detected using this method of inspection.  Any scratches that are detected will be measured using a calibrated eyepiece.

2. Further inspection will be done with a 6X eyeglass using the same illumination conditions, again with two observers. Sleeks down to 0.5 micrometers wide can be detected using this method. The surface will be scanned along one or two chords from center to edge, then at ten positions around the edge, and ten to fifteen positions near the center.

3. An inspection is then carried out with a dark field microscope with a similar sampling frequency as described in section 2.

Surface Figure, measured over the central 30 mm diameter

All specified quantities refer to the physical surface of the optic.

Surface 1: Spherical, concave.
Radius of curvature > 17.87 meters concave +0.25 meters, -0.35 meters 
Astigmatism:  < 100 nanometers (surface peak to valley)

Surface 2: Nominally flat.

(/4)

Surface Errors, Surface 1

All specified quantities refer to the physical surface of the optic.
The following root mean square standard deviation (rms) values are calculated from the phase maps which are to be provided with each optic.  rms is defined as the square root of the mean of the square of each pixel value.  Known bad pixels are excluded from this calculation.

Low Spatial Frequency Band:  < 4.3 cm -1
With piston, tip, tilt, power (best fit spherical surface) and astigmatism removed over the central 30 mm diameter aperture: 
rms  <5 nanometers

High Spatial Frequency Band:  4.3 - 7,500 cm -1
rms  < 0.4 nanometers 
 
Measured at the following locations:
1.  The center of the mirror substrate.
2.  Four positions equally spaced along the circumference of a centered, 30 mm diameter circle.

Table 1 Certification Data Requirements

| Specification | Test Method | Frequency of Inspection | Data Delivered |
| --- | --- | --- | --- |
| Physical Dimensions | Visual Inspection | 100% | Diameter, Thickness, Bevel dimension, Wedge angle. |
| Side and Bevel Polish | Visual Inspection | 100% | Inspection Report included with Certification |
| Scratches and Point defects | Visual Inspection | 100% | Hand sketch including scratch/pit dimensions |
| Surface Figure | Interferometry | 100% | Surface Map |
| Surface Errors - Low Spatial Frequency | Interferometry | 100% | Surface Map |
| Surface Errors - High Spatial Frequency | High resolution Surface Map | 100% | Surface maps for 3 central locations. Numerical values included with Certification |

Orientation:  For the purpose of full surface phase maps the substrate shall be oriented such that the point of minimum thickness shall be at the top center of the data.

Format:  All Data shall be delivered according to Table 1.  In addition to the hard copy an electronic data set of the phase maps shall be delivered in either ASCII or Vision.OPD format.