| AUTHOR: | CHECKED: | DATE | APPROVALS |
| --- | --- | --- | --- |
|  |  |  | DCN NO. | REV | DATE |
| G. Billingsley | Coyne | 6-16-06 | E060179-00 | A |  |

Applicable Documents

LIGO-D040431-B	Quad ETM Silica Test Mass
LIGO-D030265-A 	Fused Silica Blank LASTI Test Mass
LIGO-E030309-A	Fused Silica Blank, LASTI Test Mass, R&D

Requirements

Physical Configuration

Shape and Polish only according to
LIGO-D040431-B	Quad ETM Silica Test Mass, X-dimension 200.6 ± 0.2 mm

Fabricate from
LIGO-E030309-A	Fused Silica Blank, LASTI Test Mass, R&D
LIGO-D030265-A 	Fused Silica Blank LASTI Test Mass

Registration Marks

Registration marks shall be etched, ground or sandblasted and located per LIGO-D040431

Serial Number

Serial Number “LEM 1” shall be etched, ground or sandblasted and located per LIGO-D040431

Side and Bevel Polish

Sides and Bevels shall be polished using a progression of smaller grit sizes.  The last step before final polish shall be equal to or less than a five micrometer grit finish. These surfaces shall appear transparent with no grey, scuffs or scratches visible to the naked eye when viewed in normal room light against a black background.

Surface 1 and 2 as defined in LIGO-D040431

Surfaces 1 and 2 be polished using a progression of smaller grit sizes.  The last step before final polish shall be equal to or less than a 15 micrometer grit finish.

Mounting Flat Figure, measured over the Bond Area per D040431, 2 places

Flat:  Peak to valley maximum deviation over the bond area, with tilt subtracted: < 60 nm
Surface rms error over the bond area: < 10 nm

Inspections

| Specification | Test Method | Data Delivered |
| --- | --- | --- |
| Outside Diameter | Physical inspection | Measured dimension |
| Wedge | Physical inspection | Measured dimension |
| Thickness | Physical inspection | Measured X- dimension per D040431-A |
| Bond area flatness | Interferometry | Surface map for each bond area |