# LASER INTERFEROMETER GRAVITATIONAL WAVE OBSERVATORY

# -LIGO-

# CALIFORNIA INSTITUTE OF TECHNOLOGY

# MASSACHUSETTS INSTITUTE OF TECHNOLOGY

| Document Type Technical Report | DCC Number LIGO-T010036-00-R | Date March 16, 2001 |
| --- | --- | --- |
| HYSTERESIS REPORT OF THE TAMA-SAS FILTERS |
| Jose Edwin Ugas, Riccardo DeSalvo, Akiteru Takamori, Virginio Sannibale, Kenji Numata, Tatsuo Yoda |

Distribution of this draft: NSF reviewers, LIGO scientists

This is an internal working note of the LIGO Laboratory

California Institute of Technology	Massachusetts Institute of Technology

LIGO Project – MS 18-33	LIGO Project – MS 20B-145

Pasadena, CA 91125	Cambridge, MA 01239

Phone (626) 395-2129	Phone (617) 253-4824

Fax (626) 304-9834	Fax (617) 253-7014

E-mail: info@ligo.caltech.edu	E-mail: info@ligo.mit.edu

www:  http://www.ligo.caltech.edu/

### HYSTERESIS REPORT

OF THE TAMA-SAS FILTERS.

Evidence of Hysteresis in the TAMA-SAS Filters.

While tuning the MGAS blades, in particular the blades located in Filter 0, it was repeatedly observed, on top of the well expected Bi-stable equilibrium condition in the measurements when K< 0, that there was some very small scale bi-stability.  The peculiar aspect of this observation is that, in all cases, the difference between the two equilibrium points was very small, about 1mm.  Furthermore, all measurements fell along the customary position versus frequency curve without deviations.  Despite this fact, two separate and repeatable equilibrium positions can be obtained if the working point is approached slowly from far above or far below, without allowing large oscillations.  This small-scale anomaly was rapidly identified as Hysteresis.

When measuring the MGAS blades utilized at SAS, one can certify that the blades are in Bi-stable equilibrium condition (K<0) when a “gap” is observed between two equilibrium points pertaining to the same payload.  Sometimes, bi-stable conditions might not be quite obvious at first.  However, the distance between the equilibrium points increase until a large gap is observed, this indicates that the blades were in bi-stable condition all along, due to the particular compression.  And, upon plotting the measurements taken for various payloads, the measured points of a position versus frequency graph form two separate convex curves leading to the gap, at which point both curves begin to open outwards in a hyperbolic form.  In the case of normal stability, when the blades are said to be in Mono-stable equilibrium condition, this same curve forms a single concave curve, which is parabolic in nature with its minimum point representing the minimum resonant frequency of the MGAS blade.

In this experiment, what we initially thought was bi-stable condition was observed in many measurements, all falling on the same concave curve (parabola), a characteristic of a Filter mono-stable operation.  Upon more careful considerations of the mechanical structure of the attachments on the Filter, it was realized that the problem could be traced to Hysteresis. The main suspects were the blades and the clamp-wedge-filter attachment system.  The hysteresis is totally obvious if the data is plotted as a function of payload versus frequency or payload versus equilibrium position, graphs that are rarely used in tuning MGAS Filters for low frequencies.

In principle, any mechanical system will have some type of hysteresis, this should also happen in the MGAS blades themselves.  Note that the system could, in principle, be at the same time in bi-stable condition and have hysteresis along the two hyperbolic arms.  In fact, upon analyzing the measurements, we found evidence of both types of phenomena.

Following the hysteresis measurements, we performed hardness measurements on all the components of the attachments.  We first suspected soft blades due perhaps to a bad precipitation hardening process.  A hardness test (Rockwell) showed that the blades had the expected hardness of 51 to 54, a hardness level which in previous measurements of the blade’s maraging steel proved it to be virtually free from hysteresis.  Instead, it was found that the wedge and clamp had a hardness of about 4 to 5 Rockwell.  It became obvious that this extreme mismatch was the most possible cause of the observed hysteresis.

The MGAS tuning experiment.

Figure 1:  Pictures of the experimental set-up.

A note on the experimental procedure.

The three main parameters of interest are radial compression, load, vertical equilibrium position, and vertical resonant frequency.  With these quantities, the optimal condition of the system can be achieved by a process called tuning, both of the radial compression, the load, and the vertical equilibrium position (the working point of the filter). The compression parameter, like load, is determined and adjusted directly; the position and frequency parameters are consequences, which are measured indirectly.  For each compression level, vertical equilibrium position and vertical resonant frequency are recorded for many different load levels.  The radial compression is adjusted until the required vertical resonant frequency of the MGAS is obtained at the filter working point (the minimum in the vertical position versus vertical resonant frequency curve).  There are, of course, other “hidden variables” that affect the frequency measured and also the position; one example is damping by friction.  Note that the most crucial part of this experiment is tuning the radial compression, since the vertical resonant frequency and the vertical equilibrium position can both be tuned by adding or subtracting suitable amounts of load.  The radial compression has to be in the right level, so as to produce the optimal condition of the attenuation filter without any bi-stable conditions.

The standard load consists of:

1.)	The wooden cage, including the iron threaded rods, bolts, washers, etc.

3.)	The iron disks

4.)	A device, made of aluminum, used to measure the position.

The total weight of the standard load for Filter 0 is 95 Kg.

Experimental Data:

First Set of measurements.

# TALBE 1

Load[Kg]	Position[mm]		frequency[mHz]

98.020		100.00			711.20

98.148		99.000			660.50

98.148		98.900			654.00

98.277		97.500			605.00

98.277		97.100			592.10

98.406		96.000			545.60

98.406		95.500			519.50

98.534		93.900			457.90

98.534		92.500			419.10

98.662		90.400			341.40

98.662 	86.500			300.90

98.790		82.900			322.60

98.790 	80.600			365.00

98.919		78.400			461.50

98.919		76.400			491.20

Figure 2:  Load vs. Vertical Resonant Frequency.

This graph shows how for each load, there is two different frequency points.

Figure 3:  Vertical Position vs. Vertical Resonant frequency.

Notice how the distance in position in each measurement, the two points corresponding to the same load, increases as the system attains it’s minimum resonant frequency.

Then, about 2 hours later, the same measurements were repeated for the same MGAS blade with the same initial conditions.  As one can see, different data was obtained:

Second Set of measurements.

TABLE 2

Load[Kg]	Position[mm]	frequency[mHz]

98.020		99.500		699.80

98.020		99.250		679.30

98.148		98.500		641.40

98.148		98.000		620.00

98.277		96.900		581.10

98.277		96.250		556.50

98.406		95.250		518.90

98.406		94.250		478.50

98.534		92.800		422.80

98.534		90.800		369.50

98.662		88.500		311.40

98.662		83.500		309.90

98.790		80.600		360.60

98.790		79.000		414.20

98.919		77.500		468.80

98.919		76.600		492.90

Figure 4:  Second measurement of Load vs. Vertical Resonant Frequency.

Figure 5:  Second measurement of Vertical Position vs. Vertical Resonant frequency.

The following are graphs corresponding to the two sets of measurements, a total of 4 individual points per each independent parameter.

Figure 6:  A comparison of the data of  Figure 2 and Figure 4.

Figure 7:  A comparison of the data in Figure 3 and Figure 5.

Figure 8:  A comparison of the data obtained from the two tables, for Load vs. Vertical Position.

As can be seen, in the second set of measurements, a lower minimum vertical resonant frequency was obtained.  It is possible that the measurements differ because the blades, and therefore the attachments, were subject to longer periods of stress and thereby altering the mechanical configuration, causing a different level of hysteresis.

Previous evidence of Hysteresis in SAS and TAMA/SAS Filters.

In previous towers, for example the first prototype “LIGO SAS Tower”, hysteresis was never observed directly.  One reason why the hysteresis was not observed before is that, due to the nature of this particular cause of hysteresis, it is mostly observed at low frequencies and with very small amplitude of oscillation.  That is to say, the mechanical configuration of the attachments and, the strength of the materials which compose it, do not allow for measurable hysteresis.  In fact, one needs to know how to identify this type of hysteresis on a system like the SAS towers, specially on the data plots, since it could be passed without noticed.

We went back to the old LIGO-SAS tower and found no visible mechanical hysteresis.  The remaining wedges of this tower were measured for hardness and were found to have a Rockwell hardness above 20.

Figure 9:  An underneath photograph of Filter Zero that shows the MGAS blades with the attachments used, in the TAMA-SAS attenuation towers.

## Conclusion

We have now concluded that the problem lies principally with the wedge’s hardness in the TAMA-SAS towers.  Harder wedges and clamps are under construction to cross check this solution on a spare MGAS Filter.  Further analysis will be done in this system to minimize the possibility of hysteresis, and it is hoped that this observation will shed light on future similar suspension systems.