*LIGO Laboratory / LIGO Scientific Collaboration*

LIGO-T040220-D-C *LIGO* 6 March 2006

Test Results for C30642 LSC Diode Elements

Richard Abbott

Distribution of this document:

LIGO Science Collaboration

This is an internal working note

of the LIGO Project.

|  |  |
| --- | --- |
| **California Institute of Technology****LIGO Project – MS 18-34****1200 E. California Blvd.****Pasadena, CA 91125**Phone (626) 395-2129Fax (626) 304-9834E-mail: info@ligo.caltech.edu | **Massachusetts Institute of Technology****LIGO Project – NW17-161****175 Albany St****Cambridge, MA 02139**Phone (617) 253-4824Fax (617) 253-7014E-mail: info@ligo.mit.edu |
| **LIGO Hanford Observatory****P.O. Box 1970****Mail Stop S9-02****Richland WA 99352**Phone 509-372-8106Fax 509-372-8137 | **LIGO Livingston Observatory****P.O. Box 940****Livingston, LA 70754**Phone 225-686-3100Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

Measurements were taken to quantify series capacitance and series resistance for the C30642 InGaAs, 2mm photodiodes used in the LSC RF Photodetector. Data taken in January of 2003 is included for comparison.

All measurements were taken using an HP4195A and associated Impedance Test Adapter. The data was taken at 24.5 MHz with 7 volts reverse bias applied

# Diodes from LHO and LLO, November 2004

|  |  |  |
| --- | --- | --- |
| **Serial Number** | **Series Capacitance (pF)** | **Series Resistance, Ohms** |
| A6578 | 105 | 11.9 |
| A6579 | 105 | 11.4 |
| A6580 | 105 | 11.6 |
| A6581 | 106 | 10.8 |
| A6582 | 105 | 11.8 |
| A6584 | 108 | 10.9 |
| A6585 | 106 | 11.3 |
| A6590 | 105 | 11.5 |
| A6591 | 105 | 11.7 |
| A6592 | 105 | 11.6 |
| A6594 | 107 | 10.6 |
| A6595 | 106 | 11.3 |
| A6596 | 105 | 11.4 |
| A6598 | 107 | 10.9 |
| A6603 | 107 | 10.6 |
| A6604 | 106 | 11.5 |
| A6605 | 107 | 10.7 |
| A6606 | 105 | 11.9 |
| A6607 | 106 | 11.1 |
|  |  |  |
| **Average** | 105.8421053 | 11.28947368 |
| **Std. Dev.** | 0.932633955 | 0.420394222 |

# C30642 Diodes from LHO 10 December 2004

|  |  |  |
| --- | --- | --- |
| **Serial Number** | **Series Capacitance (pF)** | **Series Resistance, Ohms** |
| A6721 | 125.7 | 14.4 |
| A6722 | 111.7 | 12.1 |
| A6723 | 125.5 | 14.3 |
| A6725 | 106.8 | 12.4 |
| A6714 | 130.6 | 13.2 |
| A6713 | 125.8 | 13.8 |
| A6712 | 105.3 | 11.8 |
| A6711 | 105.6 | 11.8 |
| A6719 | 107 | 11.1 |
|  |  |  |
| **Average** | 116 | 12.76666667 |
| **Std. Dev** | 10.00177762 | 1.130388331 |

# C30642 Diodes from CIT measured in January 2003

|  |  |  |
| --- | --- | --- |
| **Serial Number** | **Series Capacitance (pF)** | **Series Resistance, Ohms** |
| Cut off | 137 | 7.6 |
| Cut off | 138 | 7.4 |
| Cut off | 69 | 15.4 |
| Cut off | 138 | 7.5 |
| Cut off | 138 | 7.6 |
| Cut off | 138 | 7.4 |
| Cut off | 137 | 7.3 |
| A5890 | 139 | 7.5 |
| A5898 | 139 | 9.1 |
| A2486 | 95 | 11.6 |
| A5897 | 139 | 7.4 |
| Cut off | 136 | 8.8 |
| A5910 | 138 | 7.6 |
| A5947 | 137 | 7.4 |
|  |  |  |
| **Average** | 129.8571429 | 8.542857143 |
| **Std. Dev** | 20.16311039 | 2.209626434 |

# Diodes from the Wilson House shop measured 5 January 2005

|  |  |  |
| --- | --- | --- |
| **Serial Number** | **Series Capacitance (pF)** | **Series Resistance, Ohms** |
| A5437 | 115 | 12.7 |
| A5414 | 118 | 13.2 |
| A5434 | 118 | 12.8 |
| A5412 | 118 | 12.5 |
| A5446 | 118 | 13.1 |
| A5423 | 117 | 13 |
| A5395 | 117 | 13.3 |
| A5427 | 117 | 13.1 |
| A5418 | 118 | 12.7 |
| A5464 | 115 | 13.3 |
| A6378 | 127.8 | 9.8 |
| A6379 | 133.7 | 9.6 |
| A6380 | 133.7 | 8.5 |
| A6381 | 128.7 | 9.9 |
| A6366 | 127.7 | 10 |
| A6367 | 128.6 | 9.8 |
| A6368 | 136 | 9.5 |
| A6369 | 132.7 | 9.3 |
| A6354 | 131.9 | 9.3 |
| A6355 | 133.2 | 9.7 |
| A6356 | 131.3 | 7.4 |
| A6357 | 131.3 | 7.8 |
|  |  |  |
| **Average** | 124.8909091 | 10.92272727 |
| **Std. Dev** | 7.399072982 | 1.972523031 |

# Diodes from LLO 12 August 2005

|  |  |  |
| --- | --- | --- |
| **Serial Number** | **Series Capacitance (pF)** | **Series Resistance, Ohms** |
| A6717 | 125 | 13.5 |
| A6726 | 104.8 | 12.3 |
| A6732 | 104.5 | 12.5 |
| A6733 | 105.4 | 12.5 |
| A6737 | 104.7 | 11.9 |
| A6330 | 127.8 | 8.9 |
| A6328 | 128.6 | 8.9 |
| A6370 | 133 | 9.7 |

# Batch of 40 Diodes received 10 October 2005

|  |  |  |  |
| --- | --- | --- | --- |
| **Serial Number** | **Series Capacitance (pF)** | **Series Resistance, Ohms** | **Dark Leakage Current (nA) 7V Bias** |
| A6835 |  |  | <10nA |
| A6836 |  |  | <10nA |
| A6837 |  |  | <10nA |
| A6838 |  |  | 260nA |
| A6839 |  |  | 40nA |
| A6840 |  |  | <10nA |
| A6841 |  |  | <10nA |
| A6843 |  |  | <10nA |
| A6844 |  |  | <10nA |
| A6848 |  |  | 940nA |
| A6849 |  |  | <10nA |
| A6850 |  |  | <10nA |
| A6851 |  |  | <10nA |
| A6851 |  |  | <10nA |
| A6852 |  |  | <10nA |
| A6854 |  |  | 590nA |
| A6856 |  |  | 60nA |
| A6890 |  |  | <10nA |
| A6891 |  |  | <10nA |
| A6892 |  |  | <10nA |
| A6893 |  |  | <10nA |
| A6895 |  |  | <10nA |
| A6896 |  |  | <10nA |
| A6897 |  |  | <10nA |
| A6898 |  |  | <10nA |
| A6899 |  |  | 10nA |
| A6900 |  |  | <10nA |
| A6901 |  |  | <10nA |
| A6902 |  |  | <10nA |
| A6903 |  |  | <10nA |
| A6904 |  |  | <10nA |
| A6905 |  |  | <10nA |
| A6906 |  |  | <10nA |
| A6908 |  |  | <10nA |
| A6909 |  |  | <10nA |
| A6910 |  |  | <10nA |
| A6911 |  |  | <10nA |
| A6913 |  |  | <10nA |
| A6914 |  |  | <10nA |
| A6915 |  |  | <10nA |

# GAP2000 Diodes measured in January 2003. Units have no serial numbers

|  |  |
| --- | --- |
| Capacitance | Resistance |
| 193 | 9.5 |
| 211 | 8.8 |
| 209 | 9.1 |
| 202 | 8.9 |
| 205 | 8.9 |
| 200 | 9 |
| 200 | 8.6 |
| 191 | 8.4 |
| 197 | 8.5 |
| 201 | 8.9 |
| 215 | 8.7 |
| 192 | 8.4 |
| 192 | 8.5 |
| 197 | 8.5 |
| 201 | 8.9 |
| 191 | 8.5 |
| 195 | 8.4 |
| 192 | 8.2 |
| 189 | 8.3 |
| 189 | 8.3 |
| 195 | 8.5 |
| 196 | 8.5 |
| 191 | 8.4 |
| 199 | 8.5 |
| 197 | 8.7 |
| 197 | 8.5 |
| 202 | 8.7 |
| 190 | 8.4 |
| 189 | 8.3 |
| 192 | 8.4 |
| 201 | 8.6 |
| 211 | 8.9 |
| 197 | 8.5 |
| 200 | 8.9 |
| 192 | 8.5 |
| 206 | 8.7 |
| 199 | 8.9 |
| 215 | 8.7 |
| 214 | 8.6 |
| 198 | 8.8 |
| 196 | 8.8 |
| 191 | 8.6 |
| 195 | 8.7 |
| 213 | 9.2 |
| 200 | 9.3 |
| 213 | 9.1 |
| **Serial Number** | **Measured Voltage (mV)** | **Calculated Current (nA)** |
| 6838 | 96 | 1056 |
| 2039 | 25 | 275 |
| 2293 | 7.5 | 83 |
| 2041 | 1003 | 11034 |
| 2044 | 4 | 44 |
| 2047 | 110 | 1210 |
| 2045 | 232 | 2552 |
| 2042 | 2 | 22 |
| 2046 | 2 | 22 |
| 6508 | 0.6 | 7 |
| 6482 | 0.7 | 8 |
| 6658 | 1.2 | 13 |
| 6511 | 0.9 | 10 |
| 6484 | 0.7 | 8 |
| 6660 | 1.1 | 12 |
| 6656 | 0.8 | 9 |
| 6509 | 0.9 | 10 |
| 6657 | 1.2 | 13 |
| 6652 | 1.1 | 12 |
| 6510 | 1 | 11 |
| 6533 | 0.8 | 9 |
| 6530 | 0.8 | 9 |
| 6655 | 1 | 11 |
| 6534 | 0.9 | 10 |
| 6106 | 1.5 | 17 |
| 6126 | 1 | 11 |
| 6141 | 7.1 | 78 |
| 6107 | 1.4 | 15 |
| 6094 | 1.5 | 17 |
| 6125 | 180 | 1980 |
| 6083 | 3 | 33 |
| 6081 | 1 | 11 |
| 6093 | 2.5 | 28 |