Author: Robert Schofield

Date: March 19, 2004

Subject: Magnetic Field Measurement

DCC Number: T050087-00-W

Hi Janeen,

This is the email I mentioned that is a summary of the presentation

I gave in Dec. It was a response to an email that Dennis sent me, but,

I could not find it in my sent-mail file so it may not have made it

to Dennis, hence I am also sending it to Dennis.

Let me give you a summary of the short presentation that

I made in Dec.:

1) Original scope: determine if generated field is of some

concern or of no concern. If of some concern, schedule

more detailed measurements. So, if we stick to the original

scope, we should schedule more measurements and not rely

heavily on what we got done in 2 days.

2) We generated fields with each of 4 actuators (1 of the

3 sets).

3) We measured V across one ohm resistor in series with actuator in order

to set current through actuator to 1A (stage 0-1) or 0.3A (stage 1-2)

Two different frequencies were tried, 380 Hz and 11 Hz.

4) I scanned the fields above the surface of the table

(using a Bartington magnetometer set on foam

that raised it about 10 cm above the surface of the table) and recorded

only the highest value that I found. These values were then scaled

according to the expected actuator currents that I obtained from

Brian L. Thus, the highest measured field for stage 0-1 horizontal

actuator driven with 1A at 11 Hz was about 1uT. Scaled down

by 20, the expected worst field is about 50pT.

This was the largest expected field that I obtained, and the largest

relative to the ambient background, which is generally in the tens of

pT at this frequency.

5) The scale of spatial field variations was smaller than the minimum

distance between 2 magnetometers, so I obtained gradients from differences

between sequential field measurements using a single magnetometer at

slightly different positions. A quick survey this way is impossible, so I

assumed that the worst gradients would be found where the fields were

greatest, and made gradient measurements only for the 0-1 horizontal

actuator (at both 380 and 11 Hz). After scaling my measured values by

B.L.'s expected currents, I got a worst expected gradient of 100pT/m at 11

Hz. This is about 10 times the ambient field gradient backgrounds that I

have measured at the sites.

6) The gradients above the actuators were so large that the optic

support structure- produced gradients did not overwhelm them. Thus

the optic support structures did not increase the estimated maximum

expected gradient.

7) An approx 8.5x11 inch sheet of 0.004 inch mu-metal reduced fields

at the level of the optic magnets by a factor of about 5 either when it

was inserted into a space above the 0-1 horizontal actuator, or when it

was placed on the bottom of an optic support structure.

8) Things to worry about:

a) I noticed that the table seemed to produce fields at about 59 Hz that

were one or two orders of magnitude above ambient background 10 cm above

the table. My interpretation was that this field (that was 1 or 2 orders

of magnitude smaller if I moved a meter or so from the table) was produced

by motor-induced table motion relative to the earths magnetic field

(thereby modulating the earths field at 59 Hz). My concern is that there

are large structures in the isolation system that can deform the earths

magnetic field and these structures are close to, and moving relative to,

the magnets on the pendula. For example, in the HAM chamber,

the magnets on the optics will be only

decimeters away from stage 1 of the active isolation system. Stage 1

will be moving with the ground (albeit with some attenuation), while

the optic (along with the earths field) will be in a much more inertial

frame. Since stage 1 will be distorting the earths magnetic field and

moving, it will be modulating the earths field at the position of the

magnets. I dont know how important this "short" will be, but I think it is

worth considering. Is there a possibility of a similar seismic "short"

associated with the static charge actuators?

b) I have not calculated the expected actuator fields at the resonant

frequencies of the pendula.

Robert