LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T070010-C	LIGO	01/22/2007

10-W Laser Head Interface Test Procedure

Peter King

Distribution of this document:

LIGO Science Collaboration

This is an internal working note

of the LIGO Project.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW17-161 175 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

This document outlines the tests to be performed on the 10-W Laser Head Interface board D070001-0A-C.  This board is a copy of the Lightwave Electronics 4-2436 Rev. B. LIGO Head Interface board.

The Head Interface board is the interface between the 10-W laser power supply and the NPRO master oscillator and power amplifier.  Signals from the two EOT-3020 InGaAs photodetectors that monitor the power of the NPRO and 10-W laser, called 126MON and AMPMON respectively, are processed by the Head Interface board and are output on the 25-pin D-sub connector.  The signals from the temperature sensors located under the amplifier are passed through the Head Interface board to the 10-W laser power supply and the Neslab RTE-140M refrigerated bath.  The other function of the Head Interface board is to supply power to the shutter and for the current shunt.

# Visual Check

Inspect the board for soldering mistakes.  Check that shunts are installed at J4, J5, J6, J7, and that four shunts are either installed at JPZT1, JPZT2, JCM1, JSL1 or in jumper storage.

# Required Equipment

The equipment required to test the Head Interface board is listed below:

- Dual DC power supply capable of 0-30 V, 0-1 A.

- Digital multimeter (DMM).

- Martel current calibrator or equivalent device.

# On-board Power

Through the 25-pin D-sub connector, apply (+24.0 ( 0.5) V to pin 24, ground pins 2, 13 and 25 and, (-12.0 ( 0.5) V to pin 22.  Using a DMM, check the voltages at the following locations and record the measured values in the tables below.

| Location | Nominal Value | Measured Value | Pass/Fail |
| --- | --- | --- | --- |
| Q1 pin 3 | (+5.00 ( 0.25) V |  |  |

| Location | Nominal Value | Measured Value | Pass/Fail |
| --- | --- | --- | --- |
| U4 pin 3 | (+12.0 ( 0.5) V |  |  |

| Location | Nominal Value | Measured Value | Pass/Fail |
| --- | --- | --- | --- |
| J2 pin 1 | (+24.0 ( 0.5) V |  |  |
| J2 pin 4 | ground (0 V) |  |  |

| Location | Nominal Value | Measured Value | Pass/Fail |
| --- | --- | --- | --- |
| J8 pin 1 | ground (0 V) |  |  |
| J8 pin 2 | (+5.00 ( 0.25) V |  |  |

| Location | Nominal Value | Measured Value | Pass/Fail |
| --- | --- | --- | --- |
| JS1 pin 1 | ground (0 V) |  |  |
| JS1 pin 2 | ground (0 V) |  |  |
| JS1 pin 12 | (-12.0 ( 0.5) V |  |  |
| JS1 pin 14 | (+12.0 ( 0.5) V |  |  |
| JS1 pin 15 | (+12.0 ( 0.5) V |  |  |
| JS1 pin 16 | ground (0 V) |  |  |

| Location | Nominal Value | Measured Value | Pass/Fail |
| --- | --- | --- | --- |
| U1 pin 4 | (+12.0 ( 0.5) V |  |  |
| U1 pin 11 | (-12.0 ( 0.5) V |  |  |

| Location | Nominal Value | Measured Value | Pass/Fail |
| --- | --- | --- | --- |
| U2 pin 4 | (-12.0 ( 0.5) V |  |  |
| U2 pin 7 | (+12.0 ( 0.5) V |  |  |

| Location | Nominal Value | Measured Value | Pass/Fail |
| --- | --- | --- | --- |
| U3 pin 4 | (-12.0 ( 0.5) V |  |  |
| U3 pin 7 | (+12.0 ( 0.5) V |  |  |

# Circuit Continuity Checks

Using the continuity function of the DMM, check the following connections for continuity.

| From | To | Pass/Fail |
| --- | --- | --- |
| J14 pin 1 | J8 pin 7 |  |
| J14 pin 4 | J8 pin 8 |  |
| J14 pin 7 | J8 pin 9 |  |
| J14 pin 9 | J8 pin 10 |  |

| From | To | Pass/Fail |
| --- | --- | --- |
| J13 pin 7 | J8 pin 3 |  |
| J13 pin 8 | J8 pin 5 |  |
| J13 pin 19 | J8 pin 4 |  |
| J13 pin 20 | J8 pin 6 |  |

Making sure that no shunts are installed at JPZT1, JPZT2, JCM1 and JSL1, check the following connections.

| From | To | Pass/Fail |
| --- | --- | --- |
| J10 pin 2 | J1 pin 9 |  |

| From | To | Pass/Fail |
| --- | --- | --- |
| J11 pin 1 | J1 pin 15 |  |
| J11 pin 2 | J1 pin 14 |  |

| From | To | Pass/Fail |
| --- | --- | --- |
| J12 pin 2 | J1 pin 11 |  |

Install shunts on JPZT1, JPZT2, JCM1 and JSL1.  Check the following connections.

| From | To | Pass/Fail |
| --- | --- | --- |
| J15 pin 1 | J1 pin 1 |  |
| J15 pin 2 | J1 pin 2 |  |
| J15 pin 3 | J1 pin 3 |  |
| J15 pin 4 | J1 pin 4 |  |
| J15 pin 5 | J1 pin 5 |  |
| J15 pin 6 | J1 pin 6 |  |
| J15 pin 7 | J1 pin 7 |  |
| J15 pin 8 | J1 pin 8 |  |
| J15 pin 9 | J1 pin 9 |  |
| J15 pin 10 | J1 pin 10 |  |
| J15 pin 11 | J1 pin 11 |  |
| J15 pin 12 | J1 pin 12 |  |
| J15 pin 13 | J1 pin 13 |  |
| J15 pin 14 | J1 pin 14 |  |
| J15 pin 15 | J1 pin 15 |  |

With the shunts installed, the following connections should be open circuit.

| From | To | Pass/Fail |
| --- | --- | --- |
| J15 pin 1 | J15 pin 2 |  |
| J15 pin 1 | J15 pin 9 |  |
| J15 pin 2 | J15 pin 3 |  |
| J15 pin 2 | J15 pin 9 |  |
| J15 pin 2 | J15 pin 10 |  |
| J15 pin 3 | J15 pin 4 |  |
| J15 pin 3 | J15 pin 10 |  |
| J15 pin 3 | J14 pin 11 |  |
| J15 pin 4 | J15 pin 5 |  |
| J15 pin 4 | J15 pin 11 |  |
| J15 pin 4 | J15 pin 12 |  |
| J15 pin 5 | J15 pin 6 |  |
| J15 pin 5 | J15 pin 12 |  |
| J15 pin 5 | J15 pin 13 |  |
| J15 pin 6 | J15 pin 7 |  |
| J15 pin 6 | J15 pin 13 |  |
| J15 pin 6 | J15 pin 14 |  |
| J15 pin 7 | J15 pin 8 |  |
| J15 pin 7 | J15 pin 14 |  |
| J15 pin 7 | J15 pin 15 |  |
| J15 pin 8 | J15 pin 15 |  |

# Laser Power Monitors

A pair of ET-3020, 3 mm diameter InGaAs photodetectors are used to monitor the NPRO power and the 10-W laser power.  These have a response of 0.8 A/W and typically in a new 10-W laser about 3 mW falls incident on each photodetector, resulting in a photocurrent of 2.4 mA.

If using a Martel current calibrator, set up the calibrator to supply 2.4 mA.  Apply the Martel output to J3 and measure the voltage at TP1.  Adjust the trimming potentiometer P1 so that the voltage at TP1 is (8.0 ( 0.1) V.  Record the final voltage at TP1 in the table below.

Now apply the Martel output to J9 and measure the voltage at TP2.  Adjust the trimming potentiometer P2 so that the voltage at TP2 is (8.0 ( 0.1) V.  Record the final voltage at TP2 in the table below.

| Location | Nominal Value | Measured Value | Pass/Fail |
| --- | --- | --- | --- |
| TP1 | (8.0 ( 0.1) V |  |  |
| TP2 | (8.0 ( 0.1) V |  |  |