LIGO Laboratory / LIGO Scientific Collaboration

LIGO- T1100113-v3	LIGO	10/3/11

Mitigation Of Wide Angle Scatter From Hr Surface Of The Test Mass with the Arm Cavity Baffle

Michael Smith

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

Table of Contents

1	INTRODUCTION	6

2	SCATTERED LIGHT NOISE THEORY	6

3	WIDE ANGLE SCATTERING SURFACES	7

3.1	Arm Cavity Baffle Wide Angle Box Structure	7

3.2	Wide Angle Scattered Light from the Quad SUS Structure	13

4	RESULTS OF WIDE ANGLE SCATTER NOISE CALCULATIONS	14

4.1	ZEMAX Lambertian Scattering Data	15

4.2	Scattered Power Calculations	16

4.2.1	Scattered Power Parameters	16

4.2.2	Incident Power	17

4.2.3	Power Scattered into IFO Mode	18

4.2.4	Scattered Light Displacement Noise	22

4.2.4.1	SUS Parts	22

4.2.4.2	Arm Cavity Back and Wide Angle Box	23

4.2.4.3	BSC Chamber Walls	25

4.2.4.4	Spoolpiece and Manifold	26

5	INTERFACES	27

Table of Tables

Table 1: Wide Angle Scattered Light from H1 ITMX,  ZEMAX Ray Trace	16

Table 2:  Wide Angle Scattered Power	22

Table of Figures

Figure 1: ACB with Wide Angle Extendable Top, Bottom, and Sides	8

Figure 2: Extender Box Assembly	9

Figure 4: Removable Extender baffle	10

Figure 4: Removable Fixed Baffle	11

Figure 5: Captured plate with tapped holes	12

Figure 6: Lower Quad SUS Structure	13

Figure 7: ZEMAX Lambertian Scatter Ray Trace from H1 ITMX HR, ACB with Wide Angle Baffle Sides, top view	14

Figure 8: ZEMAX Lambertian Scatter Ray Trace from H1 ITMX HR, ACB with Wide Angle Baffle Sides, side view	15

Figure 9: SUS Wide Angle Scatter	23

Figure 10: Wide Angle Displacement Noise from Various Parts of the ACB	25

Figure 11: Total Components of TM Wide Angle Scattering	27

Abstract:

The wide angle scatter from the test mass HR surface can be totally mitigated by placing an extendable box structure on the back side of the Arm Cavity Baffle that catches most of the scattered light.

The revised ACB design used oxidized polished stainless steel instead of porcelainized steel.

# INTRODUCTION

Defects or dust on the test mass (TM) cavity mirrors scatter power into wide angles ; see T070089-02, Wide Angle Scatter from LIGO Arm Cavities. The angular scatter distribution is presumed to be Lambertian (i.e. “flat”), such that the probability of a cavity photon scattering to angle  off axis is given by  where  is a constant, nominally the “Lambertian reflectance.” This is reasonable for small point defects, and approximately consistent with Kells’ measurements out to . The value for  has been estimated by Hiro Yamamoto in G070240 to be 10 ppm of loss for wide-angle (point-defect) scatter.

This wide angle scatter can be totally mitigated by placing an extendable box structure on the back side of the Arm Cavity Baffle that catches all of the scattered light from the TM HR surface.

# SCATTERED LIGHT NOISE THEORY

The wide angle scattered light from the TM that hits an adjacent surface in the chamber will scatter from that surface back onto the TM and then re-scatter from the TM into the mode of the interferometer (IFO).

The power scattered by the adjacent surface onto the TM that will re-scatter into the IFO mode is given by

where Pinc is the power incident on the adjacent scattering surface, BRDFs is the scattering probability per solid angle of the adjacent surface, wifo is the mode size of the arm cavity mode, and L is the distance from the adjacent surface to the TM.

The power re-scattered into the IFO by the TM is proportional to the power incident on the TM, to the wide-angle scatter loss fraction, to the cosine of the incident angle on the TM, and to the solid angle of the interferometer mode.

By combining these two equations we get the following:

The displacement noise, in m/rtHz, caused by scatter from the moving scattering surface is proportional to the DARM transfer function (see T060073 ), to the square root of the scattered power relative to the PSL power, and to the seismic motion spectrum of the scattering surface.

# WIDE ANGLE SCATTERING SURFACES

## Arm Cavity Baffle Wide Angle Box Structure

A new Arm Cavity Baffle concept that completely mitigates the wide angle scattered light displacement noise from the TM is shown in Figure 1. The extendable sides, and top and bottom of the box structure can be moved as close as practicable to the TM SUS structure to capture most of the wide angle scattered light from the TM.

The wide angle box structure can be hinged out of the way for ease of access to the HR surface of the TM. In addition, the top, bottom, and sides of the box structure are removable to allow further access to the TM in the quad SUS for maintenance.

This design will eliminate the need for the Wide Angle Plate Baffles mounted to the BSC chamber walls.

Figure 1: ACB with Wide Angle Extendable Top, Bottom, and Sides

QUESTIONS
1) You mention extend-able sides, need to see detailed layout of new design wrt each ETM / ITM? (i.e. in all chambers)

The wide-angle box baffle structure is comprised of a fixed set of top and bottom shelves, and side walls that mount to the back of the arm cavity baffle; and a sliding extender box structure that fits over the fixed portion and is extended to bring the box structure as close as practicable to each TM suspension installation, as shown in Figure 2. The dimensions of the extended wide angle box for each of the TM installations are shown in D0901376_AdLIGO_AOS_SLC ARM Cavity Baffle Final Assy-v2 ; see sheets 2, 3, 4, and 5.

Figure 2: Extender Box Assembly

2) Are removable side sections removed / replaced from the outside?
The extender box and the fixed blue box can be disassembled from the main arm cavity baffle while it is installed inside the BSC chamber, as shown in Figure 4 and Figure 5.

Figure 4: Removable Extender baffle

Figure 4: Removable Fixed Baffle

3) What are used in place of tapped holes in the sheet metal to connect the sections?
The extender box sides, shown in green, are attached to the fixed blue box sides by means of the captured plate with tapped holes, shown in Figure 5 .

Figure 5: Captured plate with tapped holes

4) how does this addition affect your balance masses and counter masses (used to bring blade to correct position?) please provide detail of balance arm

The counter balance mass for the extended baffle box extends beyond the front of the arm cavity baffle and can be placed at different locations on the extended bars to balance the different extensions of the baffle box, as shown in Figure 1.

5) How does this affect your blade and flexure?

The suspension blade thickness will be increased appropriately to support the added weight of the arm cavity baffle with the extender box attached.

## Wide Angle Scattered Light from the Quad SUS Structure

The quad SUS lower structure, shown in Figure 2, has an octagonal structural ring that surrounds the HR surface of the TM and will catch some of the wide angle light scattered. A structural plate above the ring with an extended ledge will block the wide angle scattered light from hitting the upper stage of the SUS.

Figure 6: Lower Quad SUS Structure

# RESULTS OF WIDE ANGLE SCATTER NOISE CALCULATIONS

The wide angle scattering from the H1 ITMX HR surface was analyzed using ZEMAX ray tracing by applying a Lambertian scattering property to the TM HR surface and detecting the scattered light with detector surfaces. Almost all of the scattered rays from the TM are caught by the ACB structure, as shown in Figure 3, and Figure 4; some rays pass through the opening in the ACB into the adjacent spool piece, or into the adjacent manifold or BSC chamber.

Figure 7: ZEMAX Lambertian Scatter Ray Trace from H1 ITMX HR, ACB with Wide Angle Baffle Sides, top view

Figure 8: ZEMAX Lambertian Scatter Ray Trace from H1 ITMX HR, ACB with Wide Angle Baffle Sides, side view

## ZEMAX Lambertian Scattering Data

Table 1 summarizes the results of the ZEMAX measurements for the fractional wide angle scatter from the TM onto each adjacent scattering surface.

Table 1: Wide Angle Scattered Light from H1 ITMX,  ZEMAX Ray Trace

|  | Power | fractional scatter | x | y | z | length to TM | incident angle, deg | incident angle, rad |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ITM HR surface |  |  | 80 | -200 | 5001 |  |  |  |
| Rectangular Volume 1130: WIDE ANGLE BAF TOP LEDGE ITMX | 9.95E-01 | 8.31E-02 | -445 | -200 | 5315 | 612 | 59 | 1.03 |
| Rectangular Volume 1131: WIDE ANGLE BAF BOTTOM LEDGE ITMX | 2.64E+00 | 2.20E-01 | 300 | -200 | 5330 | 396 | 34 | 0.59 |
| Rectangular Volume 1132: WIDE ANGLE BAF PLATE 1 ITMX | 2.34E-02 | 1.96E-03 | 80 | 1046 | 5308 | 1283 | 76 | 1.33 |
| Rectangular Volume 1133: WIDE ANGLE BAF PLATE 2 ITMX | 1.37E-02 | 1.15E-03 | 80 | 718 | 5618 | 1106 | 56 | 0.98 |
| Rectangular Volume 1134: WIDE ANGLE BAF PLATE 3 ITMX | 4.49E-02 | 3.75E-03 | 80 | -1019 | 5327 | 881 | 68 | 1.19 |
| Rectangular Volume 1135: WIDE ANGLE BAF PLATE 4 ITMX | 9.77E-03 | 8.15E-04 | 80 | -662 | 5665 | 809 | 35 | 0.61 |
| Poly Object 1179: ACB_WIDE-ANGLE-BAFFLE-SIDE.POB | 8.21E-01 | 6.86E-02 | 80 | 400 | 5342 | 690 | 60 | 1.05 |
| Poly Object 1180: ACB_WIDE-ANGLE-BAFFLE-SIDE.POB | 2.87E+00 | 2.39E-01 | 80 | -400 | 5342 | 395 | 30 | 0.53 |
| Rectangle 1181: Side Extender | 8.50E-02 | 7.09E-03 | 80 | -428 | 5075 | 240 | 72 | 1.26 |
| Rectangle 1181: Plate above TM | 8.01E-02 | 6.69E-03 | -159 | -200 | 5042 | 242 | 80 | 1.40 |
| Rectangular Pipe 1182: SUS ring around TM | 1.18E-01 | 9.87E-03 | 254 | -200 | 5012 | 174 | 86 | 1.51 |
| Rectangle 1141: total ACB baffle | 3.22E+00 | 2.69E-01 | -70 | 0 | 5760 | 799 | 18 | 0.32 |
| Rectangle 1143: UPPER BSC | 4.00E-02 | 3.34E-03 | -2000 | -200 | 5364 | 2111 | 80 | 1.40 |
| Rectangle 1144: FLOOR BSC | 8.50E-02 | 7.09E-03 | 896 | -200 | 5120 | 825 | 82 | 1.43 |
| Detector Surf 1146: ADAPTER FLANGE OPENING (leaves BSC) | 5.39E-01 | 4.50E-02 |  |  |  |  |  |  |
| Detector Surf 1147: 3A-15 VIEWPORT FLANGE DETECTOR (enters manifold) | 4.00E-01 | 3.34E-02 | 80 | -620 | 10000 | 5017 | 5 | 0.08 |
| total scatter | 1.20E+01 | 1.00E+00 |  |  |  |  |  |  |
| derived spoolpiece scatter | 1.39E-01 | 1.16E-02 | 80 | -600 | 7300 | 2334 | 10 | 0.17 |

## Scattered Power Calculations

### Scattered Power Parameters

The arm cavity power was taken from T070303 with an arm cavity gain of 13000 referenced to the PSL input power of 125 W.

The following parameters were used for the scattered light noise calculations:

BRDF of chamber wall, sr^-1

BRDF of oxidized un-polished steel, sr^-1

laser wavelength, m

wave number, m^-1

hemispherical scattering loss fraction

TM wide angle (ref: T070089)

arm power, W

input laser power, W

arm cavity length, m

The values for the incident angle, fractional scattered power from TM, and the distance from the scattering surface to the TM surface are taken from Table 1.

### Incident Power

ACB BACK, W

WIDE ANGLE BAF TOP LEDGE, W

WIDE ANGLE BAF BOTTOM LEDGE, W

WIDE ANGLE BAF PLATE 1 ITMX, W

WIDE ANGLE BAF PLATE 2 ITMX, W

WIDE ANGLE BAF PLATE 3 ITMX, W

WIDE ANGLE BAF PLATE 4 ITMX, W

WIDE ANGLE BAF SIDE right, W

WIDE ANGLE BAF SIDE left, W

WIDE ANGLE BAF SIDE Extender, W

PEEK AND CABLING, W

### Power Scattered into IFO Mode

ACB BACK, W

WIDE ANGLE BAF TOP LEDGE, W

WIDE ANGLE BAF BOTTOM LEDGE, W

WIDE ANGLE BAF SIDE right, W

WIDE ANGLE BAF SIDE left, W

WIDE ANGLE BAF SIDE Extender, W

TOTAL ACB BACK AND BOX, W

PEEK AND CABLING, W

PLATE above TM, W

SUS ring around TM, W

TOTAL SUS, W

UPPER BSC, W

FLOOR BSC, W

SPOOL, W

MANIFOLD, W

Table 2 summarizes the calculated incident powers and the power scattered into the IFO mode.

Table 2:  Wide Angle Scattered Power

| SURFACE | INCIDENT POWER, W | POWER SCATTERED INTO IFO, W | MOTION SPECTRUM |
| --- | --- | --- | --- |
| WIDE ANGLE BAF TOP LEDGE ITMX | 0.674 | 2.19E-19 | x-ACB |
| WIDE ANGLE BAF BOTTOM LEDGE ITMX | 1.787 | 1.99E-18 | x-ACB |
| WIDE ANGLE BAF PLATE 1 ITMX | 0.016 | 5.56E-22 | x-MANIFOLD |
| WIDE ANGLE BAF PLATE 2 ITMX | 0.00975 | 9.41E-22 | x-MANIFOLD |
| WIDE ANGLE BAF PLATE 3 ITMX | 0.031 | 2.99E-21 | x-MANIFOLD |
| WIDE ANGLE BAF PLATE 4 ITMX | 0.0065 | 1.81E-21 | x-MANIFOLD |
| ACB_WIDE-ANGLE-BAFFLE-SIDE RIGHT | 0.561 | 1.23E-19 | x-ACB |
| ACB_WIDE-ANGLE-BAFFLE-SIDE LEFT | 1.95 | 2.27E-18 | x-ACB |
| Side Extender | 0.057 | 6.53E-20 | x-ACB |
| PEAK AND CABLING | 0.044 | 4.66-20 | x-ACB |
| total ACB Back | 2.186 | 7.01E-19 | x-ACB |
| Plate above TM | 0.057 | 1.21E-19 | x-ISI |
| SUS ring around TM | 0.081 | 1.43E-19 | x-ISI |
| UPPER BSC | 0.024 | 7.47E-22 | x-MANIFOLD |
| FLOOR BSC | 0.057 | 1.09E-20 | x-MANIFOLD |
| SPOOL BSC | 0.366 | 4.91E-20 | x-MANIFOLD |
| MANIFOLD | 0.268 | 7.71E-21 | x-MANIFOLD |

### Scattered Light Displacement Noise

#### SUS Parts

SUS

The SUS structure has the same seismic motion spectrum as the ISI table, which is assumed to be the same as the BSC ISI requirement. The calculated wide angle scattering displacement noise from the SUS octagonal ring and ledge are shown in  Figure 9.

Figure 9: SUS Wide Angle Scatter

#### Arm Cavity Back and Wide Angle Box

Wide Angle Baffle top

Wide Angle Baffle bottom

Wide Angle Baffle Sides

ACB BACK

ACB BACK & BOX

PD AND CABLING, W

The wide angle scattering displacement noise from the various ACB surfaces are shown in Figure 10.

Figure 10: Wide Angle Displacement Noise from Various Parts of the ACB

#### BSC Chamber Walls

Upper BSC

Floor BSC

#### Spoolpiece and Manifold

Spoolpiece

Manifold

The scattered light displacement noise from all of the components that re-scatter the wide angle scattered light from the TM are shown in Figure 11. Notice that the scattering from the BSC walls is acceptable, without the need for the Wide Angle Flat Baffles.

Figure 11: Total Components of TM Wide Angle Scattering

# INTERFACES

The Arm Cavity Baffle with the wide angle box structure can be hinged out of the way for ease of access to the HR surface of the TM. In addition, the top, bottom, and sides of the Arm Cavity Baffle wide angle box are removable to allow further access to the TM in the quad SUS for maintenance.

# Text boxes

Octagonal Structural Ring

Structural Plate