## SLC 6in no wedge summary
SLC 6" Dia, No-Wedge, Optical Viewport
o-ring |  |  | compression |  |  | gland fill |  |  | clearance (in) |  |  | air leak rate (torr-liter/sec) |  | Gland
designation | ID (in) | cross-section dia. (in) | nom | min | max | nom | min | max | nom | min | max | min | max | ID (in) | OD (in) | mechanical clear aperture (in) | comment
Ideal |  |  | 50% |  |  | 75% | [10 empty] | for high vacuum
Criteria |  |  |  | ≥ 20% |  |  | ≥ 60% | ≤ 90% |  | ≥ .005 | ≤ ~.031 |  |  |  | ≤ 5.895 | ≥ 5.200
#2-252 | 5.234 | 0.139 | 29.5% | 20.14% | 38.85% | 73.74% | 61.41% | 89.98% | 0.018 | 0.005 | 0.031 | 1.4089e-06 | 2.5342e-06 | 5.234 |  | 5.134 | mech clear aperture < optical clear aperture
#2-253 | 5.359 | [10 empty] | 1.4416e-06 | 2.5932e-06 | 5.3381 | 5.8041 | 5.2382 | The only solution which meets all criteria
#2-254 | 5.484 | [10 empty] | 1.4713e-06 | 2.9906e-06 | 5.484 | 5.904 | 5.384 | gland OD too close to optic edge
#2-355 | 5.225 | 0.21 | 32.86% | 26.67% | 39.05% | 77.98% | 68.04% | 90.04% |  |  |  | 1.4332e-06 | 2.1771e-06 | 5.225 | 5.855 | 5.125 | mech clear aperture < optical clear aperture
#2-356 | 5.35 | [10 empty] | 1.4662e-06 | 2.2272e-06 | 5.35 | 5.98 | 5.25 | gland OD exceeds optic face OD
#2-357 | 5.475 | [10 empty] | 1.4991e-06 | 2.2772e-06 | 5.475 | 6.105 | 5.375 | gland OD exceeds optic face OD
#2-431 | 5.225 | 0.275 | 38.18% | 33.45% | 42.91% | 79.41% | 70.58% | 89.77% |  |  |  | 1.2926e-06 | 1.6913e-06 | 5.225 |  | 5.125 | mech clear aperture < optical clear aperture
#2-432 | 5.35 | [10 empty] | 1.322e-06 | 1.7297e-06 | 5.35 | 6.23 | 5.25
#2-433 | 5.475 | [10 empty] | 1.3514e-06 | 1.7682e-06 | 5.475 | 6.355 | 5.375 | gland OD exceeds optic face OD
1) All critical dimensions toleranced at ± .001 in
2) Trading mechanical aperture for leak rate: An air leak rate of 2.9e-6 vs 1.7 t-l/s results in an increase of 1 nTorr in the vertex/diagonal sections or 2 nTorr in the end stations.
Compare to a pressure of 20 nTorr in these volumes (10 nTorr in the LLO end stations). This is a weak function
3) Note that the < 5 arc minute wedge requirement (E1100267-v1) corresponds to < .0087 in over the 6 in dia., which is within the thickness tolerance; These are not additive tolerances.
0.02 | 0.77 | 0.080639 | 0.0020948
| 0.75 | 0.078544
-0.02 | 0.73 | 0.07645 | -0.0020947

## SLC 6in no wedge #2-253 WHAT IF
Viewport Calculations | [14 empty] | .139 Cross Section
[16 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
value | ± tol | units | value | units | description | notes | [8 empty] | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
[5 empty] | OPTIC | [9 empty] | 0% | 0 | 0 | 0 | 0
6 | 0.01 | in |  |  | optic diameter | glass spec E1100267-v1 | [8 empty] | 5% | 0.95 | 6 | 2.5 | 10
0.03 | 0.01 | in |  |  | bevel width | glass spec E1100267-v1 (bevel dimension not specific!) | [8 empty] | 10% | 2 | 15 | 4.5 | 20
5.94 | 0.02 | in |  |  | optic face diameter | [9 empty] | 20% | 4.5 | 30 | 9 | 45
5.2 |  | in |  |  | clear aperture diameter | glass spec E1100267-v1 | [8 empty] | 30% | 10 | 70 | 20 | 90
0.75 | 0.001 | in |  |  | t, optic thickness | WHAT IF OPTIC THICKNESS WAS TIGHTLY CONTROLLED | [8 empty] | 40% | 25 | 116 | 40 | 190
fused silica |  | -- |  |  | optic material | [9 empty] | 50% | 39.0375 |  |  | 287.2115
10616762.4067 |  | psi | 73200000000 | Pa | E, optic elastic modulus | P. Yoder, Opto-mechanical Systems Design, Table 3.5
0.167 |  | -- |  |  | nu, optic Poison's ratio | P. Yoder, Opto-mechanical Systems Design, Table 3.5
7251.8869 |  | psi | 50000000 | Pa | optic ultimate strength | P. Yoder, Opto-mechanical Systems Design, Table 3.5
WINDOW CLAMP
7.97 | 0.01 |  |  |  | Dco, window clamp outer diameter
5.2 | 0.01 |  |  |  | Dci, window clamp inner diameter | standard dimension?
6.05 | 0.01 |  |  |  | Dcr, window clamp recess diameter
0.625 | 0.01 |  |  |  | tc, window clamp thickness | standard dimension?
0.5 | 0.001 |  |  |  | bc, window clamp recess depth
FLANGE
9.97 | 0.01 |  |  |  | Dfo, flange outer diameter | standard dimension?
5.259 | 0.01 |  |  |  | Dfi, flange inner diameter | check that this is larger than clear aperture
6.05 | 0.01 |  |  |  | Dfr, flange recess diameter
0.97 | 0.01 |  |  |  | tf, flange thickness | standard dimension?
0.384 | 0.001 |  |  |  | bf, flange recess depth
0.05 |  |  |  |  | tw, gland min wall thickness
VACUUM SEAL O-RING
#2-253 |  | -- |  |  | o-ring designation
5.359 | 0.035 | in |  |  | D, o-ring I.D.
0.139 | 0.004 | in |  |  | o-ring cross-section diameter
5.498 |  | in |  |  | o-ring center diameter
75 | 5 | -- |  |  | o-ring shore A hardness
50% |  | -- |  |  | % compression, ideal | 30% min, 50% ideal for vacuum
1.2612e-06 |  | torr-liter/sec |  |  | leak rate, max
1.104e-06 |  | torr-liter/sec |  |  | leak rate, min
163.1245 | 124.087 | lb/in |  |  | o-ring compression force per length | [9 empty] | .210 Cross Section
2817.5645 | 2143.2902 | lb |  |  | o-ring total compression force | [10 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
674.2743 |  | lb |  |  | min o-ring force | [9 empty] | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
4960.8547 |  | lb |  |  | max o-ring force | [9 empty] | 0% | 0 |  |  | 0
0.031 |  | in |  |  | g_allowable, allowable clearance gap | Parker Handbook, Fig 3-2, 70 shore A, < 100 psi | 55.49% | higher than recommended | [6 empty] | 5% | 3 |  |  | 15
[5 empty] | GLAND | [9 empty] | 10% | 7 |  |  | 28
5.359 | 0.001 |  |  |  | gland I.D. | Parker Handbook, design chart 4-3, vacuum | [8 empty] | 20% | 16 |  |  | 55
0.255 | 0.001 |  |  |  | gland width (should be .158 to .164) | Parker Handbook, design chart 4-3, vacuum | [8 empty] | 30% | 35 |  |  | 100
5.869 |  |  |  |  | gland OD (reference dimension) | [9 empty] | 40% | 68 |  |  | 250
0.065 | 0.001 |  |  |  | gland depth | Parker Handbook, design chart 4-3, vacuum | [8 empty] | 50% | 103.6425 |  |  | 375.911
0.01785 |  | in^2 |  |  | gland volume, min
0.01989 |  | in^2 |  |  | gland volume, max
0.01887 |  | in^2 |  |  | gland volume, nominal
75% |  | -- |  |  | gland fill, ideal | 75% ideal, 60% to 90% allowed |  |  | 0.47382
71.97% |  |  |  |  | gland fill, min |  | 0.052357 | 0.026179
89.98% |  |  |  |  | gland fill, max |  |  |  | 0.158
80.42% |  |  |  |  | gland fill, nominal |  |  |  | 0.052667
[5 empty] | CLAMPING RING/WASHER |  |  |  | 0.051
PEEK |  |  |  |  | material
5.99 |  |  |  |  | Do, ring O.D.
5.43 |  | in |  |  | Di, ring I.D.
0.56 |  | in |  |  | ring width
5.71 |  | in |  |  | ring center diameter | stair-step the ring cross-section to reduce clamping moment? |  |  |  |  | squeeze | He Leak Rate (cc/sec/in) | He Leak Rate (torr-liter/sec/in) | Air Leak Rate (torr-liter/sec/in)
0.125 | 0.001 | in |  |  | ring thickness | [5 empty] | 15% | 9.6e-07 | 7.296e-07 | 2.6628e-07
500000 |  | psi |  |  | elastic modulus | [5 empty] | 30% | 4.8e-07 | 3.648e-07 | 1.3314e-07
0.00011769 |  | in |  |  | max compression | [5 empty] | 50% | 2.6e-07 | 1.976e-07 | 7.2117e-08
[5 empty] | CLAMPING TOLERANCES |  |  |  |  | min | 43.88% |  |  | 7.3017e-08
| 0.001 | in |  |  | window clamp, recess for clamp ring, depth |  |  |  |  | max | 49.64% |  |  | 6.3918e-08
| 0.001 | in |  |  | flange , recess for optic, depth
| 0.001 | in |  |  | o-ring, gland, depth
| 0.001 | in |  |  | clamp ring, thickness
| 0.00011769 | in |  |  | clamp ring, max. compression
| 0.004 | in |  |  | o-ring, cross-section diameter
| 0.001 | in |  |  | optic, thickness
| 0.0091177 | in |  |  | clamp compression tolerance
0.065 |  | 46.76% |  |  | nominal squeeze (compression)
0.061 |  | 43.88% |  |  | min squeeze (compression)
0.069 |  | 49.64% |  |  | max squeeze (compression)
0.009 |  | in |  |  | nominal gap
0.013 |  | in |  |  | maximum gap
0.005 |  | in |  |  | minimum gap
WINDOW CLAMP BOLT LOADS
6 |  |  |  |  | number of bolts
5/16 - 24 UNRF, silver plated |  |  |  |  | bolt type
0.3125 |  | in |  |  | bolt major diameter
0.2614 |  | in |  |  | bolt minor diameter
0.053666 |  | in^2 |  |  | tensile stress area
826.8091 |  | lb |  |  | max load per bolt to compress o-ring
15406.5079 |  | psi |  |  | max stress per bolt to compress o-ring
5510 |  | lb |  |  | tensile strength | Unbrako, SHCS, Stainless steel
30000 |  | psi |  |  | min tensile yield stress | Unbrako, SHCS, Stainless steel
95000 |  | psi |  |  | min tensile ultimate stress | Unbrako, SHCS, Stainless steel
190 |  | in-lb |  |  | seating torque | Unbrako, SHCS, Stainless steel
415.6327 |  | lb |  |  | pressure force on optic
5.264e-09 |  | std. cc cm/cm^2 sec bar |  |  | F, permeability rate of air through elastomer (table 3-24)
14.7 |  | psi |  |  | P, q, pressure
1 |  | -- |  |  | Q, factor depending on % squeeze (dry) (Fig 3-11)
0.005 |  | -- |  |  | S, squeeze fraction (squeeze percent/100)
2.8738e-07 |  | std. cc/sec |  |  | L=.7*F*D*P*Q*(1-S)^2, permeability rate through o-ring
|  | torr-liter/sec |  |  | L, permeability rate through o-ring
|  | OPTIC STRESS |  |  | Young, Roark's Formulas for Stress & Strain, 6th ed., Table 24, case 10a
simply supported, flat circular plate of constant thickness
with uniformly distributed pressure (r0 = 0)
2.749 |  | in |  |  | a, simply supported radius
383953.6363 |  | in-lb |  |  | D=E*t^3/[12*(1-nu^2), plate constant
0.015625 |  | -- |  |  | L11 (for r0 = 0), loading constant
0.19794 |  | -- |  |  | L17 (for r0 = 0), loading constant
-0.00015126 |  | in |  |  | yc = (-q a^4) (5+nu)/(64D(1+nu)), Center deflection
21.9885 |  | in-lb/in |  |  | Mc = q a^2 (3+nu)/16, unit bending moment at center
234.5436 |  | psi | 1617121.3148 | Pa | sc = 6 Mc/t^2, bending stress at plate center (outer fiber)
3 |  | -- |  |  | FS, Factor of Safety
9.3064 |  | -- |  |  | Safety Margin
See Parker catalog 5705A, Vacuum Sealing
Fluorocarbon (FKM)
Perfluoroelastomer (FFKM)
compare Parker o-ring leak rate with PSI measurements
Parker recommends Gask-O-Seals -- what is this?
Surface finish of 16 rms common, but 32 have been successful
Use leak rate approx. method in section on gasses
Design chart 4-2. If difficult to assemble, use design chart 4-1 and design table 4-1
Vacuum seal glands, design chart 4-7
For viewport seals, dimension I.D. and width (not O.D.)
[7 empty] | 0.63855 | 0.10483 | 0.16183

## SLC - 6in no wedge #2-253
Viewport Calculations | [16 empty] | .139 Cross Section
[18 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
value | ± tol | units | value | units | description | notes | [10 empty] | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
[5 empty] | OPTIC | [11 empty] | 0% | 0 | 0 | 0 | 0
6 | 0.01 | in |  |  | optic diameter | glass spec E1100267-v1 | [10 empty] | 5% | 0.95 | 6 | 2.5 | 10
0.03 | 0.01 | in |  |  | bevel width | glass spec E1100267-v1 (bevel dimension not specific! Should prefer smaller ID glands) | [10 empty] | 10% | 2 | 15 | 4.5 | 20
5.94 | 0.02 | in |  |  | optic face diameter | [11 empty] | 20% | 4.5 | 30 | 9 | 45
5.2 |  | in |  |  | clear aperture diameter | glass spec E1100267-v1 | [10 empty] | 30% | 10 | 70 | 20 | 90
0.75 | 0.01 | in |  |  | t, optic thickness | [11 empty] | 40% | 25 | 116 | 40 | 190
fused silica |  | -- |  |  | optic material | [11 empty] | 50% | 39.0375 |  |  | 287.2115
10674777.5018 |  | psi | 73600000000 | Pa | E, optic elastic modulus | P. Yoder, Opto-mechanical Systems Design, Table 3.5
0.167 |  | -- |  |  | nu, optic Poison's ratio | P. Yoder, Opto-mechanical Systems Design, Table 3.5
7251.8869 |  | psi | 50000000 | Pa | optic ultimate strength | P. Yoder, Opto-mechanical Systems Design, Table 3.5
WINDOW CLAMP
7.97 | 0.01 |  |  |  | Dco, window clamp outer diameter
5.2382 | 0.01 |  |  |  | Dci, window clamp inner diameter | standard dimension?
6.05 | 0.01 |  |  |  | Dcr, window clamp recess diameter
0.625 | 0.01 |  |  |  | tc, window clamp thickness | standard dimension?
0.5 | 0.001 |  |  |  | bc, window clamp recess depth
FLANGE
9.97 | 0.01 |  |  |  | Dfo, flange outer diameter | standard dimension?
5.2382 | 0.01 |  |  |  | Dfi, flange inner diameter | check that this is larger than clear aperture
6.05 | 0.01 |  |  |  | Dfr, flange recess diameter
0.97 | 0.01 |  |  |  | tf, flange thickness | standard dimension?
0.393 | 0.001 |  |  |  | bf, flange recess depth
0.05 |  |  |  |  | tw, gland min wall thickness
VACUUM SEAL O-RING
#2-253 |  | -- |  |  | o-ring designation
5.359 | 0.035 | in | 136.1186 | mm | D, o-ring I.D.
0.139 | 0.004 | in | 3.5306 | mm | d, o-ring cross-section diameter |  | 3.6322
5.498 |  | in | 139.6492 | mm | o-ring center diameter |  | 3.429
75 | 5 | -- |  |  | o-ring shore A hardness |  | 1.7145
50% |  | -- |  |  | % compression, ideal | 30% min, 50% ideal for vacuum | 1.8161 | 0.0095493
2.5932e-06 |  | torr-liter/sec |  |  | leak rate, max |  | 0.63361 |  |  | 5.47
1.4416e-06 |  | torr-liter/sec |  |  | leak rate, min |  | 0.67055 |  |  | 5.95
163.1245 | 124.087 | lb/in |  |  | o-ring compression force per length |  | 287.2115 |  |  | -0.48 | [6 empty] | .210 Cross Section
2817.5645 | 2143.2902 | lb |  |  | o-ring total compression force |  | 0.12587 | [10 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
674.2743 |  | lb |  |  | min o-ring force |  | -0.13333 | [9 empty] | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
4960.8547 |  | lb |  |  | max o-ring force | [11 empty] | 0% | 0 |  |  | 0
1152.2833 |  | psi |  |  | max pressure on clamping ring | [11 empty] | 5% | 3 |  |  | 15
2863.1116 |  | psi |  |  | max pressure on o-ring | [11 empty] | 10% | 7 |  |  | 28
0.031 |  | in | 0.7874 | mm | g_allowable, allowable clearance gap | Parker Handbook, Fig 3-2, 70 shore A, < 100 psi | 42.07% | higher than recommended | [8 empty] | 20% | 16 |  |  | 55
[5 empty] | GLAND | [11 empty] | 30% | 35 |  |  | 100
5.3381 | 0.001 |  |  |  | gland I.D. | A. Roth, Vacuum Sealing Techniques, Table 3-30 | [10 empty] | 40% | 68 |  |  | 250
0.233 | 0.001 | in | 5.9182 | mm | gland width (should be .158 to .164) | Parker Handbook, design chart 4-3, vacuum |  |  |  | 251 | [6 empty] | 50% | 103.6425 |  |  | 375.911
5.8041 |  |  |  |  | gland OD (reference dimension) |  |  | 5.5711 |  | 4 | 30
0.072 | 0.001 | in | 1.8288 | mm | gland depth | Set to 1/2 the maximum o-ring cross-section diameter |  | 2.7856 |  | 55.7778
0.017632 |  | in^2 |  |  | gland area, min
0.024336 |  | in^2 |  |  | gland area, max |  |  |  |  | 4.1833 | 10.8
0.02097 |  | in^2 |  |  | gland area, nominal |  |  |  | A | 3.8615 | 51.6
75% |  | -- |  |  | gland fill, ideal | 75% ideal, 60% to 90% allowed |  | 0.015175 | h
58.82% |  |  |  |  | gland fill, min |  |  | 0.09 | L
91.09% |  |  |  |  | gland fill, max |  |  | 0.097922
72.36% |  |  |  |  | gland fill, nominal
[5 empty] | CLAMPING RING/WASHER |  |  | 0.067
PEEK |  |  |  |  | material |  |  | 0.42675
5.99 |  |  |  |  | Do, ring O.D.
5.43 |  | in |  |  | Di, ring I.D. | [7 empty] | squeeze | He Leak Rate (cc/sec/in) | He Leak Rate (torr-liter/sec/in) | Air Leak Rate (torr-liter/sec/in)
0.56 |  | in |  |  | ring width | [7 empty] | 15% | 9.6e-07 | 7.296e-07 | 2.6628e-07
5.71 |  | in |  |  | ring center diameter | stair-step the ring cross-section to reduce clamping moment? | [6 empty] | 30% | 4.8e-07 | 3.648e-07 | 1.3314e-07
0.125 | 0.001 | in |  |  | ring thickness | [7 empty] | 50% | 2.6e-07 | 1.976e-07 | 7.2117e-08
500000 |  | psi |  |  | elastic modulus | [6 empty] | min | 25.9% |  |  | 1.5013e-07
0.00011769 |  | in |  |  | max compression | [6 empty] | max | 44.6% |  |  | 8.3464e-08
CLAMPING TOLERANCES
| 0.001 | in |  |  | window clamp, recess for clamp ring, depth
| 0.001 | in |  |  | flange , recess for optic, depth
| 0.001 | in |  |  | o-ring, gland, depth
| 0.001 | in |  |  | clamp ring, thickness
| 0.00011769 | in |  |  | clamp ring, max. compression
| 0.004 | in |  |  | o-ring, cross-section diameter
| 0.01 | in |  |  | optic, thickness
| 0.018118 | in |  |  | clamp compression tolerance
0.049 |  | 35.25% |  |  | nominal squeeze (compression) |  | 0.1295
0.036 |  | 25.9% |  |  | min squeeze (compression)
0.062 |  | 44.6% |  |  | max squeeze (compression)
0.018 |  | in |  |  | nominal gap
0.031 |  | in |  |  | maximum gap
0.005 |  | in |  |  | minimum gap
WINDOW CLAMP BOLT LOADS
6 |  |  |  |  | number of bolts
5/16 - 24 UNRF, silver plated |  |  |  |  | bolt type
0.3125 |  | in |  |  | bolt major diameter
0.2614 |  | in |  |  | bolt minor diameter
0.053666 |  | in^2 |  |  | tensile stress area
826.8091 |  | lb |  |  | max load per bolt to compress o-ring
15406.5079 |  | psi |  |  | max stress per bolt to compress o-ring
5510 |  | lb |  |  | tensile strength | Unbrako, SHCS, Stainless steel
30000 |  | psi |  |  | min tensile yield stress | Unbrako, SHCS, Stainless steel
95000 |  | psi |  |  | min tensile ultimate stress | Unbrako, SHCS, Stainless steel
190 |  | in-lb |  |  | seating torque | Unbrako, SHCS, Stainless steel
415.6327 |  | lb |  |  | pressure force on optic
5.264e-09 |  | std. cc cm/cm^2 sec bar |  |  | F, permeability rate of air through elastomer (table 3-24)
14.7 |  | psi |  |  | P, q, pressure
1.35 |  | -- |  |  | Q, factor depending on % squeeze (dry) (Fig 3-11)
0.2 |  | -- |  |  | S, squeeze fraction (squeeze percent/100)
2.508e-07 |  | std. cc/sec |  |  | L=.7*F*D*P*Q*(1-S)^2, permeability rate through o-ring
1.9061e-07 |  | torr-liter/sec |  |  | L, permeability rate through o-ring
|  | OPTIC STRESS |  |  | Young, Roark's Formulas for Stress & Strain, 6th ed., Table 24, case 10a
simply supported, flat circular plate of constant thickness
with uniformly distributed pressure (r0 = 0)
2.736 |  | in |  |  | a, simply supported radius | 2.7855
386051.7436 |  | in-lb |  |  | D=E*t^3/[12*(1-nu^2), plate constant | 2.749
0.015625 |  | -- |  |  | L11 (for r0 = 0), loading constant
0.19794 |  | -- |  |  | L17 (for r0 = 0), loading constant
-0.00014761 |  | in |  |  | yc = (-q a^4) (5+nu)/(64D(1+nu)), Center deflection
21.781 |  | in-lb/in |  |  | Mc = q a^2 (3+nu)/16, unit bending moment at center
232.3306 |  | psi | 1601862.7704 | Pa | sc = 6 Mc/t^2, bending stress at plate center (outer fiber)
3 |  | -- |  |  | FS, Factor of Safety
9.4046 |  | -- |  |  | Safety Margin
See Parker catalog 5705A, Vacuum Sealing
Fluorocarbon (FKM)
Perfluoroelastomer (FFKM)
compare Parker o-ring leak rate with PSI measurements
Parker recommends Gask-O-Seals -- what is this?
Surface finish of 16 rms common, but 32 have been successful
Use leak rate approx. method in section on gasses
Design chart 4-2. If difficult to assemble, use design chart 4-1 and design table 4-1
Vacuum seal glands, design chart 4-7
For viewport seals, dimension I.D. and width (not O.D.)
0.16183
[7 empty] | 0.63855 | 0.10483
0.0087267
0.3125
7.9375
45
0.14173
[7 empty] | 3.6 | depth | 0.23622
[7 empty] | 6 | width | 0.55118
14

## SLC - 6in no wedge #2-254
Viewport Calculations | [14 empty] | .139 Cross Section
[16 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
value | ± tol | units | value | units | description | notes | [8 empty] | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
[5 empty] | OPTIC | [9 empty] | 0% | 0 | 0 | 0 | 0
6 | 0.01 | in |  |  | optic diameter | glass spec E1100267-v1 | [8 empty] | 5% | 0.95 | 6 | 2.5 | 10
0.03 | 0.01 | in |  |  | bevel width | glass spec E1100267-v1 (bevel dimension not specific!) | [8 empty] | 10% | 2 | 15 | 4.5 | 20
5.94 | 0.02 | in |  |  | optic face diameter | [9 empty] | 20% | 4.5 | 30 | 9 | 45
5.2 |  | in |  |  | clear aperture diameter | glass spec E1100267-v1 | [8 empty] | 30% | 10 | 70 | 20 | 90
0.75 | 0.01 | in |  |  | t, optic thickness | [9 empty] | 40% | 25 | 116 | 40 | 190
fused silica |  | -- |  |  | optic material | [9 empty] | 50% | 39.0375 |  |  | 287.2115
10616762.4067 |  | psi | 73200000000 | Pa | E, optic elastic modulus | P. Yoder, Opto-mechanical Systems Design, Table 3.5
0.167 |  | -- |  |  | nu, optic Poison's ratio | P. Yoder, Opto-mechanical Systems Design, Table 3.5
7251.8869 |  | psi | 50000000 | Pa | optic ultimate strength | P. Yoder, Opto-mechanical Systems Design, Table 3.5
WINDOW CLAMP
7.97 | 0.01 |  |  |  | Dco, window clamp outer diameter
5.2 | 0.01 |  |  |  | Dci, window clamp inner diameter | standard dimension?
6.05 | 0.01 |  |  |  | Dcr, window clamp recess diameter
0.625 | 0.01 |  |  |  | tc, window clamp thickness | standard dimension?
0.5 | 0.001 |  |  |  | bc, window clamp recess depth
FLANGE
9.97 | 0.01 |  |  |  | Dfo, flange outer diameter | standard dimension?
5.384 | 0.01 |  |  |  | Dfi, flange inner diameter | check that this is larger than clear aperture
6.05 | 0.01 |  |  |  | Dfr, flange recess diameter
0.97 | 0.01 |  |  |  | tf, flange thickness | standard dimension?
0.393 | 0.001 |  |  |  | bf, flange recess depth
0.05 |  |  |  |  | tw, gland min wall thickness
VACUUM SEAL O-RING
#2-254 |  | -- |  |  | o-ring designation
5.484 | 0.035 | in |  |  | D, o-ring I.D.
0.139 | 0.004 | in |  |  | o-ring cross-section diameter
5.623 |  | in |  |  | o-ring center diameter
75 | 5 | -- |  |  | o-ring shore A hardness
50% |  | -- |  |  | % compression, ideal | 30% min, 50% ideal for vacuum
2.9906e-06 |  | torr-liter/sec |  |  | leak rate, max
1.4713e-06 |  | torr-liter/sec |  |  | leak rate, min
163.1245 | 124.087 | lb/in |  |  | o-ring compression force per length | [9 empty] | .210 Cross Section
2881.6234 | 2192.0191 | lb |  |  | o-ring total compression force | [10 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
689.6043 |  | lb |  |  | min o-ring force | [9 empty] | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
5073.6424 |  | lb |  |  | max o-ring force | [9 empty] | 0% | 0 |  |  | 0
0.031 |  | in |  |  | g_allowable, allowable clearance gap | Parker Handbook, Fig 3-2, 70 shore A, < 100 psi | 28.05% | higher than recommended | [6 empty] | 5% | 3 |  |  | 15
[5 empty] | GLAND | [9 empty] | 10% | 7 |  |  | 28
5.484 | 0.001 |  |  |  | gland I.D. | Parker Handbook, design chart 4-3, vacuum | [8 empty] | 20% | 16 |  |  | 55
0.21 | 0.001 |  |  |  | gland width (should be .158 to .164) | Parker Handbook, design chart 4-3, vacuum | [8 empty] | 30% | 35 |  |  | 100
5.904 |  |  |  |  | gland OD (reference dimension) | Too Close to Optic Flat Diameter! | [8 empty] | 40% | 68 |  |  | 250
0.08 | 0.001 |  |  |  | gland depth | Parker Handbook, design chart 4-3, vacuum | [8 empty] | 50% | 103.6425 |  |  | 375.911
0.01785 |  | in^2 |  |  | gland volume, min
0.02331 |  | in^2 |  |  | gland volume, max
0.02058 |  | in^2 |  |  | gland volume, nominal
75% |  | -- |  |  | gland fill, ideal | 75% ideal, 60% to 90% allowed |  |  | 0.47382
61.41% |  |  |  |  | gland fill, min |  | 0.052357 | 0.026179
89.98% |  |  |  |  | gland fill, max
73.74% |  |  |  |  | gland fill, nominal
CLAMPING RING/WASHER
PEEK |  |  |  |  | material
5.99 |  |  |  |  | Do, ring O.D.
5.43 |  | in |  |  | Di, ring I.D.
0.56 |  | in |  |  | ring width
5.71 |  | in |  |  | ring center diameter | stair-step the ring cross-section to reduce clamping moment? |  |  |  |  | squeeze | He Leak Rate (cc/sec/in) | He Leak Rate (torr-liter/sec/in) | Air Leak Rate (torr-liter/sec/in)
0.125 | 0.001 | in |  |  | ring thickness | [5 empty] | 15% | 9.6e-07 | 7.296e-07 | 2.6628e-07
500000 |  | psi |  |  | elastic modulus | [5 empty] | 30% | 4.8e-07 | 3.648e-07 | 1.3314e-07
0.00012036 |  | in |  |  | max compression | [5 empty] | 50% | 2.6e-07 | 1.976e-07 | 7.2117e-08
[5 empty] | CLAMPING TOLERANCES |  |  |  |  | min | 20.14% |  |  | 1.693e-07
| 0.001 | in |  |  | window clamp, recess for clamp ring, depth |  |  |  |  | max | 38.85% |  |  | 8.329e-08
| 0.001 | in |  |  | flange , recess for optic, depth
| 0.001 | in |  |  | o-ring, gland, depth
| 0.001 | in |  |  | clamp ring, thickness
| 0.00012036 | in |  |  | clamp ring, max. compression
| 0.004 | in |  |  | o-ring, cross-section diameter
| 0.01 | in |  |  | optic, thickness
| 0.01812 | in |  |  | clamp compression tolerance
0.041 |  | 29.5% |  |  | nominal squeeze (compression)
0.028 |  | 20.14% |  |  | min squeeze (compression)
0.054 |  | 38.85% |  |  | max squeeze (compression)
0.018 |  | in |  |  | nominal gap
0.031 |  | in |  |  | maximum gap
0.005 |  | in |  |  | minimum gap
WINDOW CLAMP BOLT LOADS
6 |  |  |  |  | number of bolts
5/16 - 24 UNRF, silver plated |  |  |  |  | bolt type
0.3125 |  | in |  |  | bolt major diameter
0.2614 |  | in |  |  | bolt minor diameter
0.053666 |  | in^2 |  |  | tensile stress area
845.6071 |  | lb |  |  | max load per bolt to compress o-ring
15756.7832 |  | psi |  |  | max stress per bolt to compress o-ring
5510 |  | lb |  |  | tensile strength | Unbrako, SHCS, Stainless steel
30000 |  | psi |  |  | min tensile yield stress | Unbrako, SHCS, Stainless steel
95000 |  | psi |  |  | min tensile ultimate stress | Unbrako, SHCS, Stainless steel
190 |  | in-lb |  |  | seating torque | Unbrako, SHCS, Stainless steel
415.6327 |  | lb |  |  | pressure force on optic
5.264e-09 |  | std. cc cm/cm^2 sec bar |  |  | F, permeability rate of air through elastomer (table 3-24)
14.7 |  | psi |  |  | P, q, pressure
1 |  | -- |  |  | Q, factor depending on % squeeze (dry) (Fig 3-11)
0.005 |  | -- |  |  | S, squeeze fraction (squeeze percent/100)
2.9409e-07 |  | std. cc/sec |  |  | L=.7*F*D*P*Q*(1-S)^2, permeability rate through o-ring
|  | torr-liter/sec |  |  | L, permeability rate through o-ring
|  | OPTIC STRESS |  |  | Young, Roark's Formulas for Stress & Strain, 6th ed., Table 24, case 10a
simply supported, flat circular plate of constant thickness
with uniformly distributed pressure (r0 = 0)
2.8115 |  | in |  |  | a, simply supported radius
383953.6363 |  | in-lb |  |  | D=E*t^3/[12*(1-nu^2), plate constant
0.015625 |  | -- |  |  | L11 (for r0 = 0), loading constant
0.19794 |  | -- |  |  | L17 (for r0 = 0), loading constant
-0.00016549 |  | in |  |  | yc = (-q a^4) (5+nu)/(64D(1+nu)), Center deflection
22.9997 |  | in-lb/in |  |  | Mc = q a^2 (3+nu)/16, unit bending moment at center
245.3298 |  | psi | 1691489.4658 | Pa | sc = 6 Mc/t^2, bending stress at plate center (outer fiber)
3 |  | -- |  |  | FS, Factor of Safety
8.8532 |  | -- |  |  | Safety Margin
See Parker catalog 5705A, Vacuum Sealing
Fluorocarbon (FKM)
Perfluoroelastomer (FFKM)
compare Parker o-ring leak rate with PSI measurements
Parker recommends Gask-O-Seals -- what is this?
Surface finish of 16 rms common, but 32 have been successful
Use leak rate approx. method in section on gasses
Design chart 4-2. If difficult to assemble, use design chart 4-1 and design table 4-1
Vacuum seal glands, design chart 4-7
For viewport seals, dimension I.D. and width (not O.D.)
[7 empty] | 0.63855 | 0.10483 | 0.16183

## SLC - 6in no wedge #2-355
Viewport Calculations | [14 empty] | .210 Cross Section
[16 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
value | ± tol | units | value | units | description | notes | [8 empty] | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
[5 empty] | OPTIC | [9 empty] | 0% | 0 |  |  | 0
6 | 0.01 | in |  |  | optic diameter | [9 empty] | 5% | 3 |  |  | 15
0.03 | 0.01 | in |  |  | bevel width | glass spec E1100267-v1 (bevel dimension not specific!) | [8 empty] | 10% | 7 |  |  | 28
5.94 | 0.02 | in |  |  | optic face diameter | [9 empty] | 20% | 16 |  |  | 55
5.38 |  | in |  |  | clear aperture diameter | [9 empty] | 30% | 35 |  |  | 100
0.75 | 0.01 | in |  |  | t, optic thickness | [9 empty] | 40% | 68 |  |  | 250
fused silica |  | -- |  |  | optic material | [9 empty] | 50% | 103.6425 |  |  | 375.911
10616762.4067 |  | psi | 73200000000 | Pa | E, optic elastic modulus | P. Yoder, Opto-mechanical Systems Design, Table 3.5
0.167 |  | -- |  |  | nu, optic Poison's ratio | P. Yoder, Opto-mechanical Systems Design, Table 3.5
7251.8869 |  | psi | 50000000 | Pa | optic ultimate strength | P. Yoder, Opto-mechanical Systems Design, Table 3.5
WINDOW CLAMP
7.97 | 0.01 |  |  |  | Dco, window clamp outer diameter
5.38 | 0.01 |  |  |  | Dci, window clamp inner diameter | standard dimension?
6.05 | 0.01 |  |  |  | Dcr, window clamp recess diameter
0.625 | 0.01 |  |  |  | tc, window clamp thickness | standard dimension?
0.5 | 0.001 |  |  |  | bc, window clamp recess depth
FLANGE
9.97 | 0.01 |  |  |  | Dfo, flange outer diameter | standard dimension?
5.125 | 0.01 |  |  |  | Dfi, flange inner diameter | check that this is larger than clear aperture
6.05 | 0.01 |  |  |  | Dfr, flange recess diameter
0.97 | 0.01 |  |  |  | tf, flange thickness | standard dimension?
0.397 | 0.001 |  |  |  | bf, flange recess depth
0.05 |  |  |  |  | tw, gland min wall thickness
VACUUM SEAL O-RING
#2-355 |  | -- |  |  | o-ring designation
5.225 | 0.037 | in |  |  | D, o-ring I.D.
0.21 | 0.005 | in |  |  | o-ring cross-section diameter
5.435 |  | in |  |  | o-ring center diameter
75 | 5 | -- |  |  | o-ring shore A hardness
50% |  | -- |  |  | % compression, ideal | 30% min, 50% ideal for vacuum
2.1771e-06 |  | torr-liter/sec |  |  | leak rate, max
1.4332e-06 |  | torr-liter/sec |  |  | leak rate, min
239.7767 | 136.1343 | lb/in |  |  | o-ring compression force per length
4094.0816 | 2324.4319 | lb |  |  | o-ring total compression force
1769.6497 |  | lb |  |  | min o-ring force
6418.5134 |  | lb |  |  | max o-ring force
0.035 |  | in |  |  | g_allowable, allowable clearance gap | Parker Handbook, Fig 3-2, 70 shore A, < 100 psi
GLAND
5.225 | 0.001 |  |  |  | gland I.D. | Parker Handbook, design chart 4-3, vacuum
0.315 | 0.001 |  |  |  | gland width | Parker Handbook, design chart 4-3, vacuum
5.855 |  |  |  |  | gland OD (reference dimension) | exceeds OD of optic face
0.12 | 0.001 |  |  |  | gland depth | Parker Handbook, design chart 4-3, vacuum
0.040635 |  | in^2 |  |  | gland volume, min
0.048825 |  | in^2 |  |  | gland volume, max
0.04473 |  | in^2 |  |  | gland volume, nominal
75% |  | -- |  |  | gland fill, ideal | 75% ideal, 60% to 90% allowed
67.6% |  |  |  |  | gland fill, min
89.34% |  |  |  |  | gland fill, max
77.43% |  |  |  |  | gland fill, nominal
CLAMPING RING/WASHER
PEEK |  |  |  |  | material
5.99 |  |  |  |  | Do, ring O.D.
5.43 |  | in |  |  | Di, ring I.D.
0.56 |  | in |  |  | ring width
5.71 |  | in |  |  | ring center diameter | stair-step the ring cross-section to reduce clamping moment? |  |  |  |  | squeeze | He Leak Rate (cc/sec/in) | He Leak Rate (torr-liter/sec/in) | Air Leak Rate (torr-liter/sec/in)
0.125 | 0.001 | in |  |  | ring thickness | [5 empty] | 15% | 9.6e-07 | 7.296e-07 | 2.6628e-07
500000 |  | psi |  |  | elastic modulus | [5 empty] | 30% | 4.8e-07 | 3.648e-07 | 1.3314e-07
0.00015227 |  | in |  |  | max compression | [5 empty] | 50% | 2.6e-07 | 1.976e-07 | 7.2117e-08
[5 empty] | CLAMPING TOLERANCES |  |  |  |  | min | 26.19% |  |  | 1.275e-07
| 0.001 | in |  |  | window clamp, recess for clamp ring, depth |  |  |  |  | max | 38.57% |  |  | 8.3937e-08
| 0.001 | in |  |  | flange , recess for optic, depth
| 0.001 | in |  |  | o-ring, gland, depth
| 0.001 | in |  |  | clamp ring, thickness
| 0.00015227 | in |  |  | clamp ring, max. compression
| 0.005 | in |  |  | o-ring, cross-section diameter
| 0.01 | in |  |  | optic, thickness
| 0.019152 | in |  |  | clamp compression tolerance
0.068 |  | 32.38% |  |  | nominal squeeze (compression)
0.055 |  | 26.19% |  |  | min squeeze (compression)
0.081 |  | 38.57% |  |  | max squeeze (compression)
0.022 |  | in |  |  | nominal gap
0.035 |  | in |  |  | maximum gap
0.009 |  | in |  |  | minimum gap
WINDOW CLAMP BOLT LOADS
6 |  |  |  |  | number of bolts
5/16 - 24 UNRF, silver plated |  |  |  |  | bolt type
0.3125 |  | in |  |  | bolt major diameter
0.2614 |  | in |  |  | bolt minor diameter
0.053666 |  | in^2 |  |  | tensile stress area
1069.7522 |  | lb |  |  | max load per bolt to compress o-ring
19933.4356 |  | psi |  |  | max stress per bolt to compress o-ring
5510 |  | lb |  |  | tensile strength | Unbrako, SHCS, Stainless steel
30000 |  | psi |  |  | min tensile yield stress | Unbrako, SHCS, Stainless steel
95000 |  | psi |  |  | min tensile ultimate stress | Unbrako, SHCS, Stainless steel
190 |  | in-lb |  |  | seating torque | Unbrako, SHCS, Stainless steel
415.6327 |  | lb |  |  | pressure force on optic
5.264e-09 |  | std. cc cm/cm^2 sec bar |  |  | F, permeability rate of air through elastomer (table 3-24)
14.7 |  | psi |  |  | P, q, pressure
1 |  | -- |  |  | Q, factor depending on % squeeze (dry) (Fig 3-11)
0.005 |  | -- |  |  | S, squeeze fraction (squeeze percent/100)
2.802e-07 |  | std. cc/sec |  |  | L=.7*F*D*P*Q*(1-S)^2, permeability rate through o-ring
|  | torr-liter/sec |  |  | L, permeability rate through o-ring
OPTIC STRESS & DEFORMATION
PLATE BENDING DUE TO 1 ATMOSPHERE
Young, Roark's Formulas for Stress & Strain, 6th ed., Table 24, case 10a
simply supported, flat circular plate of constant thickness
with uniformly distributed pressure (r0 = 0)
2.7175 |  | in |  |  | a, simply supported radius
383953.6363 |  | in-lb |  |  | D=E*t^3/[12*(1-nu^2), plate constant
0.015625 |  | -- |  |  | L11 (for r0 = 0), loading constant
0.19794 |  | -- |  |  | L17 (for r0 = 0), loading constant
-0.00014445 |  | in |  |  | yc = (-q a^4) (5+nu)/(64D(1+nu)), Center deflection
21.4874 |  | in-lb/in |  |  | Mc = q a^2 (3+nu)/16, unit bending moment at center
229.1993 |  | psi | 1580273.3905 | Pa | sc = 6 Mc/t^2, bending stress at plate center (outer fiber)
4 |  | -- |  |  | FS, Factor of Safety
6.91 |  | -- |  |  | Safety Margin
O-RING COMPRESSION
6418.5134 |  | lb |  |  | maximum total force from o-ring compression
1170.7503 |  | psi |  |  | max pressure from o-ring
5351892.5946 |  | psi | 36900000000 | Pa | bulk modulus
0.00016407 |  | in |  |  | compressive displacement
1277.8811
See Parker catalog 5705A, Vacuum Sealing
Fluorocarbon (FKM)
Perfluoroelastomer (FFKM)
compare Parker o-ring leak rate with PSI measurements
Parker recommends Gask-O-Seals -- what is this?
Surface finish of 16 rms common, but 32 have been successful
|  | Use leak rate approx. method in section on gasses |  |  |  |  | 0.63855 | 0.10483 | 0.16183
Design chart 4-2. If difficult to assemble, use design chart 4-1 and design table 4-1
Vacuum seal glands, design chart 4-7
For viewport seals, dimension I.D. and width (not O.D.)

## SLC - 6in no wedge #2-356
Viewport Calculations | [9 empty] | .210 Cross Section
[11 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
value | ± tol | units | value | units | description | notes |  |  |  | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
[5 empty] | OPTIC |  |  |  |  | 0% | 0 |  |  | 0
6 | 0.01 | in |  |  | optic diameter |  |  |  |  | 5% | 3 |  |  | 15
0.03 | 0.01 | in |  |  | bevel width | glass spec E1100267-v1 (bevel dimension not specific!) |  |  |  | 10% | 7 |  |  | 28
5.94 | 0.02 | in |  |  | optic face diameter |  |  |  |  | 20% | 16 |  |  | 55
5.38 |  | in |  |  | clear aperture diameter |  |  |  |  | 30% | 35 |  |  | 100
0.75 | 0.01 | in |  |  | t, optic thickness |  |  |  |  | 40% | 68 |  |  | 250
fused silica |  | -- |  |  | optic material |  |  |  |  | 50% | 103.6425 |  |  | 375.911
10616762.4067 |  | psi | 73200000000 | Pa | E, optic elastic modulus | P. Yoder, Opto-mechanical Systems Design, Table 3.5
0.167 |  | -- |  |  | nu, optic Poison's ratio | P. Yoder, Opto-mechanical Systems Design, Table 3.5
7251.8869 |  | psi | 50000000 | Pa | optic ultimate strength | P. Yoder, Opto-mechanical Systems Design, Table 3.5
WINDOW CLAMP
7.97 | 0.01 |  |  |  | Dco, window clamp outer diameter
5.38 | 0.01 |  |  |  | Dci, window clamp inner diameter | standard dimension?
6.05 | 0.01 |  |  |  | Dcr, window clamp recess diameter
0.625 | 0.01 |  |  |  | tc, window clamp thickness | standard dimension?
0.5 | 0.001 |  |  |  | bc, window clamp recess depth
FLANGE
9.97 | 0.01 |  |  |  | Dfo, flange outer diameter | standard dimension?
5.25 | 0.01 |  |  |  | Dfi, flange inner diameter | check that this is larger than clear aperture
6.05 | 0.01 |  |  |  | Dfr, flange recess diameter
0.97 | 0.01 |  |  |  | tf, flange thickness | standard dimension?
0.397 | 0.001 |  |  |  | bf, flange recess depth
0.05 |  |  |  |  | tw, gland min wall thickness
VACUUM SEAL O-RING
#2-356 |  | -- |  |  | o-ring designation
5.35 | 0.037 | in |  |  | D, o-ring I.D.
0.21 | 0.005 | in |  |  | o-ring cross-section diameter
5.56 |  | in |  |  | o-ring center diameter
75 | 5 | -- |  |  | o-ring shore A hardness
50% |  | -- |  |  | % compression, ideal | 30% min, 50% ideal for vacuum
leak rate (cc/sec/in)
leak rate (torr-Liter/s)
239.7767 | 136.1343 | lb/in |  |  | o-ring compression force per length
4188.2417 | 2377.8917 | lb |  |  | o-ring total compression force
1810.35 |  | lb |  |  | min o-ring force
6566.1334 |  | lb |  |  | max o-ring force
0.035 |  | in |  |  | g_allowable, allowable clearance gap | Parker Handbook, Fig 3-2, 70 shore A, < 100 psi
GLAND
5.35 | 0.001 |  |  |  | gland I.D. | Parker Handbook, design chart 4-3, vacuum
0.315 | 0.001 |  |  |  | gland width | Parker Handbook, design chart 4-3, vacuum
5.98 |  |  |  |  | gland OD (reference dimension) | exceeds OD of optic face
0.12 | 0.001 |  |  |  | gland depth | Parker Handbook, design chart 4-3, vacuum
0.040635 |  | in^2 |  |  | gland volume, min
0.048825 |  | in^2 |  |  | gland volume, max
0.04473 |  | in^2 |  |  | gland volume, nominal
75% |  | -- |  |  | gland fill, ideal | 75% ideal, 60% to 90% allowed
67.6% |  |  |  |  | gland fill, min
89.34% |  |  |  |  | gland fill, max
77.43% |  |  |  |  | gland fill, nominal
CLAMPING RING/WASHER
PEEK |  |  |  |  | material
5.99 |  |  |  |  | Do, ring O.D.
5.43 |  | in |  |  | Di, ring I.D.
0.56 |  | in |  |  | ring width
5.71 |  | in |  |  | ring center diameter | stair-step the ring cross-section to reduce clamping moment?
0.125 | 0.001 | in |  |  | ring thickness
500000 |  | psi |  |  | elastic modulus
0.00015577 |  | in |  |  | max compression
CLAMPING TOLERANCES
| 0.001 | in |  |  | window clamp, recess for clamp ring, depth
| 0.001 | in |  |  | flange , recess for optic, depth
| 0.001 | in |  |  | o-ring, gland, depth
| 0.001 | in |  |  | clamp ring, thickness
| 0.00015577 | in |  |  | clamp ring, max. compression
| 0.005 | in |  |  | o-ring, cross-section diameter
| 0.01 | in |  |  | optic, thickness
| 0.019156 | in |  |  | clamp compression tolerance
0.068 |  | 32.38% |  |  | nominal squeeze (compression)
0.055 |  | 26.19% |  |  | min squeeze (compression)
0.081 |  | 38.57% |  |  | max squeeze (compression)
0.022 |  | in |  |  | nominal gap
0.035 |  | in |  |  | maximum gap
0.009 |  | in |  |  | minimum gap
WINDOW CLAMP BOLT LOADS
6 |  |  |  |  | number of bolts
5/16 - 24 UNRF, silver plated |  |  |  |  | bolt type
0.3125 |  | in |  |  | bolt major diameter
0.2614 |  | in |  |  | bolt minor diameter
0.053666 |  | in^2 |  |  | tensile stress area
1094.3556 |  | lb |  |  | max load per bolt to compress o-ring
20391.8863 |  | psi |  |  | max stress per bolt to compress o-ring
5510 |  | lb |  |  | tensile strength | Unbrako, SHCS, Stainless steel
30000 |  | psi |  |  | min tensile yield stress | Unbrako, SHCS, Stainless steel
95000 |  | psi |  |  | min tensile ultimate stress | Unbrako, SHCS, Stainless steel
190 |  | in-lb |  |  | seating torque | Unbrako, SHCS, Stainless steel
415.6327 |  | lb |  |  | pressure force on optic
5.264e-09 |  | std. cc cm/cm^2 sec bar |  |  | F, permeability rate of air through elastomer (table 3-24)
14.7 |  | psi |  |  | P, q, pressure
1 |  | -- |  |  | Q, factor depending on % squeeze (dry) (Fig 3-11)
0.005 |  | -- |  |  | S, squeeze fraction (squeeze percent/100)
2.869e-07 |  | std. cc/sec |  |  | L=.7*F*D*P*Q*(1-S)^2, permeability rate through o-ring
|  | torr-liter/sec |  |  | L, permeability rate through o-ring
OPTIC STRESS & DEFORMATION
PLATE BENDING DUE TO 1 ATMOSPHERE
Young, Roark's Formulas for Stress & Strain, 6th ed., Table 24, case 10a
simply supported, flat circular plate of constant thickness
with uniformly distributed pressure (r0 = 0)
2.78 |  | in |  |  | a, simply supported radius
383953.6363 |  | in-lb |  |  | D=E*t^3/[12*(1-nu^2), plate constant
0.015625 |  | -- |  |  | L11 (for r0 = 0), loading constant
0.19794 |  | -- |  |  | L17 (for r0 = 0), loading constant
-0.0001582 |  | in |  |  | yc = (-q a^4) (5+nu)/(64D(1+nu)), Center deflection
22.4872 |  | in-lb/in |  |  | Mc = q a^2 (3+nu)/16, unit bending moment at center
239.8633 |  | psi | 1653798.9566 | Pa | sc = 6 Mc/t^2, bending stress at plate center (outer fiber)
4 |  | -- |  |  | FS, Factor of Safety
6.5584 |  | -- |  |  | Safety Margin
O-RING COMPRESSION
6566.1334 |  | lb |  |  | maximum total force from o-ring compression
1171.2493 |  | psi |  |  | max pressure from o-ring
5351892.5946 |  | psi | 36900000000 | Pa | bulk modulus
0.00016414 |  | in |  |  | compressive displacement
1307.2712
See Parker catalog 5705A, Vacuum Sealing
Fluorocarbon (FKM)
Perfluoroelastomer (FFKM)
compare Parker o-ring leak rate with PSI measurements
Parker recommends Gask-O-Seals -- what is this?
Surface finish of 16 rms common, but 32 have been successful
|  | Use leak rate approx. method in section on gasses |  |  |  |  | 0.63855 | 0.10483 | 0.16183
Design chart 4-2. If difficult to assemble, use design chart 4-1 and design table 4-1
Vacuum seal glands, design chart 4-7
For viewport seals, dimension I.D. and width (not O.D.)

## SLC - 6in no wedge #2-357
Viewport Calculations | [9 empty] | .210 Cross Section
[11 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
value | ± tol | units | value | units | description | notes |  |  |  | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
[5 empty] | OPTIC |  |  |  |  | 0% | 0 |  |  | 0
6 | 0.01 | in |  |  | optic diameter |  |  |  |  | 5% | 3 |  |  | 15
0.03 | 0.01 | in |  |  | bevel width | glass spec E1100267-v1 (bevel dimension not specific!) |  |  |  | 10% | 7 |  |  | 28
5.94 | 0.02 | in |  |  | optic face diameter |  |  |  |  | 20% | 16 |  |  | 55
5.38 |  | in |  |  | clear aperture diameter |  |  |  |  | 30% | 35 |  |  | 100
0.75 | 0.01 | in |  |  | t, optic thickness |  |  |  |  | 40% | 68 |  |  | 250
fused silica |  | -- |  |  | optic material |  |  |  |  | 50% | 103.6425 |  |  | 375.911
10616762.4067 |  | psi | 73200000000 | Pa | E, optic elastic modulus | P. Yoder, Opto-mechanical Systems Design, Table 3.5
0.167 |  | -- |  |  | nu, optic Poison's ratio | P. Yoder, Opto-mechanical Systems Design, Table 3.5
7251.8869 |  | psi | 50000000 | Pa | optic ultimate strength | P. Yoder, Opto-mechanical Systems Design, Table 3.5
WINDOW CLAMP
7.97 | 0.01 |  |  |  | Dco, window clamp outer diameter
5.38 | 0.01 |  |  |  | Dci, window clamp inner diameter | standard dimension?
6.05 | 0.01 |  |  |  | Dcr, window clamp recess diameter
0.625 | 0.01 |  |  |  | tc, window clamp thickness | standard dimension?
0.5 | 0.001 |  |  |  | bc, window clamp recess depth
FLANGE
9.97 | 0.01 |  |  |  | Dfo, flange outer diameter | standard dimension?
5.375 | 0.01 |  |  |  | Dfi, flange inner diameter | check that this is larger than clear aperture
6.05 | 0.01 |  |  |  | Dfr, flange recess diameter
0.97 | 0.01 |  |  |  | tf, flange thickness | standard dimension?
0.393 | 0.001 |  |  |  | bf, flange recess depth
0.05 |  |  |  |  | tw, gland min wall thickness
VACUUM SEAL O-RING
#2-357 |  | -- |  |  | o-ring designation
5.475 | 0.037 | in |  |  | D, o-ring I.D.
0.21 | 0.005 | in |  |  | o-ring cross-section diameter
5.685 |  | in |  |  | o-ring center diameter
75 | 5 | -- |  |  | o-ring shore A hardness
50% |  | -- |  |  | % compression, ideal | 30% min, 50% ideal for vacuum
leak rate (cc/sec/in)
leak rate (torr-Liter/s)
239.7767 | 136.1343 | lb/in |  |  | o-ring compression force per length
4282.4018 | 2431.3515 | lb |  |  | o-ring total compression force
1851.0503 |  | lb |  |  | min o-ring force
6713.7533 |  | lb |  |  | max o-ring force
0.031 |  | in |  |  | g_allowable, allowable clearance gap | Parker Handbook, Fig 3-2, 70 shore A, < 100 psi
GLAND
5.475 | 0.001 |  |  |  | gland I.D. | Parker Handbook, design chart 4-3, vacuum
0.315 | 0.001 |  |  |  | gland width | Parker Handbook, design chart 4-3, vacuum
6.105 |  |  |  |  | gland OD (reference dimension) | exceeds OD of optic
0.123 | 0.001 |  |  |  | gland depth | Parker Handbook, design chart 4-3, vacuum
0.04032 |  | in^2 |  |  | gland volume, min
0.04851 |  | in^2 |  |  | gland volume, max
0.044415 |  | in^2 |  |  | gland volume, nominal
75% |  | -- |  |  | gland fill, ideal | 75% ideal, 60% to 90% allowed
68.04% |  |  |  |  | gland fill, min
90.04% |  |  |  |  | gland fill, max
77.98% |  |  |  |  | gland fill, nominal
CLAMPING RING/WASHER
PEEK |  |  |  |  | material
5.99 |  |  |  |  | Do, ring O.D.
5.43 |  | in |  |  | Di, ring I.D.
0.56 |  | in |  |  | ring width
5.71 |  | in |  |  | ring center diameter | stair-step the ring cross-section to reduce clamping moment?
0.125 | 0.001 | in |  |  | ring thickness
500000 |  | psi |  |  | elastic modulus
0.00015927 |  | in |  |  | max compression
CLAMPING TOLERANCES
| 0.001 | in |  |  | window clamp, recess for clamp ring, depth
| 0.001 | in |  |  | flange , recess for optic, depth
| 0.001 | in |  |  | o-ring, gland, depth
| 0.001 | in |  |  | clamp ring, thickness
| 0.00015927 | in |  |  | clamp ring, max. compression
| 0.005 | in |  |  | o-ring, cross-section diameter
| 0.01 | in |  |  | optic, thickness
| 0.019159 | in |  |  | clamp compression tolerance
0.069 |  | 32.86% |  |  | nominal squeeze (compression)
0.056 |  | 26.67% |  |  | min squeeze (compression)
0.082 |  | 39.05% |  |  | max squeeze (compression)
0.018 |  | in |  |  | nominal gap
0.031 |  | in |  |  | maximum gap
0.005 |  | in |  |  | minimum gap
WINDOW CLAMP BOLT LOADS
6 |  |  |  |  | number of bolts
5/16 - 24 UNRF, silver plated |  |  |  |  | bolt type
0.3125 |  | in |  |  | bolt major diameter
0.2614 |  | in |  |  | bolt minor diameter
0.053666 |  | in^2 |  |  | tensile stress area
1118.9589 |  | lb |  |  | max load per bolt to compress o-ring
20850.337 |  | psi |  |  | max stress per bolt to compress o-ring
5510 |  | lb |  |  | tensile strength | Unbrako, SHCS, Stainless steel
30000 |  | psi |  |  | min tensile yield stress | Unbrako, SHCS, Stainless steel
95000 |  | psi |  |  | min tensile ultimate stress | Unbrako, SHCS, Stainless steel
190 |  | in-lb |  |  | seating torque | Unbrako, SHCS, Stainless steel
415.6327 |  | lb |  |  | pressure force on optic
5.264e-09 |  | std. cc cm/cm^2 sec bar |  |  | F, permeability rate of air through elastomer (table 3-24)
14.7 |  | psi |  |  | P, q, pressure
1 |  | -- |  |  | Q, factor depending on % squeeze (dry) (Fig 3-11)
0.005 |  | -- |  |  | S, squeeze fraction (squeeze percent/100)
2.936e-07 |  | std. cc/sec |  |  | L=.7*F*D*P*Q*(1-S)^2, permeability rate through o-ring
|  | torr-liter/sec |  |  | L, permeability rate through o-ring
OPTIC STRESS & DEFORMATION
PLATE BENDING DUE TO 1 ATMOSPHERE
Young, Roark's Formulas for Stress & Strain, 6th ed., Table 24, case 10a
simply supported, flat circular plate of constant thickness
with uniformly distributed pressure (r0 = 0)
2.8425 |  | in |  |  | a, simply supported radius
383953.6363 |  | in-lb |  |  | D=E*t^3/[12*(1-nu^2), plate constant
0.015625 |  | -- |  |  | L11 (for r0 = 0), loading constant
0.19794 |  | -- |  |  | L17 (for r0 = 0), loading constant
-0.00017291 |  | in |  |  | yc = (-q a^4) (5+nu)/(64D(1+nu)), Center deflection
23.5097 |  | in-lb/in |  |  | Mc = q a^2 (3+nu)/16, unit bending moment at center
250.7697 |  | psi | 1728996.3183 | Pa | sc = 6 Mc/t^2, bending stress at plate center (outer fiber)
4 |  | -- |  |  | FS, Factor of Safety
6.2296 |  | -- |  |  | Safety Margin
O-RING COMPRESSION
6713.7533 |  | lb |  |  | maximum total force from o-ring compression
1171.7269 |  | psi |  |  | max pressure from o-ring
5351892.5946 |  | psi | 36900000000 | Pa | bulk modulus
0.0001642 |  | in |  |  | compressive displacement
1336.6613
See Parker catalog 5705A, Vacuum Sealing
Fluorocarbon (FKM)
Perfluoroelastomer (FFKM)
compare Parker o-ring leak rate with PSI measurements
Parker recommends Gask-O-Seals -- what is this?
Surface finish of 16 rms common, but 32 have been successful
|  | Use leak rate approx. method in section on gasses |  |  |  |  | 0.63855 | 0.10483 | 0.16183
Design chart 4-2. If difficult to assemble, use design chart 4-1 and design table 4-1
Vacuum seal glands, design chart 4-7
For viewport seals, dimension I.D. and width (not O.D.)

## SLC - 6in no wedge #2-432
Viewport Calculations | [14 empty] | .210 Cross Section
[16 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
value | ± tol | units | value | units | description | notes | [8 empty] | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
[5 empty] | OPTIC | [9 empty] | 0% | 0 |  |  | 0
6 | 0.01 | in |  |  | optic diameter | [9 empty] | 5% | 2.5 |  |  | 15
0.03 | 0.01 | in |  |  | bevel width | glass spec E1100267-v1 (bevel dimension not specific!) | [8 empty] | 10% | 7 |  |  | 30
5.94 | 0.02 | in |  |  | optic face diameter | [9 empty] | 20% | 22 |  |  | 65
5.38 |  | in |  |  | clear aperture diameter | [9 empty] | 30% | 45 |  |  | 135
0.75 | 0.01 | in |  |  | t, optic thickness | [9 empty] | 40% | 110 |  |  | 300
fused silica |  | -- |  |  | optic material | [9 empty] | 50% | 172.484 |  |  | 457.8035
10616762.4067 |  | psi | 73200000000 | Pa | E, optic elastic modulus | P. Yoder, Opto-mechanical Systems Design, Table 3.5
0.167 |  | -- |  |  | nu, optic Poison's ratio | P. Yoder, Opto-mechanical Systems Design, Table 3.5
7251.8869 |  | psi | 50000000 | Pa | optic ultimate strength | P. Yoder, Opto-mechanical Systems Design, Table 3.5
WINDOW CLAMP
7.97 | 0.01 |  |  |  | Dco, window clamp outer diameter
5.38 | 0.01 |  |  |  | Dci, window clamp inner diameter | standard dimension?
6.05 | 0.01 |  |  |  | Dcr, window clamp recess diameter
0.625 | 0.01 |  |  |  | tc, window clamp thickness | standard dimension?
0.5 | 0.001 |  |  |  | bc, window clamp recess depth
FLANGE
9.97 | 0.01 |  |  |  | Dfo, flange outer diameter | standard dimension?
5.25 | 0.01 |  |  |  | Dfi, flange inner diameter | check that this is larger than clear aperture
6.05 | 0.01 |  |  |  | Dfr, flange recess diameter
0.97 | 0.01 |  |  |  | tf, flange thickness | standard dimension?
0.393 | 0.001 |  |  |  | bf, flange recess depth
0.05 |  |  |  |  | tw, gland min wall thickness
VACUUM SEAL O-RING
#2-432 |  | -- |  |  | o-ring designation
5.35 | 0.037 | in |  |  | D, o-ring I.D.
0.275 | 0.006 | in |  |  | o-ring cross-section diameter
5.625 |  | in |  |  | o-ring center diameter
75 | 5 | -- |  |  | o-ring shore A hardness
50% |  | -- |  |  | % compression, ideal | 30% min, 50% ideal for vacuum
1.7297e-06 |  | torr-liter/sec |  |  | leak rate, max
1.322e-06 |  | torr-liter/sec |  |  | leak rate, min
315.1438 | 142.6597 | lb/in |  |  | o-ring compression force per length
5569.0498 | 2521.0059 | lb |  |  | o-ring total compression force
3048.0439 |  | lb |  |  | min o-ring force
8090.0556 |  | lb |  |  | max o-ring force
0.031 |  | in |  |  | g_allowable, allowable clearance gap | Parker Handbook, Fig 3-2, 70 shore A, < 100 psi | 40.13% | higher than recommended
GLAND
5.35 | 0.001 |  |  |  | gland I.D. | Parker Handbook, design chart 4-3, vacuum
0.44 | 0.001 |  |  |  | gland width (should be .309 to .314) | Parker Handbook, design chart 4-3, vacuum
6.23 |  |  |  |  | gland OD (reference dimension) | exceeds OD of optic
0.152 | 0.001 |  |  |  | gland depth | Parker Handbook, design chart 4-3, vacuum
0.06908 |  | in^2 |  |  | gland volume, min
0.08052 |  | in^2 |  |  | gland volume, max
0.0748 |  | in^2 |  |  | gland volume, nominal
75% |  | -- |  |  | gland fill, ideal | 75% ideal, 60% to 90% allowed |  |  | 0.47382
70.58% |  |  |  |  | gland fill, min |  | 0.052357 | 0.026179
89.77% |  |  |  |  | gland fill, max
79.41% |  |  |  |  | gland fill, nominal |  |  | 0.1375
[5 empty] | CLAMPING RING/WASHER |  |  | .
PEEK |  |  |  |  | material
5.99 |  |  |  |  | Do, ring O.D.
5.43 |  | in |  |  | Di, ring I.D.
0.56 |  | in |  |  | ring width
5.71 |  | in |  |  | ring center diameter | stair-step the ring cross-section to reduce clamping moment? |  |  |  |  | squeeze | He Leak Rate (cc/sec/in) | He Leak Rate (torr-liter/sec/in) | Air Leak Rate (torr-liter/sec/in)
0.125 | 0.001 | in |  |  | ring thickness | [5 empty] | 15% | 9.6e-07 | 7.296e-07 | 2.6628e-07
500000 |  | psi |  |  | elastic modulus | [5 empty] | 30% | 4.8e-07 | 3.648e-07 | 1.3314e-07
0.00019192 |  | in |  |  | max compression | [5 empty] | 50% | 2.6e-07 | 1.976e-07 | 7.2117e-08
[5 empty] | CLAMPING TOLERANCES |  |  |  |  | min | 33.45% |  |  | 9.7884e-08
| 0.001 | in |  |  | window clamp, recess for clamp ring, depth |  |  |  |  | max | 42.91% |  |  | 7.4811e-08
| 0.001 | in |  |  | flange , recess for optic, depth
| 0.001 | in |  |  | o-ring, gland, depth
| 0.001 | in |  |  | clamp ring, thickness
| 0.00019192 | in |  |  | clamp ring, max. compression
| 0.006 | in |  |  | o-ring, cross-section diameter
| 0.01 | in |  |  | optic, thickness
| 0.020192 | in |  |  | clamp compression tolerance
0.105 |  | 38.18% |  |  | nominal squeeze (compression)
0.092 |  | 33.45% |  |  | min squeeze (compression)
0.118 |  | 42.91% |  |  | max squeeze (compression)
0.018 |  | in |  |  | nominal gap
0.031 |  | in |  |  | maximum gap
0.005 |  | in |  |  | minimum gap
WINDOW CLAMP BOLT LOADS
6 |  |  |  |  | number of bolts
5/16 - 24 UNRF, silver plated |  |  |  |  | bolt type
0.3125 |  | in |  |  | bolt major diameter
0.2614 |  | in |  |  | bolt minor diameter
0.053666 |  | in^2 |  |  | tensile stress area
1348.3426 |  | lb |  |  | max load per bolt to compress o-ring
25124.6031 |  | psi |  |  | max stress per bolt to compress o-ring
5510 |  | lb |  |  | tensile strength | Unbrako, SHCS, Stainless steel
30000 |  | psi |  |  | min tensile yield stress | Unbrako, SHCS, Stainless steel
95000 |  | psi |  |  | min tensile ultimate stress | Unbrako, SHCS, Stainless steel
190 |  | in-lb |  |  | seating torque | Unbrako, SHCS, Stainless steel
415.6327 |  | lb |  |  | pressure force on optic
5.264e-09 |  | std. cc cm/cm^2 sec bar |  |  | F, permeability rate of air through elastomer (table 3-24)
14.7 |  | psi |  |  | P, q, pressure
1 |  | -- |  |  | Q, factor depending on % squeeze (dry) (Fig 3-11)
0.005 |  | -- |  |  | S, squeeze fraction (squeeze percent/100)
2.869e-07 |  | std. cc/sec |  |  | L=.7*F*D*P*Q*(1-S)^2, permeability rate through o-ring
|  | torr-liter/sec |  |  | L, permeability rate through o-ring
|  | OPTIC STRESS |  |  | Young, Roark's Formulas for Stress & Strain, 6th ed., Table 24, case 10a
simply supported, flat circular plate of constant thickness
with uniformly distributed pressure (r0 = 0)
2.8125 |  | in |  |  | a, simply supported radius
383953.6363 |  | in-lb |  |  | D=E*t^3/[12*(1-nu^2), plate constant
0.015625 |  | -- |  |  | L11 (for r0 = 0), loading constant
0.19794 |  | -- |  |  | L17 (for r0 = 0), loading constant
-0.00016573 |  | in |  |  | yc = (-q a^4) (5+nu)/(64D(1+nu)), Center deflection
23.016 |  | in-lb/in |  |  | Mc = q a^2 (3+nu)/16, unit bending moment at center
245.5044 |  | psi | 1692692.9446 | Pa | sc = 6 Mc/t^2, bending stress at plate center (outer fiber)
3 |  | -- |  |  | FS, Factor of Safety
8.8462 |  | -- |  |  | Safety Margin
See Parker catalog 5705A, Vacuum Sealing
Fluorocarbon (FKM)
Perfluoroelastomer (FFKM)
compare Parker o-ring leak rate with PSI measurements
Parker recommends Gask-O-Seals -- what is this?
Surface finish of 16 rms common, but 32 have been successful
Use leak rate approx. method in section on gasses
Design chart 4-2. If difficult to assemble, use design chart 4-1 and design table 4-1
Vacuum seal glands, design chart 4-7
For viewport seals, dimension I.D. and width (not O.D.)
[7 empty] | 0.63855 | 0.10483 | 0.16183

## SLC - 6in no wedge #2-433
Viewport Calculations | [14 empty] | .210 Cross Section
[16 empty] | 70 shore A | 70 shore A | 80 shore A | 80 shore A
value | ± tol | units | value | units | description | notes | [8 empty] | compression | min F (lb/in) | max F (lb/in) | min F (lb/in) | max F (lb/in)
[5 empty] | OPTIC | [9 empty] | 0% | 0 |  |  | 0
6 | 0.01 | in |  |  | optic diameter | [9 empty] | 5% | 2.5 |  |  | 15
0.03 | 0.01 | in |  |  | bevel width | glass spec E1100267-v1 (bevel dimension not specific!) | [8 empty] | 10% | 7 |  |  | 30
5.94 | 0.02 | in |  |  | optic face diameter | [9 empty] | 20% | 22 |  |  | 65
5.38 |  | in |  |  | clear aperture diameter | [9 empty] | 30% | 45 |  |  | 135
0.75 | 0.01 | in |  |  | t, optic thickness | [9 empty] | 40% | 110 |  |  | 300
fused silica |  | -- |  |  | optic material | [9 empty] | 50% | 172.484 |  |  | 457.8035
10616762.4067 |  | psi | 73200000000 | Pa | E, optic elastic modulus | P. Yoder, Opto-mechanical Systems Design, Table 3.5
0.167 |  | -- |  |  | nu, optic Poison's ratio | P. Yoder, Opto-mechanical Systems Design, Table 3.5
7251.8869 |  | psi | 50000000 | Pa | optic ultimate strength | P. Yoder, Opto-mechanical Systems Design, Table 3.5
WINDOW CLAMP
7.97 | 0.01 |  |  |  | Dco, window clamp outer diameter
5.38 | 0.01 |  |  |  | Dci, window clamp inner diameter | standard dimension?
6.05 | 0.01 |  |  |  | Dcr, window clamp recess diameter
0.625 | 0.01 |  |  |  | tc, window clamp thickness | standard dimension?
0.5 | 0.001 |  |  |  | bc, window clamp recess depth
FLANGE
9.97 | 0.01 |  |  |  | Dfo, flange outer diameter | standard dimension?
5.375 | 0.01 |  |  |  | Dfi, flange inner diameter | check that this is larger than clear aperture
6.05 | 0.01 |  |  |  | Dfr, flange recess diameter
0.97 | 0.01 |  |  |  | tf, flange thickness | standard dimension?
0.393 | 0.001 |  |  |  | bf, flange recess depth
0.05 |  |  |  |  | tw, gland min wall thickness
VACUUM SEAL O-RING
#2-433 |  | -- |  |  | o-ring designation
5.475 | 0.037 | in |  |  | D, o-ring I.D.
0.275 | 0.006 | in |  |  | o-ring cross-section diameter
5.75 |  | in |  |  | o-ring center diameter
75 | 5 | -- |  |  | o-ring shore A hardness
50% |  | -- |  |  | % compression, ideal | 30% min, 50% ideal for vacuum
1.7682e-06 |  | torr-liter/sec |  |  | leak rate, max
1.3514e-06 |  | torr-liter/sec |  |  | leak rate, min
315.1438 | 142.6597 | lb/in |  |  | o-ring compression force per length
5692.8064 | 2577.0282 | lb |  |  | o-ring total compression force
3115.7782 |  | lb |  |  | min o-ring force
8269.8346 |  | lb |  |  | max o-ring force
0.031 |  | in |  |  | g_allowable, allowable clearance gap | Parker Handbook, Fig 3-2, 70 shore A, < 100 psi | 40.13% | higher than recommended
GLAND
5.475 | 0.001 |  |  |  | gland I.D. | Parker Handbook, design chart 4-3, vacuum
0.44 | 0.001 |  |  |  | gland width (should be .309 to .314) | Parker Handbook, design chart 4-3, vacuum
6.355 |  |  |  |  | gland OD (reference dimension) | exceeds OD of optic
0.152 | 0.001 |  |  |  | gland depth | Parker Handbook, design chart 4-3, vacuum
0.06908 |  | in^2 |  |  | gland volume, min
0.08052 |  | in^2 |  |  | gland volume, max
0.0748 |  | in^2 |  |  | gland volume, nominal
75% |  | -- |  |  | gland fill, ideal | 75% ideal, 60% to 90% allowed |  |  | 0.47382
70.58% |  |  |  |  | gland fill, min |  | 0.052357 | 0.026179
89.77% |  |  |  |  | gland fill, max
79.41% |  |  |  |  | gland fill, nominal |  |  | 0.1375
[5 empty] | CLAMPING RING/WASHER |  |  | .
PEEK |  |  |  |  | material
5.99 |  |  |  |  | Do, ring O.D.
5.43 |  | in |  |  | Di, ring I.D.
0.56 |  | in |  |  | ring width
5.71 |  | in |  |  | ring center diameter | stair-step the ring cross-section to reduce clamping moment? |  |  |  |  | squeeze | He Leak Rate (cc/sec/in) | He Leak Rate (torr-liter/sec/in) | Air Leak Rate (torr-liter/sec/in)
0.125 | 0.001 | in |  |  | ring thickness | [5 empty] | 15% | 9.6e-07 | 7.296e-07 | 2.6628e-07
500000 |  | psi |  |  | elastic modulus | [5 empty] | 30% | 4.8e-07 | 3.648e-07 | 1.3314e-07
0.00019619 |  | in |  |  | max compression | [5 empty] | 50% | 2.6e-07 | 1.976e-07 | 7.2117e-08
[5 empty] | CLAMPING TOLERANCES |  |  |  |  | min | 33.45% |  |  | 9.7884e-08
| 0.001 | in |  |  | window clamp, recess for clamp ring, depth |  |  |  |  | max | 42.91% |  |  | 7.4811e-08
| 0.001 | in |  |  | flange , recess for optic, depth
| 0.001 | in |  |  | o-ring, gland, depth
| 0.001 | in |  |  | clamp ring, thickness
| 0.00019619 | in |  |  | clamp ring, max. compression
| 0.006 | in |  |  | o-ring, cross-section diameter
| 0.01 | in |  |  | optic, thickness
| 0.020196 | in |  |  | clamp compression tolerance
0.105 |  | 38.18% |  |  | nominal squeeze (compression)
0.092 |  | 33.45% |  |  | min squeeze (compression)
0.118 |  | 42.91% |  |  | max squeeze (compression)
0.018 |  | in |  |  | nominal gap
0.031 |  | in |  |  | maximum gap
0.005 |  | in |  |  | minimum gap
WINDOW CLAMP BOLT LOADS
6 |  |  |  |  | number of bolts
5/16 - 24 UNRF, silver plated |  |  |  |  | bolt type
0.3125 |  | in |  |  | bolt major diameter
0.2614 |  | in |  |  | bolt minor diameter
0.053666 |  | in^2 |  |  | tensile stress area
1378.3058 |  | lb |  |  | max load per bolt to compress o-ring
25682.9276 |  | psi |  |  | max stress per bolt to compress o-ring
5510 |  | lb |  |  | tensile strength | Unbrako, SHCS, Stainless steel
30000 |  | psi |  |  | min tensile yield stress | Unbrako, SHCS, Stainless steel
95000 |  | psi |  |  | min tensile ultimate stress | Unbrako, SHCS, Stainless steel
190 |  | in-lb |  |  | seating torque | Unbrako, SHCS, Stainless steel
415.6327 |  | lb |  |  | pressure force on optic
5.264e-09 |  | std. cc cm/cm^2 sec bar |  |  | F, permeability rate of air through elastomer (table 3-24)
14.7 |  | psi |  |  | P, q, pressure
1 |  | -- |  |  | Q, factor depending on % squeeze (dry) (Fig 3-11)
0.005 |  | -- |  |  | S, squeeze fraction (squeeze percent/100)
2.936e-07 |  | std. cc/sec |  |  | L=.7*F*D*P*Q*(1-S)^2, permeability rate through o-ring
|  | torr-liter/sec |  |  | L, permeability rate through o-ring
|  | OPTIC STRESS |  |  | Young, Roark's Formulas for Stress & Strain, 6th ed., Table 24, case 10a
simply supported, flat circular plate of constant thickness
with uniformly distributed pressure (r0 = 0)
2.875 |  | in |  |  | a, simply supported radius
383953.6363 |  | in-lb |  |  | D=E*t^3/[12*(1-nu^2), plate constant
0.015625 |  | -- |  |  | L11 (for r0 = 0), loading constant
0.19794 |  | -- |  |  | L17 (for r0 = 0), loading constant
-0.00018096 |  | in |  |  | yc = (-q a^4) (5+nu)/(64D(1+nu)), Center deflection
24.0503 |  | in-lb/in |  |  | Mc = q a^2 (3+nu)/16, unit bending moment at center
256.5369 |  | psi | 1768759.6399 | Pa | sc = 6 Mc/t^2, bending stress at plate center (outer fiber)
3 |  | -- |  |  | FS, Factor of Safety
8.4228 |  | -- |  |  | Safety Margin
See Parker catalog 5705A, Vacuum Sealing
Fluorocarbon (FKM)
Perfluoroelastomer (FFKM)
compare Parker o-ring leak rate with PSI measurements
Parker recommends Gask-O-Seals -- what is this?
Surface finish of 16 rms common, but 32 have been successful
Use leak rate approx. method in section on gasses
Design chart 4-2. If difficult to assemble, use design chart 4-1 and design table 4-1
Vacuum seal glands, design chart 4-7
For viewport seals, dimension I.D. and width (not O.D.)
[7 empty] | 0.63855 | 0.10483 | 0.16183

## volume considerations
O-RING VOLUMETRIC COMPRESSION
standard formulas and design guidance for o-ring compression (i.e. squeeze) are two-dimensional
However, the o-ring is stretched circumferrentially:
a) if the o-ring ID is less than the glad ID (e.g. due to tolerances in fabrication)
b) as the o-ring is comressed and it stretches to a greater diameter
So, one must consider the volume, not just the vertical squeeze and the cross-sectional area
value | ± tol | units | value | ± tol | units | symbol | description
5.359 | 0.035 | in |  |  | mm | Di | o-ring unstretched, un-compressed inner diamter (ID)
5.498 | [5 empty] | Dc | nom, o-ring unstretched, un-compressed, center diamter (ID)
5.459 | [6 empty] | min, o-ring unstretched, un-compressed, center diamter (ID)
5.537 | [6 empty] | max, o-ring unstretched, un-compressed, center diamter (ID)
0.139 | 0.004 | in |  |  | mm | d | nom, o-ring unstretched, un-comressed cross-sectional diameter
0.135 | [6 empty] | min, o-ring unstretched, un-comressed cross-sectional diameter
0.143 | [6 empty] | max, o-ring unstretched, un-comressed cross-sectional diameter
0.015175 |  | in^2 |  |  | mm^2 | Ao | o-ring unstretched, un-compressed cross-sectional area
0.2621 |  | in^3 |  |  | mm^3 | Vo | nom, o-ring unstretched, un-compressed volume
0.24548 | [6 empty] | min, o-ring unstretched, un-compressed volume
0.27937 | [6 empty] | max, o-ring unstretched, un-compressed volume
Consider the case where the gland ID is set equal to the nominal o-ring ID (Parker)
5.359 |  | in |  |  | mm | Dg | gland inner diameter
0.13457 |  | in |  |  | mm | dos | initial cross-section dia, min volume o-ring | -5.4528e-07
0.00042585 |  | in |  |  | mm |  | effective squeeze due to min V with Dg = Do
Consider the case where the gland ID is set equal to the nominal o-ring ID less 15% of the o-ring cross-sectional dia. (Roth)
5.3381 |  | in |  |  | mm | Dg | gland inner diameter
0.13483 |  | in |  |  | mm | dos | initial cross-section dia, min volume o-ring | -6.5231e-07
0.00017294 |  | in |  |  | mm |  | effective squeeze due to min V with Dg
GLAND FILL RATIO
The fill ratio is normally defined on the basis of the cross-sectional area of the o-ring and gland. Consider the volumes instead:
Consider the case where the gland ID is set equal to the nominal o-ring ID (Parker)
72.36% | [5 empty] | k | o-ring/gland fill ratio, area based
0.018 |  | in |  |  |  | g | gap
0.072 |  | in |  |  |  | h | gland depth
0.23301 |  | in |  |  |  | w | gland width
0.020971 |  | in^2 |  |  |  | Ag | gland cross-sectional area
0.2621 |  | in^3 |  |  |  | Vo
0.36842 |  | in^3 |  |  |  | Vg
71.14% | [5 empty] | K | o-ring/gland fill ratio, volume based
1.217% | [6 empty] | error in fill ratio using cross-sectional area

## thread engagement
0.3125 | in | D | major diameter
24 |  | n | number of threads per inch
0.041667 | 1/in | p | pitch
0.058 | in^2 | At | tensile stress area
0.25872 | in | Le | minimum thread engagement length | engineersedge.com
0.3125 |  |  | min, rule of thumb
0.46875 |  |  | ideal, rule of thumb
0.1875
0.15625
0.125
0.17188
8.96
0.14062