LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1100845 	                 LIGO                   February 21, 2012

aLIGO BSC-ISI, Pre-integration Testing report,

Phase II (before cartridge install)

E1100845 – V3

Vincent Lhuillier, Fabrice Matichard

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

Introduction	4

1.	Hardware changes	5

1.	CPS – E1100369	5

2.	GS13 – E1100740	5

3.	L4C – E1100740	6

4.	T240 – E1100740	6

5.	Electronic Inventory	7

6.	Models Modifications	7

7.	Cables – E1100822	7

8.	Mass distribution	7

9.	Hardware changes	7

2.	Basic functionalities just after installing the BSC-ISI on the teststand (After July 4th, 2011)	8

1.	Pressure sensors	8

2.	Offset CPS Unlocked vs locked	8

3.	Range of motion	9

4.	Linearity test	9

5.	Static tests	11

1.	Offset local drive	11

2.	Offset Cartesian drive	12

3.	Transfer functions and Comparison with measurements done in the staging building.	13

4.	Controls work on the ISI before ITMY and FMY	19

1.	Work prior to the control design	19

1.	Calibration filters	19

2.	Complementary filters	21

2.	Control work – Tools - Programs	21

3.	Transfer functions measurements	22

4.	Concatenating the data	23

5.	Symmetrization – Calibration – Idealization of the geophones	23

1.	Actuators symmetrization	23

2.	L4C Symmetrization	24

3.	Symmetrization of the GS13	26

4.	Symmetrization filters	28

5.	Symmetrized transfer functions	29

6.	Calibration check	33

6.	Change of Base – Local to Cartesian - Simulations	34

7.	Digitalization of the symmetrization filters	37

8.	Design of the damping filters of stage 2	37

9.	Design of the damping filters of stage 1	40

10.	Digitalization of the damping filters	42

11.	ASD with Damping loops closed	42

12.	Blending CPS and Geophones signals - Super Sensors	44

13.	Digitalization of the Blend filters	46

14.	Isolation filters – Stage 1 – Z – RX – RY	46

15.	Isolation filters – Stage 2 – Z – RX - RY	47

16.	Isolation filters – Stage 1 – X – Y - RZ	48

17.	Isolation filters – Stage 2 –X – Y - RZ	49

18.	Open loop check	53

19.	Digitalization of the isolation filters	54

20.	ASD of the isolated ISI	54

5.	Test with the QUAD (ITMY) and without FMY	56

6.	Test with the QUAD (ITMY) and FMY before cartridge install	60

1.	Instrument changes	60

2.	Sensors check	61

3.	Mass budget	64

4.	Balancing check	66

5.	Static test – Local Basis	67

6.	Static Test – Cartesian Basis	67

7.	Pressure sensors check	69

8.	Range of motion test	69

9.	Linearity-hysteresis test	70

10.	Transfer function measurements	72

11.	Damping loops	75

7.	Tests after Cartridge install	77

1.	Transfer functions after cartridge install	77

2.	Addition of Vibration absorbers – Viton pads – TMDs	80

1.	Vibration absorbers (0.25”x0.25”xThickness”)	80

2.	TMD (blade 0-1)	81

3.	Viton pads under the stage 2 stack mass	82

4.	The transfer functions comparisons	82

5.	SUS resonances – Payload History	86

6.	Suspension modes	98

8.	Last tests before closing the chamber	99

1.	Spectra	99

2.	Geophones pressure sensors	101

3.	Balancing – CPS Offset locked vs unlocked	102

4.	Static test – Local basis	102

5.	Static Test – Cartesian basis	103

6.	Range of motion – Actuators	104

7.	Linearity Test	105

8.	Transfer functions	105

9.	State of health of the T240 on February 6, 2012	105

10.	Validation	106

Conclusion	107

Introduction

The BSC-ISI testing is performed in three phases:

1) BSC-ISI, Pre-integration Testing, Phase I (post-assembly, in the staging building)

2) BSC-ISI, Pre-integration Testing, Phase II: Final tests done before insertion in the chamber
3) BSC-ISI, Integration Phase Testing: Procedure and results related to the commissioning in the chamber.

The ISI-BSC8 was moved from the Staging building to the LVEA test stand on June 30, 2011.

This document presents the series of tests (Phase II) performed on the ISI-BSC8 (ITMY) in the LVEA. Some tests were done with and without the QUAD and the folding miror. Tests started on July 8, 2011. Damping loops were closed for the first time on July 22, 2011 and Isolation loops on August 4, 2011.

All results are posted on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/H2/ITMY/

The following type of document can be found in the SVN:

- Excell spreadsheet (.xls)

- Data location

- Figures location

- Masses distribution scheme (ppt)

The dome of the chamber was put on February 2, 2012.

# Hardware changes

## CPS – E1100369

Due to flux issues on capacitive position sensors during the production process, six out of twelve installed sensors have been removed. The table below presents the list of the sensors installed during phase II. These sensors have been changed on August 17, 2011.

This table is the final configuration of the table installed on ISI-BSC8 (after August 17, 2011).

| Status | Location | Invoice | Gauge | S/N | Probe | S/N | Master/Slave | Phase | Coarse/Fine |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Installed on Unit 1 | ST1-H1 | 101612 | 8800 | 12504 | 2822-V | 12949 | Slave | 180 | Coarse |
| Installed on Unit 1 | ST1-H2 | 101406 | 8800 | 12305 | 2822-V | 12864 | Slave | 180 | Coarse |
| Installed on Unit 1 | ST1-H3 | 101612 | 8800 | 12512 | 2822-V | 12952 | Slave | 180 | Coarse |
| Installed on Unit 1 | ST1-V1 | 101406 | 8800 | 12312 | 2822-V | 12860 | Master | 0 | Coarse |
| Installed on Unit 1 | ST2-V2 | 101406 | 8800 | 12307 | 2822-V | 12863 | Slave | 0 | Coarse |
| Installed on Unit 1 | ST2-V3 | 101406 | 8800 | 12337 | 2822-V | 12865 | Slave | 0 | Coarse |
| Installed on Unit 1 | ST2-H1 | 101406 | 8800 | 12380 | 2822-V | 12867 | Slave | 180 | Fine |
| Installed on Unit 1 | ST2-H2 | 101612 | 8800 | 12378 | 2822-V | 12940 | Slave | 180 | Fine |
| Installed on Unit 1 | ST2-H3 | 101612 | 8800 | 12423 | 2822-V | 12938 | Slave | 180 | Fine |
| Installed on Unit 1 | ST2-V1 | 101612 | 8800 | 12413 | 2822-V | 12939 | Slave | 0 | Fine |
| Installed on Unit 1 | ST2-V2 | 101612 | 8800 | 12381 | 2822-V | 12950 | Slave | 0 | Fine |
| Installed on Unit 1 | ST2-V3 | 101406 | 8800 | 12373 | 2822-V | 12868 | Slave | 0 | Fine |

Table 1 - CPS sensors location after removal of the questionable ones

## GS13 – E1100740

All GS13 were changed during phase II testing. One GS13 on was opened during phase I testing and the 5 other GS13s initially installed were mounted with the wrong screws on the electrical feedthroughs.

The final installation locations are reported in the table below:

| location GS13 | Serial Number | POD | Cleaning status | Comments |
| --- | --- | --- | --- | --- |
| H1 | Not noted at LLO | 33 | Class A | Initially installed - POD 68 - S/N 822 - Replacement on September 30, 2011 |
| H2 | 801 | 31 | Class A | Initially installed - POD 85 - S/N 842 - Replacement on September 30, 2011 |
| H3 | 783 | 16 | Class A | Initially installed - POD 89 - S/N 841 - Replacement on September 30, 2011 |
| V1 | 740 | 79 | Class A | Initially installed - POD 76 - S/N 728 - Replacement on September 30, 2011 |
| V2 | 739 | 94 | Class A | Initially installed - POD 84 - S/N703 - Replacement on September 30, 2011 |
| V3 | 677 | 26 | Class A | Initially installed - POD 98 - S/N 719 - Replacement by the 26 on August 11, 2011 |

Table 2 - GS13 Location

Note: Control work was realized with the set of geophones with the following serial numbers: 822, 842 841 728, 703 and 719.

## L4C – E1100740

The L4Cs have not been changed. The X tag on L4C S/N 815 pod 22 raised concerns but the pod has been leaked checked. The L4C S/N 815 pod 22 stays on the assembly.

The final installation locations are reported in the table below:

| L4C location | Serial Number | POD | Adaptor | Cleaning status | Comments |
| --- | --- | --- | --- | --- | --- |
| H1 | 815 | 22 |  | Class A | X tag but fine |
| H2 | 1094 | 152 |  | Class A |  |
| H3 | 968 | 150 |  | Class A |  |
| V1 | 1068 | 155 |  | Class A |  |
| V2 | 1070 | 151 |  | Class A |  |
| V3 | 811 | 153 |  | Class A |  |

Table 3 - L4C location

## T240 – E1100740

After investigations on the pressure sensors of the T240, it appears that the voltage regulators of the pressure sensors electronic board were failing. The high failure rate on this board is due to the voltage peak (created by the inductance of the cable) when the T240 are unplugged without turning off the interface chassis. Consequently, two of the T240 has been replaced.

The final installation locations are reported in the table below:

| location T240 | Serial Number | POD | Cleaning status | Comments |
| --- | --- | --- | --- | --- |
| Corner 1 | 112 | 32 | Class A | Initially installed - POD 12 - S/N 138 - Replacement on September 30, 2011 |
| Corner 2 | 118 | 20 | Class A | Initially installed - POD 26 - S/N 133 - Replacement on September 30, 2011 |
| Corner 3 | 101 | 19 | Class A | Initially installed - POD 36 - S/N 109 - Replacement on September 30, 2011 |

Table 4 - T240 Location

Note: Control work was realized with the set of geophones with the following serial numbers: 138, 133 and 109.

## Electronic Inventory

This table reports the electronic equipment used in the LVEA.

| Hardware | Ligo reference | S/N |
| --- | --- | --- |
| Interface Chassis Pod 1 | D1002432 | S1102228 |
| Interface Chassis Pod 2 | D1002433 | S1102230 |
| Interface Chassis Pod 3 | D1002434 | S1102229 |
| Anti-alliasing Chassis | D1002693 | S1103404 |
| Anti-alliasing Chassis | D1002693 | S1103405 |
| Anti-image Chassis | D1002693 | S1103402 |
| Binary Input Chassis | D1001726 | S1101284 |
|  | D1001726 | S1101280 |
| Binary Output Chassis | D1001728 | S1101319 |
| T240 Interface Pod 1 | D1002694 | S1103177 |
| T240 Interface Pod 2 | D1002694 | S1103181 |
| T240 Interface Pod 3 | D1002694 | S1103180 |
| Anti-image Chassis | D1000305 | S1103502 |
| Coil driver Pod 1 | D0902744 | S1103564 |
| Coil driver Pod 2 | D0902744 | S1103563 |
| Coil driver Pod 3 | D0902744 | S1103356 |

Table 5 - Electronic inventory

The coil drivers have been modified (After modifications, BIO switches are in the upper state when they are working fine – fuse ok – no overheating). The changes were done on September 2, 2011.

## Models Modifications

The BSC-ISI model has been updated to introduce the SUS watchdogs and fix the BIO inputs after the hardware medication (coil driver). Pressure sensors alarm was added. Channels list used by RCG 2.4 to build the .ini file was also added.

## Cables – E1100822

The class B cables used for testing have been replaced on October 7, 2011 by the final class A ones as specified in the document E1100770.

## Mass distribution

The mass distribution evolved a few times during this phase of testing (masses on top of stage 2, cookie cutters, QUAD, FMY)

## Hardware changes

One locker has been re shimmed around July 14, 2011.

# Basic functionalities just after installing the BSC-ISI on the teststand (After July 4th, 2011)

## Pressure sensors

The pressure sensors of 2 T240 (corner 2 and 3) are not working. T240s were later replaced (cf inventory)

## Offset CPS Unlocked vs locked

The motion of the ISI between the unlocked and the locked position is measured using the CPS. The position of the two stages in the two different configurations is reported in the table below.

NOTE:

It appears that when the system is locked, important variations have been observed in the CPS when the ISI is balanced and when the ISI is not balanced. After transporting the ISI from the staging building to the LVEA, the gap of the capacitive position sensors have been reset. The readout of the CPS in the locked position concerned the assembly team.

By regapping the CPS, the reference position defined in the staging building (with a symmetric payload – masses on top of stage 2) has been lost.

After rebalancing, the level of the optic table has been rechecked.

Measurement realized on July 13, 2011

The results of this test can be found in the SVN at:

\seisvn\seismic\BSC-ISI\H2\ITMY\Data\Static_Tests

- LHO_ISI_BSC_Unlocked_Final_Tuning_2011_07_13.mat

- LHO_ISI_BSC_Locked_Final_Tuning_2011_07_13.mat

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | -220.1 | 7.0 | -128.2 | 16.5 | -92 | -0.11 |
| ST1 - H2 | -61.5 | 4.7 | 378.5 | 18.4 | -440 | -0.52 |
| ST1 - H3 | 0.2 | 4.0 | 592.7 | 18.1 | -593 | -0.71 |
| ST1 - V1 | -32.6 | 4.3 | 893.2 | 42.7 | -926 | -1.10 |
| ST1 - V2 | 354.9 | 3.8 | 759.7 | 58.4 | -405 | -0.48 |
| ST1 - V3 | -465.1 | 4.1 | 129.5 | 50.6 | -595 | -0.71 |
| ST2 - H1 | -1015.2 | 14.7 | 357.2 | 42.1 | -1372 | -0.41 |
| ST2 - H2 | -583.3 | 33.2 | -193.1 | 39.7 | -390 | -0.12 |
| ST2 - H3 | 1622.7 | 28.2 | 3772.0 | 47.3 | -2149 | -0.64 |
| ST2 - V1 | 6.3 | 25.3 | -1988.9 | 97.6 | 1995 | 0.59 |
| ST2 - V2 | -818.2 | 107.1 | 823.0 | 144.2 | -1641 | -0.49 |
| ST2 - V3 | 220.8 | 163.8 | -1238.7 | 122.4 | 1459 | 0.43 |

Table 6 - Locked vs Unlocked Position

Test result:	Passed:   X   	Failed:         .

## Range of motion

The range of motion of the table is measured by pushing on the table in a direction collinear to the CPS. The Static tests results can be found on the SVN at:

\seisvn\seismic\BSC-ISI\H2\ITMY\Data\Static_Tests

- LHO_ISI_BSC8_Range_Of_Motion_20110714.mat

| Sensors | Push in positive direction | Push in negative direction | Mil (positive) | Amplitude (count) | Mill |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -8199 | -137 | 8089 | 16288 | 19.4 |
| ST1 - H2 | -7783 | 371 | 8239 | 16022 | 19.1 |
| ST1 - H3 | -7605 | 591 | 8080 | 15685 | 18.7 |
| ST1 - V1 | -12246 | 932 | 14126 | 26372 | 31.4 |
| ST1 - V2 | -12278 | 831 | 13964 | 26242 | 31.2 |
| ST1 - V3 | -12808 | 236 | 13264 | 26072 | 31.0 |
| ST2 - H1 | -9015 | 349 | 9715 | 18730 | 5.6 |
| ST2 - H2 | -9571 | -203 | 9177 | 18747 | 5.6 |
| ST2 - H3 | -5351 | 3810 | 12972 | 18323 | 5.5 |
| ST2 - V1 | -13391 | -1850 | 9688 | 23079 | 6.9 |
| ST2 - V2 | -10428 | 994 | 12389 | 22817 | 6.8 |
| ST2 - V3 | -12497 | -925 | 10641 | 23138 | 6.9 |

Table 7 - Range of motion - Actuator drive in the LVEA

Note 1: Due to temporary cables in-air cables (3 extensions), the readout of ST1 horizontal CPS has a gain of 0.5. We realized that he CPS extension cable had pin1 connected to the shield (Warning: Grounded when plugged on the interface chassis but not grounded when plugged to a breakout board. The 2 sockets of a breakout board are not connected). The range of motion measured by the ST1 horizontal CPS has to be multiplied by 2.

Note 2 : Due to longer in-air cables (Test-stand to the flanges) in the LVEA than in the staging building, the voltage drop creates a fall of the actuators output force for a same output voltage on the DAC (-15% on stage 1 actuators and – 9% on stage 2actuators). These changes can also be seen in the linearity test.

Note 3: The results are similar to those measured in the staging building.

Test result:	Passed:   X   	Failed:         .

## Linearity test

The test was performed on July 19, 2011 after replacing the bad extension cables used for the CPS.

The data of the linearity test can be found on the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Linearity_Test/

- LHO_ISI_BSC8_Linearity_test_20110719.mat

The figures of the linearity test can be found on the SVN at:

seismic\BSC-ISI\H2\ITMY\Data\Figures\Linearity_Test\

- LHO_ISI_BSC8_Linearity_test_20110719.fig

Figure 1 - Linearity test - NO QUAD - No FMY - In the LVEA

|  |  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- | --- |
| Stage 1 | ST1 - H1 | 0.548 | -284 | 0.5432 | 0.85 |
|  | ST1 - H2 | 0.539 | 534 |  | -0.75 |
|  | ST1 - H3 | 0.543 | 974 |  | -0.10 |
|  | ST1 - V1 | 0.440 | 929 | 0.4380 | 0.55 |
|  | ST1 - V2 | 0.439 | 855 |  | 0.14 |
|  | ST1 - V3 | 0.435 | 296 |  | -0.69 |
| Stage 2 | ST2 - H1 | 0.313 | 325 | 0.3110 | 0.73 |
|  | ST2 - H2 | 0.314 | -223 |  | 1.09 |
|  | ST2 - H3 | 0.305 | 3723 |  | -1.82 |
|  | ST2 - V1 | 0.386 | -1851 | 0.3855 | 0.22 |
|  | ST2 - V2 | 0.383 | 945 |  | -0.78 |
|  | ST2 - V3 | 0.388 | -878 |  | 0.56 |

Table 8 - Slope – Offset Linearity test

Note:

The average slopes are lower than those measured in the staging building due to the voltage drop created by the longer cables.

Test result:	Passed:   X   	Failed:         .

## Static tests

### Offset local drive

The test was performed on July 18, 2011.

Note: Due to longer cable s, offsets measured by CPS for a 7000 count drive are slightly lower than offsets measured in the staging building.

Results of this test can be found in the SVN at:

\seisvn\seismic\BSC-ISI\H2\ITMY\Data\Static_Tests

- LHO_ISI_BSC8_Offset_Local_Drive_20110718.mat

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 3839 | 1536 | 1557 | 19 | -6 | -11 |
|  | ST1 - H2 | 1539 | 3774 | 1532 | -5 | 11 | -1 |
|  | ST1 - H3 | 1543 | 1535 | 3811 | 0 | -21 | 24 |
|  | ST1 - V1 | 74 | -134 | 106 | 3095 | -570 | -555 |
|  | ST1 - V2 | 84 | 32 | -140 | -553 | 3076 | -571 |
|  | ST1 - V3 | -156 | 84 | 38 | -565 | -538 | 3051 |

Table 9 - Static Tests – Local to Local - Stage 1

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2192 | 327 | 347 | -12 | -18 | 21 |
|  | ST2 - H2 | 326 | 2187 | 349 | -10 | -2 | -6 |
|  | ST2 - H3 | 326 | 311 | 2162 | -25 | 9 | 6 |
|  | ST2 - V1 | 69 | 114 | -182 | 2717 | 283 | -12 |
|  | ST2 - V2 | -202 | 51 | 123 | -27 | 2676 | 287 |
|  | ST2 - V3 | 120 | -209 | 97 | 294 | -7 | 2705 |

Table 10 - Static Tests – Local to Local - Stage 2

Test result:	Passed:   X   	Failed:         .

### Offset Cartesian drive

Results of this test can be found in the SVN at:

\seisvn\seismic\BSC-ISI\H2\ITMY\Data\Static_Tests

- LHO_ISI_BSC8_Offset_Cartesian_Drive_20110719.mat

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST1 - X | ST1 - Y | ST1 - Z | ST1 - RY | ST1 - RY | ST1 - RZ |
| Actuators | ST1 - X | 1529 | -1 | 12 | -8 | -20 | 18 |
|  | ST1 - Y | -28 | 1504 | 2 | 19 | -21 | -9 |
|  | ST1 - Z | 6 | -13 | 652 | 1 | -17 | -9 |
|  | ST1 - RX | -2 | 315 | -2 | 2656 | -5 | -15 |
|  | ST1 - RY | -327 | 23 | 7 | 5 | 2635 | -6 |
|  | ST1 - RZ | 14 | -8 | 18 | -3 | -2 | 2900 |

Table 11 - Static Tests - Cartesian to Cartesian - Stage 1

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - X | ST2 - Y | ST2 - Z | ST2 - RY | ST2 - RY | ST2 - RZ |
| Actuators | ST2 - X | 1234 | 8 | -6 | -3 | 35 | -14 |
|  | ST2 - Y | 11 | 1230 | -12 | -21 | -8 | 19 |
|  | ST2 - Z | 0 | 2 | 984 | -9 | 13 | -3 |
|  | ST2 - RX | 3 | -12 | -21 | 3939 | 12 | -5 |
|  | ST2 - RY | 9 | 9 | 5 | 24 | 3973 | 5 |
|  | ST2 - RZ | 1 | 13 | 3 | 0 | -6 | 2351 |

Table 12 - Static Tests - Cartesian to Cartesian - Stage 2

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - X | 1547 | -750 | -745 | 27 | -1 | -15 |
|  | ST1 - Y | 41 | 1286 | -1312 | -5 | 13 | -12 |
|  | ST1 - Z | -1 | -18 | -8 | 656 | 645 | 638 |
|  | ST1 - RX | -15 | 123 | -149 | -2548 | 2161 | 394 |
|  | ST1 - RY | -180 | 75 | 67 | -1006 | -1692 | 2709 |
|  | ST1 - RZ | 2858 | 2773 | 2787 | 16 | -6 | 17 |

Table 13 - Static Tests - Cartesian to local - Stage 2

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - X | 657 | -1254 | 610 | -15 | -7 | 64 |
|  | ST2 - Y | 1050 | 4 | -1049 | 30 | -14 | 30 |
|  | ST2 - Z | -17 | 4 | 5 | 991 | 971 | 1018 |
|  | ST2 - RX | -270 | 4 | 261 | -2295 | 2316 | -6 |
|  | ST2 - RY | 181 | -304 | 157 | -1403 | -1294 | 2738 |
|  | ST2 - RZ | 1628 | 1647 | 1610 | 20 | 2 | 37 |

Table 14 - Static Tests - Cartesian to local - Stage 2

Test result:	Passed:   X   	Failed:         .

# Transfer functions and Comparison with measurements done in the staging building.

The parameters for the measurements in the LVEA are slightly different from those in the staging building. It may be explained by the relative proximity of the cleanroom fans as well as the weaker power delivered by the actuators due to the longer in-air cables.

These parameters are summed up in the table below:

| Stage 1 | Section | Freq min | Freq max | Fres | Amplitude (H –V) | Nrep | Time Stage 1 (min) |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 1 | 0.01 | 0.1 | 0.01 | 2000 - 2000 | 3 | 32 |
|  | 2 | 0.1 | 0.7 | 0.02 | 140 - 200 | 25 | 127 |
|  | 3 | 0.7 | 10 | 0.05 | 160 – 220 | 75 | 154 |
|  | 4 | 10 | 100 | 0.1 | 120 – 120 | 50 | 53 |
|  | 5 | 100 | 500 | 0.2 | 130 - 130 | 50 | 53 |
|  | 6 | 500 | 1000 | 0.25 | 100 - 100 | 50 | 53 |
|  |  |  |  |  |  |  | 472 |
| Stage 2 | Section | Freq min | Freq max | Fres | Amplitude (H-V) | Nrep | Time Stage 2 (min) |
|  | 1 | 0.01 | 0.1 | 0.01 | 2000 - 2000 | 3 | 32 |
|  | 2 | 0.1 | 0.7 | 0.02 | 500 - 550 | 25 | 127 |
|  | 3 | 0.7 | 10 | 0.05 | 450 - 400 | 75 | 154 |
|  | 4 | 10 | 100 | 0.1 | 180 -180 | 50 | 53 |
|  | 5 | 100 | 500 | 0.2 | 200 – 200 | 50 | 53 |
|  | 6 | 500 | 1000 | 0.25 | 150 - 150 | 50 | 53 |
|  |  |  |  |  |  |  | 472 |
|  |  |  |  |  |  | Overall ID | 944 |

Measurements data can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Measurements/Undamped:

- LHO_ISI_BSC8_Data_L2L_10mHz_100mHz_ST1_ST2_20110723-145046.mat

- LHO_ISI_BSC8_Data_L2L_100mHz_700mHz_ST1_ST2_20110723-101732.mat

- LHO_ISI_BSC8_Data_L2L_700mHz_10Hz_ST1_ST2_20110723-050631.mat

- LHO_ISI_BSC8_Data_L2L_10Hz_100Hz_ST1_ST2_20110722-211830.mat

- LHO_ISI_BSC8_Data_L2L_100Hz_1000Hz_ST1_ST2_20110722-184135.mat

Once the data are processed, they can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/X1/Data/BSC8/Transfer_Functions/Simulations/Undamped

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2011_06_22.mat

The script used to compare transfer function can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/H2/ITMY/Scripts/Misc

- Comparison_TF_L2L_LHO_ISI_BSC8.m

The figure that shows the comparison between the transfer functions of the staging building and the LVEA are located in the SVN at:

SeiSVN/seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Measurements/
Comparison/L2L/

- LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_H_to_ST1_CPS_H_20110622_vs_20110722.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_H_to_ST1_L4C_H_20110622_vs_20110722.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_V_to_ST1_CPS_V_20110622_vs_20110722.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_V_to_ST1_L4C_V_20110622_vs_20110722.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_H_to_ST2_CPS_H_20110622_vs_20110722.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_H_to_ST2_GS13_H_20110622_vs_20110722.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_V_to_ST1_CPS_V_20110622_vs_20110722.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_V_to_ST2_GS13_V_20110622_vs_20110722.fig

Note 1:  The transfer functions are measured from the Output filters bank excitation point to the input (IN1) of the input filters bank. The transfer functions presented below are raw transfer functions without any electronic compensation.

Note 2: The L4Cs are out of phase (should be -90 before 1Hz).  A minus sign is added in the calibration filters that convert count to nm/s.

Note 3: On the ST1-ACT-H to ST1-CPS-H transfer functions, we can see the first resonances of the LVEA teststand at 19.6Hz, 21.4Hz and 29.6Hz. The resonances of the staging building teststand are at 33.4Hz, 57.3Hz and 69.4Hz. The teststand in the LVEA has longer legs that the teststand in the staging building.

Note 4: The transfer functions measured in the LVEA and in the staging building are very similar (at LF and HF)

Figure 2 - Transfer functions comparison - ST1 ACT V to ST1 CPS V

Figure 3 - Transfer functions comparison - ST1 ACT V to ST1 CPS V

Figure 4 - Transfer functions comparison - ST1 ACT H to ST1 L4C H

Figure 5 - Transfer functions comparison - ST1 L4C V to ST1 L4C V

Figure 6 - Transfer functions comparison - ST2 ACT H to ST2 CPS H

Figure 7 - Transfer functions comparison - ST2 ACT V to ST2 CPS V

Figure 8 - Transfer functions comparison – ST2 ACT H to ST2 GS13 H

Figure 9 - Transfer functions comparison – ST2 ACT V to ST2 GS13 V

Test result:	Passed:   X   	Failed:         .

# Controls work on the ISI before ITMY and FMY

## Work prior to the control design

### Calibration filters

The scripts used to build the calibration filters used to display calibrated ASD are located:

seismic/BSC-ISI/Common/Calibration_BSC_ISI/aLIGO/

- aLIGO_BSC_ISI_Calibration.m

This script creates a mat file in which the calibration filters are stored.

- aLIGO_BSC_ISI_Calibration.mat

aLIGO_BSC_ISI_Calibration.mat is called by scripts that calculate the calibrated ASD

The document E1100524 describes the input, output filters and the BIO switches that need to be engaged when the BSC-ISI is controlled. The input filters account for: 
- the interface chassis electronic 
- the sensor calibration 
- the state of the BIO switches 
- sign inversion if the instrument are badly connected

The output filters account for: 
- the dynamic of the coil driver 
- the dynamic of the actuators 

The state of the BIO switches is also defined in E1100524.

All elements related to the input and out put filters are located in the SVN at the following location:

seismic/BSC-ISI/Common/Compensation_Filters_BSC_ISI/aLIGO/

The scripts to create the filters are the following.

- aLIGO_BSC_ISI_Input_Filters.m

- aLIGO_BSC_ISI_Output_Filters.m

The filters are stored:

- aLIGO_BSC_ISI_Input_Filters.mat

- aLIGO_BSC_ISI_Output_Filters.mat

The scripts that digitalize the filters are:

- aLIGO_BSC_ISI_Input_Filters_Digitalization.m

- aLIGO_BSC_ISI_Ouput_Filters_Digitalization.m

The text files that contain the coefficients of the input and the output filters are:

- aLIGO_BSC_ISI_Input_Filters_27-Jul-2011.txt

- aLIGO_BSC_ISI_Output_Filters_27-Jul-2011.txt

Figures that show the input and the output filters are:

- aLIGO_BSC_ISI_Input_Filters_2011_06_29.fig

- aLIGO_BSC_ISI_Output_Filters_2011_06_29.fig

The figures below shows the filters that are used to compensated for the electronics and the actuators.

Figure 10 - Input filters after gain calibration (1ct/nm) and (1ct/(nm/s))

Figure 11 - Output filters

### Complementary filters

The complementary filters that have been used to blend the capacitive position sensors and the geophones are located in the SVN at:

seismic/BSC-ISI/Common/Complementary_Filters_BSC_ISI/aLIGO/

- aLIGO_BSC_ISI_Complementary_Filters.mat

These filters are very basic. They will be retrofitted later.

## Control work – Tools - Programs

The design and the implementation of the control on the BSC-ISI are segmented into 15 steps. Each one of these step is associated with a main scripts.

For instance, Step_1_TF_L2L_10mHz_1000Hz_LHO_ISI_BSC8.m compiles the transfer functions into a unique structure (frd).

The scripts mentioned below may be subject with modifications.

These control scripts are located on the SVN at:

\seisvn\seismic\BSC-ISI\H2\ITMY\Scripts\Control_Scripts

- Step_0_Run_all_LHO_ISI_BSC8.m

- Step_1_TF_L2L_10mHz_1000Hz_LHO_ISI_BSC8.m

- Step_2_Symmetrization_Calibration_LHO_ISI_BSC8.m

- Step_3_TF_Basis_Change_L2C_LHO_ISI_BSC8.m

- Step_4_C2D_Sym_Filters_LHO_ISI_BSC8.m

- Step_5_Damping_Loops_ST2_LHO_ISI_BSC8.m

- Step_6_Damping_Loops_ST1_LHO_ISI_BSC8.m

- Step_7_C2D_Damping_Filters_LHO_ISI_BSC8.m

- Step_8_Blend_Filters_LHO_ISI_BSC8.m

- Step_9_C2D_Blend_Filters_LHO_ISI_BSC8.m

- Step_10_Isolation_Loops_ST1_Z_RX_RY_LHO_ISI_BSC8.m

- Step_11_Isolation_Loops_ST2_Z_RX_RY_LHO_ISI_BSC8.m

- Step_12_Isolation_Loops_ST1_X_Y_RZ_LHO_ISI_BSC8.m

- Step_13_Isolation_Loops_ST2_X_Y_RZ_LHO_ISI_BSC8.m

- Step_14_Open_Loop_check_LHO_ISI_BSC8.m

- Step_15_C2D_Isolation_Filters_LHO_ISI_BSC8.m

The main scripts mentioned above call lots of other functions that are common to every BSC-ISIs located in the SVN at:

\seisvn\seismic\BSC-ISI\Common\Control_Scripts_Functions_3_BSC_ISI\

- Plot_Comparison_TF_L2L_BSC_ISI.m

- Plot_Damping_Filters_MIMO_BSC_ISI.m

- Plot_Damping_Filters_MIMO_BSC_ISI_Presentation.m

- Plot_Damping_Filters_ST1_SISO_BSC_ISI.m

- Plot_Damping_Filters_ST2_SISO_BSC_ISI.m

- Plot_Full_MIMO_Suppression_BSC_ISI.m

- Plot_Isolation_Full_MIMO_BSC_ISI.m

- Plot_L4C_Fitt_BSC_ISI.m

- Plot_L4C_Fitt_all_BSC_ISI.m

- Plot_L4C_Fitt_report_display_BSC_ISI.m

- Plot_L4C_GS13_Fitt_BSC_ISI.m

- Plot_L4C_GS13_Fitt_all_BSC_ISI.m

- Plot_L4C_GS13_Fitt_with_sys_BSC_ISI.m

- Plot_L4C_GS13_Fitt_without_sys_BSC_ISI.m

- Plot_L4C_GS13_Symmetrization_Filters_BSC_ISI.m

- Plot_L4C_without_sys_L2L_vs_C2L_BSC_ISI.m

- Plot_Open_Loops_check.m

- Plot_Super_Sensors_Isolation_MIMO_BSC_ISI.m

- Plot_Super_Sensors_Isolation_SISO_BSC_ISI.m

- Plot_Super_Sensors_MIMO_BSC_ISI.m

- Plot_Super_Sensors_SISO_BSC_ISI.m

- Plot_TF_C2C_BSC_ISI.m

- Plot_TF_C2C_Fit_BSC_ISI.m

- Plot_TF_C2C_Fit_all_BSC_ISI.m

- Plot_TF_C2C_meas_BSC_ISI.m

- Plot_TF_L2L_BSC_ISI.m

- Plot_TF_L2L_Comparison_BSC_ISI.m

- Plot_TF_L4C_GS13_Symmetrization_BSC_ISI.m

- Plot_TF_L4C_GS13_Symmetrization_all_BSC_ISI.m

## Transfer functions measurements

The transfer functions are measured in the local basis from the excitation point in the output filters bank and the input IN1 of the input filters bank.

Note: The transfer functions are measured using a strong drive in order to reduce the measurements time (~13h). Due to the important drive, the T240s saturate. No work will be done will be done with the T240 at low frequency using this set of measurements.

The transfer functions used to design the controller have been measured on July 27, 2011.

## Concatenating the data

This step consists in concatenating the data into a unique frd. It also saves a parameters file that is updated throughout the design of the controllers.

The concatenated data can be found in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Simulations/Undamped/

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2011_07_27.mat

The parameter files can be found in the SVN at :

seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Simulations/Parameters/

- Parameters_LHO_ISI_BSC8_2011_07_27.mat

## Symmetrization – Calibration – Idealization of the geophones

The symmetrization step realizes the following tasks:

- Correct the actuator signs (shouldn’t be corrected if the actuators have been wired properly in the staging building – but here just in case)

- Add an actuator symmetrization gain using the CPSs as references

- Apply the Input filters (Calibration gain, whitening, BIO compensation)

- Correct the sign of the geophones (L4C)

- Symmetrize and “idealize” geophones

- Recalculate the transfer functions once all the input filters are applied. The symmetrized transfer functions are calculated from the excitation point of the output filters bank and the output point of the input filters bank.

### Actuators symmetrization

The actuator symmetrization gains are reported in the table below:

| Actuator | Gain |
| --- | --- |
| ST1 ACT H1 | 1.009 |
| ST1 ACT H2 | 1.000 |
| ST1 ACT H3 | 0.992 |
| ST1 ACT V1 | 0.989 |
| ST1 ACT V2 | 1.002 |
| ST1 ACT V3 | 1.010 |
| ST2 ACT H1 | 0.990 |
| ST2 ACT H2 | 0.990 |
| ST2 ACT H3 | 1.022 |
| ST2 ACT V1 | 1.001 |
| ST2 ACT V2 | 1.008 |
| ST2 ACT V3 | 0.991 |

Table 15 - Actuator symmetrization gains

Note: Symmetrization gains of the actuators are within 2% and with a positive sign.

### L4C Symmetrization

The L4C has a low Q and its internal resonance is in the vicinity of the BSC-ISI first suspension mode. Consequently, identifying the L4C dynamic is tricky. But, the BSC-ISI motions are also sensed by CPS. At low frequency, when driving the actuators, the platform motion of the ISI is not dominated by the ground motion or the air turbulence in the clean room. CPS of stage 1 and L4C observe approximately the same motion (one is relative, one is absolute). Consequently, the L4C transfer functions (the instrument itself) can be calculated by removing the BSC-ISI response to the measured L4C transfer functions using the CPS transfer functions. The Cartesian to local transfer functions are preferred to local to local because the CPS and the L4C are not collocated (on aLIGO BSC-ISI). Next, the L4C transfer function (the instrument itself) can be fitted.

At low frequency, we can estimate the transfer function of the instrument (L4CSensor) using:

Symmetrization filters gains: The symmetrization filters gains are computed such that the symmetrized transfer functions (H1-H2-H3 & V1-V2-V3) can be superposed to the “averaged initial transfer function” of H1-H2-H3 & V1-V2-V3.

Curve fitting the L4C after removing the BSC-ISI response

Figure 12 - Fitting the Horizontal L4C on corner 1

Figures that show the symmetrization of the L4C are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Misc/

- LHO_ISI_BSC8_Curve_Fitting_L4C_H1.fig

- LHO_ISI_BSC8_Symmetrized_Idealized_TF_ST1_ACT_H1_ST1_L4C_H1.fig

The symmetrized-idealized-calibrated transfer functions can be calculated by applying the calibration gain, the whitening and symmetrization filters to the raw transfer functions.

Note 1: The L4C has a pair of complex conjugate poles at 1Hz with a phase of 74°.

Note 2: Due to the idealization of the geophones (L4C and GS13) by 3 zeros at 0Hz and a pair of complex conjugate poles at 1Hz with a phase of 45°, the symmetrized-idealized-calibrated transfer functions are very different from the raw transfer function around 1Hz.

Note 3: A minus sign is added in the symmetrization filters to account for the bad phase measured on the raw L4C transfer functions.

Figure 13 - Applying the input filters to raw transfer functions

### Symmetrization of the GS13

The identification of the L4C and the GS13 cannot be done the same way. The identification of the GS13 is done using the raw transfer functions.

In section 7.5, you can notice that the nature and the location of the payload (on top of stage 2 or hanging from the optic table) strongly affect the dynamic of the ISI and not only at high frequencies. When the QUAD (only) was installed, the first suspension resonance of the ISI was measured at 1.15Hz instead of 1Hz with the stack of masses on top of stage 2.

The program is written to work with a real payload (first resonance at 1.15Hz). In that case it is quite easy to fit the GS13s that have a pole at 1Hz with a phase of 83degree (sharp bump before the first ISI suspension mode). In the present case (first ISI resonance at 1Hz), the commissioner tweaked a little bit (5 minutes for 6 geophones) to fit the GS13s.

The following figures present the fitting realized on the GS13 transfer functions and the recalculated transfer functions after applying the input filters.

These figures are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Misc/

- LHO_ISI_BSC8_Curve_Fitting_GS13_H1.fig

- LHO_ISI_BSC8_Curve_Fitting_GS13_V1.fig

- LHO_ISI_BSC8_Symmetrized_Idealized_TF_ST1_ACT_H1_ST2_GS13_H1.fig

Note 1: The GS13 has a pair of complex conjugate poles at 1Hz with a phase of 83°.

Note 2: When driving horizontally the stage 2 horizontal actuator in 1 corner, Stage 2 tilts. The tilt is observed by the horizontal GS13 below 100mHz.

Note 3: It is difficult to consider the GS13 well fitted (example in figure 14). But, the recalculated symmetrized-idealized transfer functions (in IV.5.5) confirm the good fitting of the GS13s.

The two figures below present the fitting step on the horizontal and the vertical GS13 in corner 1.

Figure 14 - Curve fitting ST2 GS13 H1 – Transfer functions - ST2 ACT H1 to ST2 GS13 H1

Table 16 - Curve fitting ST2 GS13 H1 – Transfer functions - ST2 ACT V1 to ST2 GS13 V1

Figure 15 - Symmetrized transfer functions

### Symmetrization filters

The symmetrization filters are presented below:

% Horizontal L4C fitting filters

Fitt_L4C_GS13(1) = zpk(-2*pi*[0 0 0],-2*pi*pair(1.03,75),1);

Fitt_L4C_GS13(2) = zpk(-2*pi*[0 0 0],-2*pi*pair(.96,73),1);

Fitt_L4C_GS13(3) = zpk(-2*pi*[0 0 0],-2*pi*pair(0.98,74),1);

% Vertical L4C fitting filters

Fitt_L4C_GS13(4) = zpk(-2*pi*[0 0 0],-2*pi*pair(1.07,76),1);

Fitt_L4C_GS13(5) = zpk(-2*pi*[0 0 0],-2*pi*pair(1.07,78),1);

Fitt_L4C_GS13(6) = zpk(-2*pi*[0 0 0],-2*pi*pair(1.02,74),1);

% Horizontal GS13 fitting filters

Fitt_L4C_GS13(7) = zpk(-2*pi*[0 0 0],-2*pi*pair(1.01,83),1);

Fitt_L4C_GS13(8) = zpk(-2*pi*[0 0 0],-2*pi*pair(0.96,85),1);

Fitt_L4C_GS13(9) = zpk(-2*pi*[0 0 0],-2*pi*pair(1.06,83),1);

% Vertical GS13 fitting filters

Fitt_L4C_GS13(10) = zpk(-2*pi*[0 0 0],-2*pi*pair(1.05,83),1);

Fitt_L4C_GS13(11) = zpk(-2*pi*[0 0 0],-2*pi*pair(0.99,83),1);

Fitt_L4C_GS13(12) = zpk(-2*pi*[0 0 0],-2*pi*pair(0.99,83),1);

Note 1: The measured resonances of the L4C are measured between 0.96Hz and 1.07Hz. The measured resonances can be considered very accurate since the identification is realized using the Cartesian to Local measurements after extracting the ISI dynamic

Note 2: The measured resonances of the GS13 are measured between 0.96Hz and 1.07Hz. Given the quality of the symmetrization, the measured GS13 internal resonance can be considered accurate.

The figure that shows the symmetrization filters are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Misc/

- LHO_ISI_BSC8_Symmetrization_Filters.fig

The continuous symmetrization-calibration-idealization filters are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Filters/Continuous/

- LHO_ISI_BSC8_Symmetrization_Filters_2011_07_27.mat

Figure 16 - Symmetrization filters

### Symmetrized transfer functions

The figures of the symmetrized-idealized-calibrated transfer functions presented hereinafter can be found in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Undamped/

- LHO_ISI_BSC8_TF_L2L_Symmetrized_from_ST1_ACT_to_ST1_CPS_2011_07_27.fig

- LHO_ISI_BSC8_TF_L2L_Symmetrized_from_ST1_ACT_to_ST1_L4C_2011_07_27.fig

- LHO_ISI_BSC8_TF_L2L_Symmetrized_from_ST2_ACT_to_ST2_CPS_2011_07_27.fig

- LHO_ISI_BSC8_TF_L2L_Symmetrized_from_ST2_ACT_to_ST2_GS13_2011_07_27.fig

The symmetrized transfer functions (data) can be found in the SVN at:
seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Simulations/Undamped/

- LHO_ISI_BSC8_TF_C2C_Sym_10mHz_1000Hz_2011_07_27.mat

Note : The sensor electronics are also compensated. For instance, in the CPS transfer functions, the zero at 0.14Hz and the pole at 1.4Hz have been compensated.

Figure 17 - Symmetrized transfer functions - ST1 ACT to ST1 CPS

Figure 18 - Symmetrized transfer functions – ST2 ACT to ST2 CPS

Figure 19 - Symmetrized transfer functions – ST1 ACT to ST1 L4C

Figure 20 - Symmetrized transfer functions - ST2 ACT to ST2 GS13

### Calibration check

The goal of this step is to verify that the input filters (symmetrization, calibration, idealization and the whitening filters) are correctly implemented.

The CPS transfer functions and the Geophones transfer functions are compared once the geophones are inverted.

The figures of the calibration check can be found in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Misc/

- LHO_ISI_BSC8_Curve_Fitting_L4C_H1.fig

- LHO_ISI_BSC8_Curve_Fitting_GS13_H1.fig

Figure 21 - Calibration check - L4C Local basis

Note 1: The stage 1 CPS transfer functions and the L4C transfer functions with the inverted geophone matches very good below 10 Hz. Since the ISI is not dominated by the ground motion when transfer functions are measured, the CPS of stage 0-1 and the L4C are sensing the same motion.

Note 2: Below the first suspension resonance, the CPS of stage 1-2 and the GS13 should sense the same motion. That is not the case. Some investigations are ongoing to find out why there is a slight gain between the CPS transfer functions and the GS13 transfer functions with the geophone inversion.

Note 3: The geophones and the CPS are not recorded at the same sample rate (CPS: 1024Hz – L4C-GS13: 2048Hz)

Figure 22 - Calibration check - Local GS13

## Change of Base – Local to Cartesian - Simulations

The change of base matrices are located in the SVN at:

seismic/BSC-ISI/Common/Base_Change_BSC_ISI/aLIGO/Y/:

- aLIGO_BSC_ISI_Matrices_Y_Direction.mat

The simulated Cartesian to Cartesian transfer functions (calculated after symmetrization) are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Undamped/

- LHO_ISI_BSC8_TF_C2C_Symmetrized_from_ST1_ACT_to_ST1_CPS_2011_07_27.fig

- LHO_ISI_BSC8_TF_C2C_Symmetrized_from_ST1_ACT_to_ST1_L4C_2011_07_27.fig

- LHO_ISI_BSC8_TF_C2C_Symmetrized_from_ST2_ACT_to_ST2_CPS_2011_07_27.fig

- LHO_ISI_BSC8_TF_C2C_Symmetrized_from_ST2_ACT_to_ST2_GS13_2011_07_27.fig

The data structure that contains the symmetrized transfer functions is located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Simulations/Undamped/

- LHO_ISI_BSC8_TF_C2C_Sym_10mHz_1000Hz_2011_07_27.mat

Figure 23 - Transfer functions in the Cartesian basis - ST1 ACT to ST1 CPS

Figure 24 - Transfer functions in the Cartesian basis - ST1 ACT to ST1 L4C

Figure 25 - Transfer functions in the Cartesian basis – ST2 ACT to ST2 CPS

Figure 26 - Transfer functions in the Cartesian basis – ST2 ACT to ST2 GS13

## Digitalization of the symmetrization filters

The digitized filters can be found in the SVN at

seismic/BSC-ISI/H2/ITMY/Filters/Digitized/

- LHO_ISI_BSC8_Symmetrization_Filters_28-Jul-2011.txt

## Design of the damping filters of stage 2

The damping filters are first designed on stage 2. The damping filters of stage 1 and stage 2 have the same poles and zeros. Only gains differ.

Here is the piece of code where to filters are defined:

L4C_GS13_Idealized_Model=zpk([0 0 0],-2*pi*[pair(1,45)],1);

Damping_Filters(ST2)=zpk(pole(L4C_GS13_Idealized_Model),-2*pi*[pair(0.1,55)],1);

Cut_Off_Filter([ST1 ST2])=zpk(-2*pi*[],-2*pi*[pair(30,65) pair(95,80)],1);

Damping_Filters(ST2)=Damping_Filters(ST2).*Cut_Off_Filter(ST2);

Only the gain is adjusted for every degree of freedom.

The data structure that contains the transfer functions of the damped ISI (by stage 2) can be found in the SVN:

seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Simulations/Damping/

- LHO_ISI_BSC8_TF_C2C_ST2_Damped_10mHz_1000Hz_2011_07_27.mat

The damping filters (continuous form) can be found in the SVN at:

seismic/BSC-ISI/H2/ITMY/Filters/Continuous/

- LHO_ISI_BSC8_Damping_Filters_2011_07_27.mat

The figures that show the effect of the damping on stage 1 when stage 2 is damped are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Misc/

- LHO_ISI_BSC8_MIMO_ST2_Damping_ONLY_TF_ST1_L4C_RZ.fig

- LHO_ISI_BSC8_MIMO_ST2_Damping_ONLY_TF_ST1_L4C_Z.fig

The figures that show how the damping filters are tuned are located in the SVN

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Damping/

- LHO_ISI_BSC8_ST2_Damping_TF_SISO_ST2_ACT_RX_to_ST2_GS13_RX_2011_07_27.fig

- LHO_ISI_BSC8_ST2_Damping_TF_SISO_ST2_ACT_RY_to_ST2_GS13_RY_2011_07_27.fig

- LHO_ISI_BSC8_ST2_Damping_TF_SISO_ST2_ACT_RZ_to_ST2_GS13_RZ_2011_07_27.fig

- LHO_ISI_BSC8_ST2_Damping_TF_SISO_ST2_ACT_X_to_ST2_GS13_X_2011_07_27.fig

- LHO_ISI_BSC8_ST2_Damping_TF_SISO_ST2_ACT_Y_to_ST2_GS13_Y_2011_07_27.fig

- LHO_ISI_BSC8_ST2_Damping_TF_SISO_ST2_ACT_Z_to_ST2_GS13_Z_2011_07_27.fig

Note 1: Once stage 2 is damped, stage 1 is also damped due to the important mass of stage 2 in comparison with stage 1.

Note 2: The motion of stage 1 around 2.5Hz is amplified on the damping loops on stage 2 are implemented.

Note 3: Once the ISI is damped on stage 2, the effect of the damping filters of stage 1 will be almost negligible.

The two following figures present the SISO responses of the ISI when the damping loops are engaged on stage 1.

Figure 27 - Damping Loops Tuning - TF SISO ST2 ACT Z to ST2 GS13 Z

Figure 28 - Damping Loops Tuning - TF SISO ST2 ACT RZ to ST2 GS13 RZ

The two following figures present some  transfer functions of the ISI on stage 1 when stage 2 is damped.

Figure 29 – Effect of stage 2 damping loops on stage 1 – MIMO transfer function - ST1 ACT Z to ST1 L4C Z

Figure 30 - Effect of stage 2 damping loops on stage 1 – MIMO transfer function - ST1 ACT RZ to ST1 L4C RZ

## Design of the damping filters of stage 1

Once stage 2 is damped, stage 1 is also. The damping filters of stage 2 have a very light effect. The gain of the damping filters are calculated using the reconstructed transfer functions (with stage 2 damped).

The figures that show how the damping filters are tuned are located in the SVN

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Damping/

- LHO_ISI_BSC8_ST1_ST2_Damping_TF_SISO_ST1_ACT_RX_to_ST1_GS13_RX_2011_07_27.fig

- LHO_ISI_BSC8_ST1_ST2_Damping_TF_SISO_ST1_ACT_RY_to_ST1_GS13_RY_2011_07_27.fig

- LHO_ISI_BSC8_ST1_ST2_Damping_TF_SISO_ST1_ACT_RZ_to_ST1_GS13_RZ_2011_07_27.fig

- LHO_ISI_BSC8_ST1_ST2_Damping_TF_SISO_ST1_ACT_X_to_ST1_GS13_X_2011_07_27.fig

- LHO_ISI_BSC8_ST1_ST2_Damping_TF_SISO_ST1_ACT_Y_to_ST1_GS13_Y_2011_07_27.fig

- LHO_ISI_BSC8_ST1_ST2_Damping_TF_SISO_ST1_ACT_Z_to_ST1_GS13_Z_2011_07_27.fig

Figure 31 - Damping Loops Tuning - TF SISO ST1 ACT Z to ST1 L4C Z

Figure 32 - Damping Loops Tuning - TF SISO ST1 ACT RZ to ST1 L4C RZ

The figure below presents the effect of the damping filters of stage 1 on stage 1 when stage 2 is not damped

Figure 33 - MIMO transfer function – ST1 ACT Z to ST1 L4C Z – Damping both stages

Figure 34 - MIMO transfer function – ST1 ACT RZ to ST1 L4C RZ – Damping both stages

The symmetrized and damped transfer functions (MIMO response) in the Cartesian basis are saved in a data structure located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Simulations/Damping/

- LHO_ISI_BSC8_TF_C2C_Damped_ST1_ST2_10mHz_1000Hz_2011_07_27.mat

## Digitalization of the damping filters

The digitized damping filters are saved in the SVN at:

seismic/BSC-ISI/H2/ITMY/Filters/Digitized/

- LHO_ISI_BSC8_Damping_Filters_28-Jul-2011.txt

## ASD with Damping loops closed

The figures of the damped ISI ASD are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Spectra/Damping/

- LHO_ISI_BSC8_Powerspectra_CT_ST1_L4C_Undamped_Damped_2011_07_29.fig

- LHO_ISI_BSC8_Powerspectra_CT_ST2_GS13_Undamped_Damped_2011_07_29.fig

- LHO_ISI_BSC8_Powerspectra_m_ST1_L4C_Undamped_Damped_2011_07_29.fig

- LHO_ISI_BSC8_Powerspectra_m_ST2_GS13_Undamped_Damped_2011_07_29.fig

The calibrated ASD data are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Spectra/Damping/

- LHO_ISI_BSC8_Calibrated_PSD_L4C_GS13_Undamped_Damped_2011_07_29.mat

Figure 35 - ASD ST1 L4C - ST1 & ST2 Damped

Figure 36 - ASD ST2 GS13 - ST1 & ST2 Damped

## Blending CPS and Geophones signals - Super Sensors

For this quick implementation of the isolation loops, basic complementary filters are used.

The complementary filters used for the blend are located in the SVN at:

seismic/BSC-ISI/Common/Complementary_Filters_BSC_ISI/aLIGO/

- aLIGO_Complementary_Filters_BSC_ISI.mat

A blend frequency of 750mHz was chosen for this test.

The figure that shows the fitting of the Cartesian to Cartesian transfer functions can be found in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Misc/

- LHO_ISI_BSC8_Fitting_TF_C2C.fig

The continuous form of the blend filters can be found in the SVN at:

seismic/BSC-ISI/H2/ITMY/Filters/Continuous/

- LHO_ISI_BSC8_Blend_Filters_2011_07_27.mat

Figure 37 - Automatic curve fitting - ST1 & ST2 CPS - ST1 L4C & ST2 GS13

The two figures below show the contribution of the CPS and the geophones in the super sensors responses. These figures can be found in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Super_Sensors/

- LHO_ISI_BSC8_Super_Sensor_SISO_ST1_Z_Blend_Freq_0.75Hz.fig

- LHO_ISI_BSC8_Super_Sensor_SISO_ST2_RZ_Blend_Freq_0.75Hz.fig

Figure 38 – Transfer function from ST1 ACT Z to ST1 Super sensors Z

Figure 39 – Transfer function from ST2 ACT RZ to ST2 Super sensor RZ

Note: The blend filters used during this test are very basic. They will need to be reworked.

## Digitalization of the Blend filters

The digitized blend filters are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Filters/Digitized/

- LHO_ISI_BSC8_Blend_Filters_04-Aug-2011.txt

## Isolation filters – Stage 1 – Z – RX – RY

The isolation filters are designed in 4 steps. Between each of these steps, the MIMO response of the ISI is calculated. Then the MIMO response is used to design the next batch of isolation filters. The design steps are the followings:

- Stage 1 – Z – RX – RY

- Stage 2 – Z – RX – RY

- Stage 1 – X – Y – RZ

- Stage 2 – X – Y – RZ

All isolation filters have been tuned respecting the following criteria:

- UUG=15Hz

- Phase margin>45 degree

- Gain margin > 20dB

This first set of isolation filters (continuous form) are stored in the SVN at:

seismic/BSC-ISI/H2/ITMY/Filters/Continuous/

- LHO_ISI_BSC8_Isolation_Filters_ST1_V_2011_07_27.mat

The figures that show the tuning of the stage 1 – Z – RX – RY isolation filters are located in the SVN:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Isolation/

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST1_ACT_Z_to_ST1_SS_Z_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST1_ACT_RX_to_ST1_SS_RX_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST1_ACT_RY_to_ST1_SS_RY_2011_07_27.fig

An example is given with the Super sensor on stage 1.

Figure 40 – Transfer function from ST1 ACT Z to Super sensor ST1 – Z - SISO - Isolation filter

## Isolation filters – Stage 2 – Z – RX - RY

These two first set of isolation filters (continuous form) are stored in the SVN at:

seismic/BSC-ISI/H2/ITMY/Filters/Continuous/

- LHO_ISI_BSC8_Isolation_Filters_ST1_ST2_V_2011_07_27.mat

The figures that show the tuning of the stage 2 – Z – RX – RY isolation filters are located in the SVN:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Isolation/

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST2_ACT_Z_to_ST1_SS_Z_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST2_ACT_RX_to_ST1_SS_RX_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST2_ACT_RY_to_ST1_SS_RY_2011_07_27.fig

Figure 41 - Transfer function from ST2 ACT Z to Super sensor ST2 –Z - SISO - Isolation filter

## Isolation filters – Stage 1 – X – Y - RZ

These three first set of isolation filters (continuous form) are stored in the SVN at:

seismic/BSC-ISI/H2/ITMY/Filters/Continuous/

- LHO_ISI_BSC8_Isolation_Filters_ST1_ST2_V_ST1_H_2011_07_27.mat

The figures that show the tuning of the stage 2 – X – Y – RZ isolation filters are located in the SVN:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Isolation/

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST1_ACT_X_to_ST1_SS_X_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST1_ACT_Y_to_ST1_SS_Y_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST1_ACT_RZ_to_ST1_SS_RZ_2011_07_27.fig

Figure 42 - Super sensor ST1 – RZ - SISO - Isolation filter

## Isolation filters – Stage 2 –X – Y - RZ

The isolation filters (continuous form) are stored in the SVN at:

seismic/BSC-ISI/H2/ITMY/Filters/Continuous/

- LHO_ISI_BSC8_Isolation_Filters_ST1_ST2_V_ST1_ST2_H_2011_07_27.mat

The figures that show the tuning of the stage 2 – X – Y – RZ isolation filters are located in the SVN:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Isolation/

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST2_ACT_X_to_ST1_SS_X_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST2_ACT_Y_to_ST1_SS_Y_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_SISO_ST2_ACT_RZ_to_ST1_SS_RZ_2011_07_27.fig

Figure 43 - Super sensor ST2 – RZ - SISO - Isolation filter

Once damping and isolation filters are implemented, the MIMO response is calculated.

The transfer functions (figures) can be found in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Isolation/

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST1_ACT_RX_to_ST1_L4C_RX_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST1_ACT_RY_to_ST1_L4C_RY_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST1_ACT_RZ_to_ST1_L4C_RZ_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST1_ACT_X_to_ST1_L4C_X_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST1_ACT_Y_to_ST1_L4C_Y_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST1_ACT_Z_to_ST1_L4C_Z_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST2_ACT_RX_to_ST2_GS13_RX_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST2_ACT_RY_to_ST2_GS13_RY_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST2_ACT_RZ_to_ST2_GS13_RZ_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST2_ACT_X_to_ST2_GS13_X_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST2_ACT_Y_to_ST2_GS13_Y_2011_07_27.fig

- LHO_ISI_BSC8_Isolation_Loops_TF_MIMO_ST2_ACT_Z_to_ST2_GS13_Z_2011_07_27.fig

The transfer functions hereinafter show how the BSC-ISI is affected by the implementation of the isolation filters.

Figure 44 - Transfer functions ST2 ACT Z to ST2 GS13 Z - Damping & Isolation on ST1 & ST2

Figure 45 - Transfer functions ST2 ACT RZ to ST2 GS13 Z - Damping & Isolation on ST1 & ST2

Figure 46 - Transfer functions - ST2 ACT X to ST2 GS13 X - Damping & Isolation on ST1 & ST2

## Open loop check

This step enables to check the importance of the order of implementation of the isolation filters. It computes open loops when isolation filters are engaged in three different configurations:

- When the damping  loops are engaged (first) and no other isolation filter is engaged

- Plant in the state used to design the isolation filters

- When all other filters are engaged

The figures “Open loop check” can be found in the SVN at

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Simulations/Open_Loop_Check/

- LHO_ISI_BSC8_Open_Loop_ST1_X_Y_RZ_2011_07_27.fig

- LHO_ISI_BSC8_Open_Loop_ST1_Z_RX_RY_2011_07_27.fig

- LHO_ISI_BSC8_Open_Loop_ST2_X_Y_RZ_2011_07_27.fig

- LHO_ISI_BSC8_Open_Loop_ST2_Z_RX_RY_2011_07_27.fig

The following figure shows an example with the stage 2 vertical isolation filters. These filters are designed in second (after stage 1 vertical isolation filters). It is clearly shown that the filters would be instable if there were engaged before isolation stage 1 vertically (blue curve).

Figure 47 - Open look check - ST2 Vertical Isolation filters

## Digitalization of the isolation filters

The digitized isolation filters can be found in the SVN.

seismic/BSC-ISI/H2/ITMY/Filters/Digitized/

- LHO_ISI_BSC8_Isolation_Filters_ST1_ST2_V_ST1_ST2_H_04-Aug-2011.txt

## ASD of the isolated ISI

Once the isolation filters were engaged, some ASD were measured. These measurements were realized in a clean room with the fans turned on.

The figures of the ASD can be found in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Spectra/Isolation/

- LHO_ISI_BSC8_Powerspectra_CT_ST1_L4C_Undamped_Isolated_FULL_2011_08_05.fig

- LHO_ISI_BSC8_Powerspectra_CT_ST2_GS13_Undamped_Isolated_FULL_2011_08_05.fig

The ASD are saved in a data structure and can be found in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Spectra/Isolation/

- LHO_ISI_BSC8_Calibrated_PSD_L4C_GS13_Damped_Isolated_ST1_ST2_2011_08_05.mat

Figure 48 – ASD ST1 L4C - Damped + Isolated (ST1 & ST2)

Figure 49 - ASD ST2 GS13 - Damped + Isolated (ST1 & ST2)

Note: The measurements are realized in a clean room with the fans turned on. It mainly explains why the performances are not good.

# Test with the QUAD (ITMY) and without FMY

The QUAD (dummy test mass) was hung under the ISI on August 17, 2011. Some measurements were done on August 31, 2011.

The segmented transfer functions are located in the SVN at

seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Measurements/Undamped/

- LHO_ISI_BSC8_Data_L2L_1Hz_10Hz_ST1_ST2_20110901-044859.mat

- LHO_ISI_BSC8_Data_L2L_10Hz_100Hz_ST1_ST2_20110901-021159.mat

- LHO_ISI_BSC8_Data_L2L_100Hz_500Hz_ST1_ST2_20110831-215702.mat

- LHO_ISI_BSC8_Data_L2L_500Hz_1000Hz_ST1_ST2_20110831-183234.mat

The transfer functions are concatenated in one data tructure located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Simulations/Undamped/

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2011_08_31.mat

The figures that show the transfer functions with the QUAD are located in the SVN.

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Measurements/Comparison/

- LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_H_to_ST1_CPS_H_20110722_vs_20110831.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_H_to_ST1_L4C_H_20110722_vs_20110831.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_V_to_ST1_CPS_V_20110722_vs_20110831.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_V_to_ST1_L4C_V_20110722_vs_20110831.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_H_to_ST2_CPS_H_20110722_vs_20110831.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_H_to_ST2_GS13_H_20110722_vs_20110831.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_V_to_ST2_CPS_V_20110722_vs_20110831.fig

- LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_V_to_ST2_GS13_V_20110722_vs_20110831.fig

Figure 50 - Transfer function comparison - ISI with/without QUAD - TF ST1 ACT H to ST1 CPS H

Figure 51 - Transfer function comparison - ISI with/without QUAD - TF ST1 ACT V to ST1 CPS V

Figure 52 - Transfer function comparison - ISI with/without QUAD - TF ST1 ACT H to ST1 L4C H

Figure 53 - Transfer function comparison - ISI with/without QUAD - TF ST1 ACT V to ST1 L4C V

Figure 54 - Transfer function comparison - ISI with/without QUAD - TF ST2 ACT H to ST2 CPS H

Figure 55 - Transfer function comparison - ISI with/without QUAD - TF ST2 ACT V to ST2 CPS V

Figure 56 – Transfer function comparison - ISI with/without QUAD - TF ST2 ACT H to ST2 CPS H

# Test with the QUAD (ITMY) and FMY before cartridge install

## Instrument changes

After the incident on Tuesday October 18, 2011 (aLOG 1587), testing activities were stopped on BSC8 for few weeks. While SUS were welding fibers on the QUAD, Seismic swapped some geophones.

GS13 and T240 installed for final testing are reported in the tables below.

| Unit | Location GS13 | Serial Number | POD | Comments |
| --- | --- | --- | --- | --- |
| BSC8 | H1 | Not noted at LLO | 33 | Class A – RGA OK |
| BSC8 | H2 | 801 | 31 | Class A – RGA OK |
| BSC8 | H3 | 783 | 16 | Class A – RGA OK |
| BSC8 | V1 | 740 | 79 | Class A – RGA OK |
| BSC8 | V2 | 739 | 94 | Class A – RGA OK |
| BSC8 | V3 | 677 | 26 | Class A – RGA OK |

Table 17 - GS13 installed for final testing

| Unit | Location T240 | Serial Number | POD | Comments |
| --- | --- | --- | --- | --- |
| BSC8 | Corner 1 | 112 | 32 | Surge protector installed |
| BSC8 | Corner 2 | 118 | 20 | Surge protector installed |
| BSC8 | Corner 3 | 101 | 19 | Surge protector installed |

Table 18 - L4C installed for final testing

Note: Rainer Weiss gave his approval on the GS13 pods after checking the RGA scans.

## Sensors check

Few ASD were measured to check the sensors and built reference amplitude spectral density. These spectra were measured in different conditions:

- Locked

- Unlocked

- ISI tilted

| State | GPS Time |
| --- | --- |
| ISI Unlocked | 1007321196 |
| ISI Locked | 1007334800 |

Table 19 - ASD - ISI Locked & unlocked

Note: The GS13 were installed once the ISI was in the LVEA. Tests were performed using the actuators to tilt the table. In that configuration, the tilt angle is not as large as it should be using masses on stage 2. However, it is sufficient to validate the functioning of the geophones.

|  | Time | Tilt | Offset drive |
| --- | --- | --- | --- |
| ISI Tilted 1 | 1007327576 | +RX | 40000 |
| ISI Tilted 2 | 1007327824 | -RX | -40000 |
| ISI Tilted 3 | 1007328175 | +RY | 32000 |
| ISI Tilted 4 | 1007328449 | -RY | -32000 |
| ISI Tilted 5 | 1007328907 |  | 25K RX-25KRY |
| ISI Tilted 6 | 1007329120 |  | 25K RX+25KRY |

Table 20 - ISI Tilted - Different angles

ASD of the unlocked, locked and titled ISI can be found in the SVN at:

/seismic/BSC-ISI/H2/ITMY/Data/Spectra/Undamped/

- LHO_ISI_BSC8_ASD_m_CPS_T240_L4C_GS13_2011_12_07_112621.mat (Unlocked)

- LHO_ISI_BSC8_ASD_m_CPS_T240_L4C_GS13_2011_12_07_151305.mat (Locked)

- LHO_ISI_BSC8_ASD_m_L4C_GS13_Stage_Tilted_2011_12_07.mat

Figures of these spectra can be found in the SVN at:

/seismic/BSC-ISI/H2/ITMY/Data/Figures/Spectra/Undamped/

- LHO_ISI_BSC8_ASD_m_CPS_T240_L4C_GS13_2011_12_07_112621.fig

- LHO_ISI_BSC8_ASD_m_CPS_T240_L4C_GS13_2011_12_07_151305.fig

- LHO_ISI_BSC8_Tilted_ASD_m_ST1_L4C_2011_12_07.fig

- LHO_ISI_BSC8_Tilted_ASD_m_ST2_GS13_2011_12_07.fig

Figure 57 - ASD ISI Unlocked

Figure 58 - ASD ISI Locked

Figure 59 - ASD of L4C when the stage 2 is tilted

Figure 60 - ASD GS13 when the stage 2 is tilted

Test result:	Passed:   X  	Failed:         .

## Mass budget

The mass budget is presented in the tables below.

| Stage 1 |  |  |  |
| --- | --- | --- | --- |
| Mass Type | Location | Weight (lb) | Weight (Kg) |
| D0902612 | C1-1 | 12 | 5.5 |
| D0902616-1,2,2,3 | C1-2 | 21.1 | 9.7 |
| D0902612 | C2-1 | 12 | 5.5 |
| D0902616--1,1,2 | C2-2 | 11.1 | 5.1 |
| D0902612 | C3-1 | 12 | 5.5 |
| D0902616--5 | C3-2 | 20.1 | 9.2 |
|  | Total | 88.3 | 40.4 |

Table 21 – Stage 1 – Mass budget

Total stage 1 payload: 40.4Kg

| Stage 2 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Mass type | Quantity | Total Weight (lb) | Total Weight (Kg) | Location |
| D1003136 | 12 | 600 | 272.2 | On Keel |
| D071200-06 | 1 | 27.2 | 12.3 | On Keel |
| D071200: 6,5,5;4,4,3,2,2,2;1,1,1,0 | 13 | 89.2 | 40.5 | Corner1 SideWall |
| D071200:6,3,3,1,0;5,1;1 | 8 | 55.7 | 25.3 | Corner2 SideWall |
| D071200:-;4,4;- | 2 | 9 | 4.1 | Corner3 SideWall |
|  | Total | 781.1 | 354.3 |  |

Table 22 - Stage 2 - Mass budget – Keel and side walls

| Stage 2 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Mass Type | Quantity | Weight | Total weight (Kg) | Total weight (lb) | Location |
| D1003163 | 6 | 10 | 60.0 | 132.3 | Optical Table |
| D1003164 | 6 | 5 | 30.0 | 66.1 | Optical Table |
| D1003123 | 6 | 3 | 18.0 | 39.7 | Optical Table |
| D1003166 | 6 | 1 | 6.0 | 13.2 | Optical Table |
| D1003167 | 6 | 0.5 | 3.0 | 6.6 | Optical Table |
| D1003168 | 12 | 0.25 | 3.0 | 6.6 | Optical Table |
|  |  | Total | 120.0 | 264.6 |  |

Table 23 - Stage 2 - Mass budget optic table

Stage 2 added masses (Keel + optical + Sidewall): 474.3Kg

SUS Payload:

44 lbs = 8 x 5.5 lbs for Vibration Absorbers
~60 lbs  = FMy Stays
~2 lbs = FMy Sheer Plate dampers
10 lbs  = Cabling
2 lbs = Cable Brackets
3 lbs  = Ring Heater + Brackets + Cables
266 lbs = ITMy QUAD Upper Structure
531 lbs = ITMy QUAD Lower Structure
44 lbs = QUAD Sleeve
2 lbs = QUAD sleeve wedges
394 lbs = FMy SUS
55 lbs = 1.26 lbs X Dog Clamps (33 long + 22 short)
TOTAL = 1413 lbs = 641 kg

Total stage 2 payload: 1115Kg

Nominal payload on stage 1: 109Kg – 240lb

Nominal payload on stage 2: 1185Kg – 2612lb

During phase I testing, the mass budget was:

- Stage 1: 35.6Kg

- Stage 2: 1082.4Kg

Note Seismic: There is a 20Kg difference in the total payload between Phase I and Phase II testing.

Note Suspension: For the duration of the SUS testing phase which occurred between Fri Dec 2nd and Wed Dec 7th, a small subset of payload was still needing to be added (not added prior due to variety of reasons).  The items are the following:

- 2 of the 4 QUAD Sleeve Wedges (the ones opposite each other perpendicular to the optic beam axis).  The last of these were installed on the QUAD on Friday Dec 9, just prior to installing the cartridge.
- The AOS FMy elliptical baffle - still waiting to install in-chamber

The total weight of the assembly was measured at 9260lbs using the load cell.

Acceptance criteria:

- Defined in E1000486

Test result:	Passed:   X  	Failed:         .

## Balancing check

The balancing of the ISI is evaluated by measuring the difference of the CPS offsets when the ISI is locked and unlocked.

The CPS offsets in the two configurations can be found in the SVN at:

/seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/

- LHO_ISI_BSC8_Locked_20111207.mat

- LHO_ISI_BSC8_Unlocked_20111207.mat

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | 260.6 | 36.6 | 417.3 | 62.7 | -156.7 | -0.19 |
| ST1 - H2 | -219.0 | 22.4 | 106.2 | 40.9 | -325.2 | -0.39 |
| ST1 - H3 | -100.5 | 16.4 | 569.4 | 73.0 | -669.9 | -0.80 |
| ST1 - V1 | -160.9 | 75.5 | 78.7 | 71.4 | -239.7 | -0.29 |
| ST1 - V2 | -60.9 | 138.5 | -312.7 | 93.0 | 251.8 | 0.30 |
| ST1 - V3 | -112.0 | 46.8 | 287.3 | 69.3 | -399.3 | -0.48 |
| ST2 - H1 | -138.1 | 25.3 | 591.0 | 98.3 | -729.1 | -0.22 |
| ST2 - H2 | 857.1 | 66.9 | 835.4 | 97.7 | 21.7 | 0.01 |
| ST2 - H3 | -345.3 | 48.4 | 836.2 | 59.0 | -1181.5 | -0.35 |
| ST2 - V1 | 923.6 | 35.6 | 1020.2 | 209.1 | -96.6 | -0.03 |
| ST2 - V2 | 706.3 | 55.7 | 468.9 | 244.2 | 237.4 | 0.07 |
| ST2 - V3 | -470.1 | 54.0 | -147.2 | 203.4 | -322.9 | -0.10 |

Table 24 – CPS Offsets – Locked vs Unlocked

Acceptance criteria:

- Defined in E1000486 – III.2

Test result:	Passed:   X  	Failed:         .

## Static test – Local Basis

The results of the static test in the local basis are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/

- LHO_ISI_BSC8_Offset_Local_Drive_20111206.mat

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 3836 | 1537 | 1535 | 19 | -8 | -9 |
|  | ST1 - H2 | 1519 | 3756 | 1509 | 0 | 15 | -4 |
|  | ST1 - H3 | 1517 | 1515 | 3756 | 2 | -5 | 13 |
|  | ST1 - V1 | 59 | -130 | 87 | 2892 | -496 | -468 |
|  | ST1 - V2 | 78 | 30 | -144 | -479 | 2909 | -500 |
|  | ST1 - V3 | -146 | 79 | 16 | -506 | -467 | 2918 |
|  |  | Sensors |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2146 | 318 | 323 | -27 | -17 | -17 |
|  | ST2 - H2 | 321 | 2147 | 318 | -19 | -13 | -27 |
|  | ST2 - H3 | 338 | 333 | 2237 | -20 | 24 | -7 |
|  | ST2 - V1 | 62 | 110 | -174 | 2609 | 370 | 4 |
|  | ST2 - V2 | -184 | 59 | 89 | 10 | 2606 | 303 |
|  | ST2 - V3 | 106 | -178 | 61 | 299 | 25 | 2585 |

Table 25 – Static Test – Local drive – Offsets in the local basis

Test result:	Passed:   X   	Failed:         .

## Static Test – Cartesian Basis

The results of the static test in the Cartesian basis are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/

- LHO_ISI_BSC8_Offset_Cartesian_Drive_20111206.mat

|  |  | ST1 - X | ST1 - Y | ST1 - Z | ST1 – RX | ST1 - RY | ST1 - RZ |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Actuators | ST1 - X | 1523 | 2 | -7 | -12 | -7 | 31 |
|  | ST1 - Y | -7 | 1506 | -25 | 11 | -9 | 3 |
|  | ST1 - Z | 27 | -1 | 656 | 5 | 5 | -7 |
|  | ST1 - RX | 5 | 305 | 0 | 2469 | -6 | -10 |
|  | ST1 - RY | -291 | 19 | -2 | -1 | 2472 | -9 |
|  | ST1 - RZ | 14 | 2 | -1 | 1 | 0 | 2884 |

|  |  | ST2 - X | ST2 - Y | ST2 - Z | ST2 – RX | ST2 - RY | ST2 - RZ |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Actuators | ST2 - X | 1229 | -19 | 57 | 25 | 5 | 15 |
|  | ST2 - Y | 0 | 1234 | -27 | 7 | 1 | -42 |
|  | ST2 - Z | 9 | 3 | 964 | 10 | 15 | -3 |
|  | ST2 - RX | 3 | -6 | -5 | 3755 | -1 | -11 |
|  | ST2 - RY | 21 | 2 | -15 | 11 | 3747 | -2 |
|  | ST2 - RZ | 12 | -33 | -27 | -5 | 16 | 2341 |

Figure 61 - Static Test – Cartesian Drive – Offsets in the Cartesian Basis

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - X | 1548 | -730 | -736 | 28 | -1 | -7 |
|  | ST1 - Y | -25 | 1298 | -1272 | 12 | 19 | 0 |
|  | ST1 - Z | -48 | -7 | -23 | 649 | 657 | 657 |
|  | ST1 - RX | -17 | 127 | -146 | -2367 | 2015 | 369 |
|  | ST1 - RY | -181 | 78 | 58 | -936 | -1580 | 2541 |
|  | ST1 - RZ | 2787 | 2758 | 2764 | 5 | -4 | -5 |
|  |  | Sensors |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - X | 617 | -1211 | 630 | -36 | -1 | 23 |
|  | ST2 - Y | 1056 | 1 | -1083 | -15 | 8 | 13 |
|  | ST2 - Z | 6 | 8 | 4 | 948 | 979 | 972 |
|  | ST2 - RX | -247 | -3 | 238 | -2198 | 2227 | -38 |
|  | ST2 - RY | 145 | -268 | 152 | -1341 | -1211 | 2546 |
|  | ST2 - RZ | 1612 | 1613 | 1661 | -18 | 18 | 21 |

Table 26 - Static Test – Cartesian Drive – Offsets in the Local Basis

Test result:	Passed:   X   	Failed:         .

## Pressure sensors check

The pressure readout are located in the SVN are:

- seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/LHO_ISI_BSC8_Pressure_Sensors_Check_20111206.mat

|  | Pressure (counts) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-D | 37 | 124 | 231 |
| ST1-L4C-P | 24642 | 24647 | 24589 |
| ST1-GS13-D | 625 | 1165 | 4065.0 |
| ST1-GS13-P | 24425 | 24324 | 24266 |
| ST1-T240-P | 14073 | 13642 | 13829 |

Figure 62 - Pressure sensor readout

Acceptance criteria:

  - The absolute pressure on the L4Cs and the GS13s must be 24700 +/- 1000 counts (100+/- 3 KPA)

  - The differential pressure on the L4Cs and the GS13s must be <5200 counts (3 KPa)

  - The absolute pressure on the T240 must be 14300 +/-600 counts (100 +/- 3 KPA)

Test result:	Passed:   X  	Failed:         .

## Range of motion test

The results of the range of motion test are located in:

- seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/LHO_ISI_BSC8_Range_Of_Motion_20111206.mat

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -15865 | 357 | 16290 | 32155 | 38.3 |
| ST1 - H2 | -16123 | 21 | 15849 | 31972 | 38.1 |
| ST1 - H3 | -15596 | 562 | 15304 | 30900 | 36.8 |
| ST1 - V1 | -12525 | -30 | 12483 | 25007 | 29.8 |
| ST1 - V2 | -12793 | -354 | 12069 | 24862 | 29.6 |
| ST1 - V3 | -12064 | 321 | 12668 | 24732 | 29.4 |
| ST2 - H1 | -8794 | 539 | 9846 | 18640 | 5.5 |
| ST2 - H2 | -8512 | 814 | 10166 | 18678 | 5.6 |
| ST2 - H3 | -8431 | 800 | 10025 | 18456 | 5.5 |
| ST2 - V1 | -10135 | 983 | 12128 | 22263 | 6.6 |
| ST2 - V2 | -10512 | 512 | 11492 | 22004 | 6.5 |
| ST2 - V3 | -11252 | -95 | 11035 | 22287 | 6.6 |

Figure 63 - Range of motion

Test result:	Passed:   X   	Failed:         .

## Linearity-hysteresis test

The linearity-hysteresis test enables to eventually detect cable contacts. It is started before measuring long transfer functions.

The data measured in this test are located in the SVN at:

/seismic/BSC-ISI/H2/ITMY/Data/Linearity_Test/

- LHO_ISI_BSC8_Linearity_test_20111206.mat

The figure of the linearity test is located in the SVN at:

/seismic/BSC-ISI/H2/ITMY/Data/Figures/Linearity_Test/

- LHO_ISI_BSC8_Linearity_test_20111206.fig

Figure 64 - Linearity Test

|  |  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- | --- |
| Stage 1 | ST1 - H1 | 0.541 | 389 | 0.5402 | 0.20 |
|  | ST1 - H2 |  |  |  |  |
|  | ST1 - H3 | 0.539 | 595 |  | -0.20 |
|  | ST1 - V1 | 0.417 | 86 | 0.4146 | 0.47 |
|  | ST1 - V2 | 0.414 | -270 |  | -0.10 |
|  | ST1 - V3 | 0.413 | 383 |  | -0.37 |
| Stage 2 | ST2 - H1 | 0.312 | 589 | 0.3105 | 0.33 |
|  | ST2 - H2 | 0.312 | 833 |  | 0.56 |
|  | ST2 - H3 | 0.308 | 803 |  | -0.89 |
|  | ST2 - V1 | 0.373 | 1077 | 0.3711 | 0.38 |
|  | ST2 - V2 | 0.368 | 510 |  | -0.75 |
|  | ST2 - V3 | 0.373 | -100 |  | 0.37 |

Table 27 - Linearity Test - Before the move

Note: The interpolation function that is used to smooth and reduce the number of points used for display failed for the ST1-ACT-H2 to ST1-CPS-H2 measurement. That’s why no data is provided for ST1-H2.

Test result:	Passed:   X   	Failed:         .

## Transfer function measurements

The segmented transfer functions are located in the SVN at

seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Measurements/Undamped/:

- LHO_ISI_BSC8_Data_L2L_10mHz_100mHz_ST1_ST2_20111206-052309.mat

- LHO_ISI_BSC8_Data_L2L_100mHz_700mHz_ST1_ST2_20111205-235951.mat

- LHO_ISI_BSC8_Data_L2L_700mHz_10Hz_ST1_ST2_20111206-191136.mat

- LHO_ISI_BSC8_Data_L2L_10Hz_100Hz_ST1_ST2_20111205-212245.mat

- LHO_ISI_BSC8_Data_L2L_100Hz_500Hz_ST1_ST2_20111205-200228.mat

- LHO_ISI_BSC8_Data_L2L_500Hz_1000Hz_ST1_ST2_20111205-185743.mat

The transfer functions are concatenated in one data structure located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Simulations/Undamped/:

- LHO_ISI_BSC8_TF_L2L_Sym_10mHz_1000Hz_2011_12_06.mat

The figures that show the transfer functions with ITMY and FMY are located in the SVN at:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LHO_ISI_BSC8_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2011_12_06.fig

- LHO_ISI_BSC8_TF_L2L_Raw_from_ST1_ACT_to_ST1_T240_2011_12_06.fig

- LHO_ISI_BSC8_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2011_12_06.fig

- LHO_ISI_BSC8_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2011_12_06.fig

- LHO_ISI_BSC8_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2011_12_06.fig

Note: The transfer functions were measured the actuator symmetrization engaged.

Figure 65 – Transfer functions – With Damped QUAD and FMY – ST1 ACT to ST1 CPS

Figure 66 - Transfer functions – With Damped QUAD and FMY – ST1 ACT to ST1 L4C

Figure 67 - Transfer functions – With Damped QUAD and FMY – ST1 ACT to ST1 T240

Figure 68 - Transfer functions – With Damped QUAD and FMY – ST2 ACT to ST2 CPS

Figure 69 - Transfer functions – With Damped QUAD and FMY – ST2 ACT to ST2 GS13

Test result:	Passed:   X   	Failed:         .

## Damping loops

The damping filters were designed using BSC-ISI transfer functions without FMY and ITMY. The goal of this test was to show the robustness of the damping filters (Payload change).

| State | GPS Time |
| --- | --- |
| ISI with the Damping Loops | 1007316891 |
| ISI Unlocked | 1007321196 |

Table 28 - GPS Time - Damping loops ON and OFF

The ASD can be found in the SVN at:

/seismic/BSC-ISI/H2/ITMY/Data/Spectra/Damping/

- LHO_ISI_BSC8_ASD_m_L4C_GS13_Undamped_vs_Damping_2011_12_07_101436.mat

The figure of the ASD when the ISI is damped and undamped can be found in the SVN at:

/seismic/BSC-ISI/H2/ITMY/Data/Figures/Spectra/Damping/

- LHO_ISI_BSC8_ASD_m_ST1_L4C_Undamped_vs_Damping_2011_12_07_101436.fig

- LHO_ISI_BSC8_ASD_m_ST2_GS13_Undamped_vs_Damping_2011_12_07_101436.fig

Figure 70 - Damping Loops stage 1

Figure 71 - Damping Stage 2

Acceptance criteria:

- Damping loops are stable

Test result:	Passed:   X  	Failed:         .

Cartridge install

After two day of testing, the ISI was validated on Wednesday 7 december 2011. The cartridge install occurred on Friday 09 December 2011 (LHO Alog 1887).

# Tests after Cartridge install

## Transfer functions after cartridge install

Reference transfer functions L2L without baffle – No TMD – No vibration absorbers – No Viton under the keel masses - ITMY and FMY damped:

- LHO_ISI_BSC8_Data_L2L_10Hz_100Hz_ST1_ST2_20120106-214333.mat

- LHO_ISI_BSC8_Data_L2L_100Hz_500Hz_ST1_ST2_20120106-190824.mat

- LHO_ISI_BSC8_Data_L2L_500Hz_1000Hz_ST1_ST2_20120106-170349.mat

L2L concatenated

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2012_01_06.mat

Reconstructed C2C with Step_3_TF_C2C_from_L2L_with_Resymmetrization and loading LHO_ISI_BSC8_Symmetrization_Filters_2011_07_27.mat

- LHO_ISI_BSC8_TF_C2C_Reconstructed_Sym_10mHz_1000Hz_2012_01_06.mat

Data before the Cartridge install used for comparison

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2012_12_06.mat

Comparison Transfer functions Teststand vs Cartridge

Comparison_TF_L2L_LHO_ISI_BSC8.m

Comparison_TF_C2C_LHO_ISI_BSC8.m

In the local basis:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Comparison/

LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_H_to_ST1_CPS_H_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_V_to_ST1_CPS_V_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_H_to_ST1_L4C_H_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_V_to_ST1_L4C_V_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_H_to_ST2_CPS_H_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_V_to_ST2_CPS_V_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_H_to_ST2_GS13_H_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_V_to_ST2_GS13_V_20111206_vs_20120106.fig

In the Cartesian basis:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Comparison/

LHO_ISI_BSC8_Comparison_TF_C2C_ST1_ACT_H_to_ST1_CPS_H_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_C2C_ST1_ACT_V_to_ST1_CPS_V_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_C2C_ST1_ACT_H_to_ST1_L4C_H_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_C2C_ST1_ACT_V_to_ST1_L4C_V_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_C2C_ST2_ACT_H_to_ST2_CPS_H_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_C2C_ST2_ACT_V_to_ST2_CPS_V_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_C2C_ST2_ACT_H_to_ST2_GS13_H_20111206_vs_20120106.fig

LHO_ISI_BSC8_Comparison_TF_C2C_ST2_ACT_V_to_ST2_GS13_V_20111206_vs_20120106.fig

Note 1:

Transfer functions are measured between 10Hz and 1Kz. The ISI was not available for low frequency measurements.

Note 2:

The transfer functions in the Cartesian basis are computed using the local to local transfer functions. Filters of the input bank are introduced during reconstruction

Note 3:

After comparing with the measurements before the cartridge install, we can notice:

- The resonances of the teststand are not visible anymore (First teststand resonance are at 19.6Hz, 21.4Hz, 29.6Hz). HEPI resonances are visible on ST1 CPS transfer functions.

- First resonance of stage 1 is 217Hz (it will be damped by the vibration absorbers on the door of stage 1) on ST1 ACT to ST1 L4C transfer functions

- Second resonance observed on stage 1 at ~250 Hz. The HAM-ISI TMDs retuned for BSC-ISI (E1100963-v1) will be subsequently installed.

- New resonances in the stage 2 transfer functions (53Hz). It is more visible in the Cartesian to Cartesian transfer functions.

Figure 72 – Comparison TF C2C reconstructed - ST1 ACT H to ST1 CPS H

Figure 73 - Comparison TF C2C reconstructed - ST1 ACT H to ST1 L4C H

Figure 74 - Comparison - TF C2C reconstructed – ST2 ACT V to ST2 GS13 V

## Addition of Vibration absorbers – Viton pads – TMDs

Transfer functions with the final configuration for the TMDs, vibration absorbers and Viton pads are presented at the end of the subsection.

### Vibration absorbers (0.25”x0.25”xThickness”)

Vibration Absorbers (217Hz resonance)
We performed tests with different thicknesses of Viton Pads with and without the lids. These configurations (S=0.25”x0.25”) are:
- Thickness: 0.0625”, 0.125”, 0.25” without lid
- Thickness: 0.0625” with lids and pads of 0.0625” and 0.125”

The best configuration:
- Bottom pads with a thickness of 0.0625”
- The top pads thickness doesn’t seem to affect the performance. Only 0.25”x0.25”x0.0625” pads are installed in the stage 1 vibration absorbers (for simplicity reasons)

Performance on the 217Hz resonance (the only resonance affected by the vibration absorbers):
- X, Y directions reduction by a factor of 7.8 on ST1 L4C
- RX, RY directions reduction by a factor of 6.5 on ST1 L4C
- the resonance is not visible on the GS13s

The figure below presents the transfer function in the Cartesian basis around the 217Hz with different pads thicknesses.

Figure 75 - Stage 1 - Vibration absorbers - Thickness test

Note: 6 vibrations absorbers were installed on stage 1 doors. After adding the lids, the stage 1 position is a little bit low (-2000 counts (2.5mils) on vertical CPS) and there is no contingency for adjustment masses. Since there is lots of room in the stage 1 lockers, the six vibration absorbers were kept to benefit from a maximal damping effect.

Dimensions:

- 4 pads at the bottom (0.25”x0.25”x0.0625”)

- 4 pads at the top (0.25”x0.25”x0.0625”)

- 2 precision washers with a 0.02” thickness

### TMD (blade 0-1)

HAM-ISI TMDs were retuned at 252Hz for BSC-ISI stage 0-1 blades. Details are given in E1100963-v1. The TMDs (base) are installed at 11 inches from the post.

Tests realized:

- Try different TMDs on the same blade

- Try different location using the same TMD

- Measurement in the morning and the afternoon

Notes and Results:

- The TMDs are tuned at 252Hz +/-1Hz

- If the TMD is too close from the blade post or the tip of the blade, the TMDs seem to be less effective. When the base of the TMD is located at 11 inches from the post, the TMDs seem to be relatively efficient. Further tests will be on BSC6.

- During testing, measurements show performance variations for a same TMD on the same blade during the day. Temperature variations are probably the source of the performance changes. The temperature may slightly change the stiffness of the TMDs, consequently changing the resonance frequency of the TMDs. In some case, the TMD and the blade may or may not be tuned.

- Tuned mass dampers effect is mainly visible on the horizontal degrees of freedom. We can see a light effect on Z.

TMDs Serial numbers in the final configuration:

Corner 1: S/N 15

Corner 2: S/N 19

Corner 3: S/N 14

Figure 76 - TMD on the blade 0-1

### Viton pads under the stage 2 stack mass

Viton pads with the dimensions (1”x1”x0.25”) were added under the stack of masses on stage 2 (8 pads per stack).

Results:

Effective on resonances at 162Hz, 180Hz and 227Hz.

### The transfer functions comparisons

All the transfer functions presented below are measured in the local basis and recombine in the Cartesian basis.

Transfer functions L2L without baffle – No TMD – No vibration absorbers – No Viton under the keel masses - ITMY and FMY damped:

- LHO_ISI_BSC8_Data_L2L_10mHz_100mHz_ST1_ST2_20111206-052309.mat

- LHO_ISI_BSC8_Data_L2L_100mHz_700mHz_ST1_ST2_20111205-235951.mat

- LHO_ISI_BSC8_Data_L2L_700mHz_10Hz_ST1_ST2_20111206-191136.mat

- LHO_ISI_BSC8_Data_L2L_10Hz_100Hz_ST1_ST2_20111205-212245.mat

- LHO_ISI_BSC8_Data_L2L_100Hz_500Hz_ST1_ST2_20111205-200228.mat

- LHO_ISI_BSC8_Data_L2L_500Hz_1000Hz_ST1_ST2_20111205-185743.mat

L2L concatenated

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2012_01_06.mat

Reconstructed C2C with Step_3_TF_C2C_from_L2L_with_Resymmetrization and loading LHO_ISI_BSC8_Symmetrization_Filters_2011_07_27.mat

- LHO_ISI_BSC8_TF_C2C_Reconstructed_Sym_10mHz_1000Hz_2012_01_06.mat

The set of that enables to evaluate the performance of the dampers is realized using the following measurements:

- LHO_ISI_BSC8_Data_L2L_100Hz_200Hz_ST1_ST2_20120130-225348.mat

- LHO_ISI_BSC8_Data_L2L_200Hz_300Hz_ST1_ST2_20120130-213347.mat

- LHO_ISI_BSC8_Data_L2L_300Hz_500Hz_ST1_ST2_20120130-201344.mat

- LHO_ISI_BSC8_Data_L2L_500Hz_750Hz_ST1_ST2_20120130-190908.mat

- LHO_ISI_BSC8_Data_L2L_750Hz_1000Hz_ST1_ST2_20120130-180434.mat

L2L concatenated

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2012_01_30.mat

Reconstructed C2C with Step_3_TF_C2C_from_L2L_with_Resymmetrization and loading LHO_ISI_BSC8_Symmetrization_Filters_2011_07_27.mat

- LHO_ISI_BSC8_TF_C2C_Reconstructed_Sym_10mHz_1000Hz_2012_01_30.mat

Stage 1 transfer functions comparison – vibration absorbers and TMDs:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Comparisons/C2C/

- LHO_ISI_BSC8_Comparison_TF_C2C_ST1_ACT_V_to_ST1_L4C_V_20120106_vs_20120130.fig

- LHO_ISI_BSC8_Comparison_TF_C2C_ST1_ACT_H_to_ST1_L4C_H_20120106_vs_20120130.fig

Stage 2 transfer functions comparison – Stage 2 pads viton pads:

seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Comparisons/C2C/

- LHO_ISI_BSC8_Comparison_TF_C2C_ST2_ACT_H_to_ST2_GS13_H_20120106_vs_20120130.fig

- LHO_ISI_BSC8_Comparison_TF_C2C_ST2_ACT_H_to_ST2_GS13_H_20120106_vs_20120130.fig

Stage 1 – Transfer functions

Figure 77 - TF C2C ST1 L4C H

Figure 78 - TF C2C ST1 L4C H

Figure 79 - TF C2C ST1 L4C V

Figure 80 - TF C2C ST1 L4C V

Stage 2 – Transfer functions

Figure 81 - TF C2C – ST2 GS13 Horizontals

Figure 82 - TF C2C – ST2 GS13 Verticals

### SUS resonances – Payload History

In this section, transfer functions of the ISI are presented for different types of payload, different stage 0 support (crossbeams vs teststand) and different dampers (vibration absorbers and TMDs).

The configurations are the following:

- LVEA teststand - Undamped - In the air - 2011 07 27

- LVEA teststand - Undamped - In the air - In The LVEA after the ST2-GS13 V3 change - QUAD installed (not damped) - Fans ON - 2011 08 31

- LVEA teststand - Undamped - In the air - FM % QUAD installed (damped) - Fans ON - 2011 12 06

- In chamber - Undamped - In the air - FM & QUAD installed (damped) - Fans ON - 2012 01 06

- H2-ITMY - ISI-BSC8 - Undamped - After cartridge install - In the air - FM & QUAD installed (damped) - Fans ON - Vibration absorbers stage 1 (0.25"x0.25"x0.25") - Vitton pads under the stage 2 top masses (8 times 1"x1"x0.25") - 2012 01 24

- H2-ITMY - ISI-BSC8 - Undamped - After cartridge install - In the air - FM & QUAD installed (damped) - Fans ON - Vibration absorbers stage 1 (0.25"x0.25"x0.0625" + Lids 0.125") - Vitton pads under the stage 2 top masses (8 times 1"x1"x0.25") - TMDs installed at 11" from base - 2012 01 25

- H2-ITMY - ISI-BSC8 - Undamped - After cartridge install - In the air - FM & QUAD installed (damped) - Fans ON - Vibration absorbers stage 1 (0.25"x0.25"x0.0625" + Lids 0.125") - Vitton pads under the stage 2 top masses (8 times 1"x1"x0.25") - TMDs installed at 11" from base - 2012 01 26

Concatenated data in the local basis can be found in the SVN at:

/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Simulations/Undamped/

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2011_07_27.mat

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2011_08_31.mat

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2011_12_06.mat

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2012_01_06.mat

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2012_01_25.mat

- LHO_ISI_BSC8_TF_L2L_Raw_10mHz_1000Hz_2012_01_26.mat

Concatenated data in the Cartesian basis can be found in the SVN at:

BSC-ISI/H2/ITMY/Data/Transfer_Functions/Simulations/Undamped/

- LHO_ISI_BSC8_TF_C2C_Reconstructed_Sym_10mHz_1000Hz_2011_07_27.mat

- LHO_ISI_BSC8_TF_C2C_Reconstructed_Sym_10mHz_1000Hz_2011_08_31.mat

- LHO_ISI_BSC8_TF_C2C_Reconstructed_Sym_10mHz_1000Hz_2011_12_06.mat

- LHO_ISI_BSC8_TF_C2C_Reconstructed_Sym_10mHz_1000Hz_2012_01_06.mat

- LHO_ISI_BSC8_TF_C2C_Reconstructed_Sym_10mHz_1000Hz_2012_01_25.mat

- LHO_ISI_BSC8_TF_C2C_Reconstructed_Sym_10mHz_1000Hz_2012_01_26.mat

ST1 ACT TO ST1 CPS

The comparison of the transfer functions are located in the SVN at:

/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Comparisons/Evolution/

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_X_TO_ST1_CPS_X.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_Y_TO_ST1_CPS_Y.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_Z_TO_ST1_CPS_Z.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_RX_TO_ST1_CPS_RX.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_RY_TO_ST1_CPS_RY.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_RZ_TO_ST1_CPS_RZ.fig

Evolution between transfer functions:

- Teststand (long legs) resonance frequencies at 19.6Hz, 21.4Hz and 29.6Hz. (2011_07_27 – 2011_08_31 – 2011_12_06)

- HEPI resonances at 9Hz, 13Hz and 35 Hz (cross beams) (2012_01_06 – 2012_01_24 – 2012_01_25 – 2012_01_26 – 2012_01_30)

- Change of the first resonance (ISI suspension modes) due to the change of inertia created by the different types of payload

- Effect of the vibration absorbers (217Hz) and the TMDs (252Hz) not visible on the CPS in the Cartesian basis

- 2012 01 06 – HEPI Unlocked – RY - X

- 2012 01 26 – HEPI locked

- QUAD resonances at 0.46Hz and 0.54Hz (SUS Model) – slightly visible in the main couplings. More visible on cross couplings (ST1 to ST2 and ST2 to ST1). See tables below.

| SUS | L2L | P2P | R2R | T2T | V2V | Y2Y |
| --- | --- | --- | --- | --- | --- | --- |
| ITMY | 0.4398 | 0.4398 | 0.4632 | 0.4629 | 0.5507 | 0.5996 |
| FMY | 0.4193 | 0.4193 | 0.4236 | 0.4236 | 1.084 | 0.4934 |

Table 29 - First modes of the QUAD (ITMY) and the folding mirror (FMY)

Figure 83 - ST1 ACT RY to ST1 CPS RY - Comparison payload

Figure 84 - ST1 ACT X to ST1 CPS X – Comparison payload

ST1 ACT TO ST1 L4C

/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Comparisons/Evolution/

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_X_TO_ST1_L4C_X.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_Y_TO_ST1_L4C_Y.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_Z_TO_ST1_L4C_Z.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_RX_TO_ST1_L4C_RX.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_RY_TO_ST1_L4C_RY.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST1_ACT_RZ_TO_ST1_L4C_RZ.fig

Evolution between transfer functions:

- Teststand resonances slightly visible at 19Hz, 21Hz and 29.6Hz (2011 07 27 – 2011 08 31 – 2011 12 06)

- Change of the first resonance (ISI suspension modes) due to the change of inertia created by the different payload

- Reduction of the resonance at 217Hz after introduction of the 6 vibration absorbers on the doors of stage 1

- Modification of the stage 0-1 blade frequency (252Hz) after introducing  the TMDs (X, Y, RX, RY)

Figure 85 - ST1 ACT Y to ST1 L4C Y

ST2 CPS

/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Comparisons/Evolution/

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_X_TO_ST2_CPS_X.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_Y_TO_ST2_CPS_Y.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_Z_TO_ST2_CPS_Z.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_RX_TO_ST2_CPS_RX.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_RY_TO_ST2_CPS_RY.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_RZ_TO_ST2_CPS_RZ.fig

Evolution between transfer functions:

- Light DC change after 2011 08 31 – Transfer functions measured with the actuator symmetrization

- Change of the first resonance (ISI suspension modes) due to the change of inertia created by the different payload

- Effect of the Vibration absorbers (217Hz)

- Barely see the QUAD resonances at 0.46Hz and 0.54Hz in the main couplings

Figure 86 - ST2 ACT RX to ST2 CPS RX

ST2 GS13

/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Comparisons/Evolution/

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_X_TO_ST2_GS13_X.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_Y_TO_ST2_GS13_Y.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_Z_TO_ST2_GS13_Z.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_RX_TO_ST2_GS13_RX.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_RY_TO_ST2_GS13_RY.fig

- LHO_ISI_BSC8_Evolution_TF_C2C_ST2_ACT_RZ_TO_ST2_GS13_RZ.fig

Evolution between transfer functions:

- Bump at 0.7Hz (end of a measurement section)

- Difficulties to measure high frequency transfer functions

- Mode at 51Hz (ST2 GS13 – No SUS) – Disappears after installation of the SUS items – After 2011 07 27.

- In general, Q at high frequencies are reduced after installing SUS items.

- Smooth resonance at 67.Hz (RX) after installing the QUAD

- New resonances around ~53 Hz after cartridge installation. Not visible in 2011 12 06.

- New resonance at 77Hz after installing  the folding mirror (RX)

- Resonance at 67Hz which is damped after installation of the SUS items. (ST2 GS13 RX)

- Resonances around 0.4Hz slightly visible on the main couplings and more visible on the cross couplings.

Figure 87  - ST2 ACT RX to ST GS13 RX

Figure 88 - ST2 ACT RX to ST GS13 RX - Zoom

Figure 89 - ST2 ACT RY to ST2 GS13 RY

Figure 90 - ST2 ACT RY to ST2 GS13 RY - Zoom

Figure 91 - ST2 ACT RZ to ST2 GS13 RZ

Figure 92 - ST2 ACT RZ to ST2 GS13 RZ - Zoom

Figure 93 - ST2 ACT X to ST2 GS13 X

Figure 94 - ST2 ACT RY to ST2 GS13 X - Zoom

Figure 95 - ST2 ACT Y to ST2 GS13 Y

Figure 96 -  ST2 ACT Y to ST2 GS13 Y - Zoom

Figure 97  - ST2 ACT Z to ST2 GS13 Z

Figure 98 -  ST2 ACT Z to ST2 GS13 Z - Zoom

### Suspension modes

SUS transfer functions are reported in the seismic SVN at:

seismic/BSC-ISI/H2/ITMY/Misc/SUS_Model/

- FMY

- ITMY

The suspension modes are reported in the table below:

| ITMY from ground |  |  |  |
| --- | --- | --- | --- |
| L2L | P2P | R2R | T2T | V2V | Y2Y |
| 0.4398 | 0.4398 | 0.4632 | 0.4629 | 0.5507 | 0.5996 |
| 0.5625 | 0.564 | 0.8745 | 0.8745 | 2.215 | 1356 |
| 0.9997 | 0.9971 | 1.048 | 1.048 | 2.812 | 2.396 |
| 2.008 | 1.313 | 2.123 | 2.123 | 3.578 | 3.051 |
| 3.4411 | 1.605 | 2.633 | 2.633 |  |  |
|  | 2.007 | 3.317 | 3.317 |  |  |
|  | 2.812 | 5.074 | 5.069 |  |  |
|  | 3.317 |  |  |  |  |
|  | 3.417 |  |  |  |  |

Table 30 - ITMY SUS pendulum modes

| FMY from ground |  |  |  |
| --- | --- | --- | --- |
| L2L | P2P | R2R | T2T | V2V | Y2Y |
| 0.4193 | 0.4193 | 0.4236 | 0.4236 | 1.084 | 0.4934 |
| 0.4684 | 0.4688 | 1.05 | 1.05 | 3.759 | 1.396 |
| 1.051 | 1.051 | 1.553 | 1.553 | 17.52 | 2.251 |
| 1.076 | 1.077 | 2.18 | 2.819 |  |  |
| 1.698 | 1.392 | 3.206 | 3.206 |  |  |
|  |  | 25.97 |  |  |  |

Table 31 - FMY SUS pendulum modes

# Last tests before closing the chamber

## Spectra

Spectra are measured in the locked and the unlocked configurations.

ISI Locked

Data:

seismic/BSC-ISI/H2/ITMY/Data/Spectra/Undamped/

LHO_ISI_BSC8_ASD_m_LOC_CPS_T240_L4C_GS13_2012_01_31_164413.mat

Figures :

seismic/BSC-ISI/H2/ITMY/Data/Figures/Spectra/Undamped/

LHO_ISI_BSC8_ASD_m_LOC_CPS_T240_L4C_GS13_2012_01_31_164413.fig

Figure 99 - Spectra - Locked position - Local basis

ISI unlocked:

Data:

seismic/BSC-ISI/H2/ITMY/Data/Spectra/Undamped/

- LHO_ISI_BSC8_ASD_m_LOC_CPS_T240_L4C_GS13_2012_02_01_102955.mat

Figures :

seismic/BSC-ISI/H2/ITMY/Data/Figures/Spectra/Undamped/

- LHO_ISI_BSC8_ASD_m_LOC_CPS_T240_L4C_GS13_2012_02_01_102955.fig

Figure 100 – Spectra - Unlocked position - Local basis

## Geophones pressure sensors

The pressure sensors data are measured in two different ways:

- Using IOP input channels

- Using the model channels after calibration

/seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/

- LHO_ISI_BSC8_Pressure_Sensors_Check_Raw_20120131.mat

|  | Pressure (counts) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-D | 81 | 260.0 | 231.0 |
| ST1-L4C-P | 24519 | 24502.0 | 24471.0 |
| ST1-GS13-D | 706 | 968.0 | 4067.0 |
| ST1-GS13-P | 24388 | 24313 | 24259 |
| ST1-T240-P | 14044 | 13622 | 13802 |

Table 32 – Pressure Raw values (Count)

/seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/

- LHO_ISI_BSC8_Pressure_Sensors_Check_Calibrated_20120131.mat

|  | Pressure (KPA) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-D | 100.8 | 100.7 | 100.7 |
| ST1-L4C-P | 0.0 | 0.2 | 0.1 |
| ST1-GS13-D | 100.4 | 100.2 | 100.1 |
| ST1-GS13-P | 0.4 | 0.6 | 2.3 |
| ST1-T240-P | 101.7 | 99.6 | 100.5 |

Table 33 - Calibrate pressure (KPa)

## Balancing – CPS Offset locked vs unlocked

CPS offsets are measured before and after locking the platform.

Results are posted in the SVN at:

/seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/

- LHO_ISI_Locked_2_2012_01_31.mat

- LHO_ISI_Unlocked_3_2012_02_01.mat

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | mil | Offset (Mean) | Std deviation | mil | Offset (Mean) | mil |
| ST1 - H1 | 1055.6 | 10.1 | 1.3 | 1412.7 | 30.8 | 1.7 | -357.2 | -0.43 |
| ST1 - H2 | -1813.4 | 11.6 | -2.2 | -1123.8 | 16.6 | -1.3 | -689.6 | -0.82 |
| ST1 - H3 | -1110.7 | 6.7 | -1.3 | 65.9 | 20.3 | 0.1 | -1176.6 | -1.40 |
| ST1 - V1 | -1311.7 | 8.1 | -1.6 | -2403.2 | 16.7 | -2.9 | 1091.5 | 1.30 |
| ST1 - V2 | -1780.9 | 11.8 | -2.1 | -1250.4 | 28.0 | -1.5 | -530.5 | -0.63 |
| ST1 - V3 | 186.2 | 6.0 | 0.2 | -1638.6 | 28.6 | -2.0 | 1824.8 | 2.17 |
| ST2 - H1 | 846.1 | 16.3 | 0.3 | 44.9 | 38.0 | 0.0 | 801.1 | 0.24 |
| ST2 - H2 | 381.5 | 17.2 | 0.1 | 1929.0 | 34.6 | 0.6 | -1547.6 | -0.46 |
| ST2 - H3 | -635.3 | 19.4 | -0.2 | 2094.9 | 44.7 | 0.6 | -2730.2 | -0.81 |
| ST2 - V1 | 2051.7 | 20.4 | 0.6 | 3881.7 | 43.6 | 1.2 | -1830.0 | -0.54 |
| ST2 - V2 | -153.5 | 15.3 | 0.0 | 64.3 | 65.3 | 0.0 | -217.8 | -0.06 |
| ST2 - V3 | 1754.8 | 22.1 | 0.5 | 1745.8 | 67.4 | 0.5 | 9.0 | 0.00 |

Figure 101 - Balancing

## Static test – Local basis

Static tests in the local basis are located in the SVN:

/seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/

- LHO_ISI_BSC8_Offset_Local_Drive_20120131.mat

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 3811 | 1531 | 1523 | 33 | -8 | -7 |
|  | ST1 - H2 | 1507 | 3727 | 1492 | -2 | 7 | -3 |
|  | ST1 - H3 | 1503 | 1493 | 3720 | 11 | -12 | 7 |
|  | ST1 - V1 | 70 | -133 | 91 | 2869 | -496 | -470 |
|  | ST1 - V2 | 77 | 28 | -139 | -469 | 2881 | -490 |
|  | ST1 - V3 | -146 | 77 | 22 | -481 | -470 | 2890 |

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2138 | 337 | 332 | 6 | -59 | 14 |
|  | ST2 - H2 | 318 | 2159 | 334 | 14 | -34 | 1 |
|  | ST2 - H3 | 340 | 348 | 2255 | -2 | -27 | 5 |
|  | ST2 - V1 | 55 | 119 | -168 | 2614 | 315 | 0 |
|  | ST2 - V2 | -174 | 75 | 99 | 16 | 2581 | 314 |
|  | ST2 - V3 | 108 | -172 | 71 | 328 | -2 | 2566 |

Table 34 – Static test local basis

## Static Test – Cartesian basis

Static tests in the local basis are located in the SVN:

/seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/

- LHO_ISI_BSC8_Offset_Cartesian_Drive_20120131.mat

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST1 - X | ST1 - Y | ST1 - Z | ST1 – RX | ST1 - RY | ST1 - RZ |
| Actuators | ST1 - X | 1503 | -2 | -1 | -13 | -13 | 36 |
|  | ST1 - Y | 1 | 1483 | -3 | 6 | -10 | 4 |
|  | ST1 - Z | 9 | 3 | 639 | 6 | -5 | -9 |
|  | ST1 - RX | 4 | 295 | -1 | 2449 | -12 | -13 |
|  | ST1 - RY | -294 | 14 | -5 | -2 | 2446 | -8 |
|  | ST1 - RZ | 19 | 4 | 2 | -3 | 2 | 2854 |
|  |  | Sensors |
|  |  | ST2 - X | ST2 - Y | ST2 - Z | ST2 – RX | ST2 - RY | ST2 - RZ |
| Actuators | ST2 - X | 1228 | -3 | 17 | 12 | 21 | 21 |
|  | ST2 - Y | -13 | 1250 | 9 | -7 | 5 | -33 |
|  | ST2 - Z | 11 | 8 | 997 | 6 | -5 | -6 |
|  | ST2 - RX | 6 | 3 | 10 | 3760 | 6 | -6 |
|  | ST2 - RY | 27 | 11 | 12 | 12 | 3724 | 6 |
|  | ST2 - RZ | 20 | -28 | 5 | -13 | 7 | 2353 |

Table 35 - Static test - Cartesian basis drive – Cartesian results

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - X | 1548 | -728 | -718 | 31 | -6 | -2 |
|  | ST1 - Y | 7 | 1284 | -1279 | 6 | 9 | 1 |
|  | ST1 - Z | 4 | -18 | -12 | 647 | 642 | 634 |
|  | ST1 - RX | -2 | 114 | -143 | -2337 | 1998 | 356 |
|  | ST1 - RY | -152 | 78 | 52 | -921 | -1569 | 2517 |
|  | ST1 - RZ | 2772 | 2734 | 2732 | 10 | 5 | 12 |
|  |  | Sensors |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - X | 607 | -1227 | 643 | 26 | 6 | 13 |
|  | ST2 - Y | 1040 | -3 | -1099 | 13 | 3 | 9 |
|  | ST2 - Z | 3 | -14 | 6 | 1005 | 979 | 964 |
|  | ST2 - RX | -239 | -6 | 236 | -2152 | 2234 | -26 |
|  | ST2 - RY | 148 | -281 | 142 | -1297 | -1225 | 2522 |
|  | ST2 - RZ | 1609 | 1601 | 1657 | 22 | -20 | 9 |

Table 36 - Static test - Cartesian basis drive - Local results

## Range of motion – Actuators

Range of motion tests results are located in the SVN at:

/seismic/BSC-ISI/H2/ITMY/Data/Static_Tests/

- LHO_ISI_BSC8_Range_Of_Motion_20120131.mat

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -14891 | 1374 | 16765 | 31657 | 37.7 |
| ST1 - H2 | -17010 | -1244 | 14703 | 31713 | 37.8 |
| ST1 - H3 | -15615 | 311 | 13784 | 29399 | 35.0 |
| ST1 - V1 | -14793 | -2503 | 9799 | 24592 | 29.3 |
| ST1 - V2 | -13525 | -1137 | 11211 | 24737 | 29.4 |
| ST1 - V3 | -13962 | -1558 | 10799 | 24762 | 29.5 |
| ST2 - H1 | -8929 | 218 | 9361 | 18290 | 5.4 |
| ST2 - H2 | -7220 | 1962 | 11144 | 18364 | 5.5 |
| ST2 - H3 | -7664 | 1877 | 11404 | 19069 | 5.7 |
| ST2 - V1 | -7189 | 3929 | 15035 | 22225 | 6.6 |
| ST2 - V2 | -10908 | 196 | 11289 | 22197 | 6.6 |
| ST2 - V3 | -9251 | 1816 | 12850 | 22101 | 6.6 |

Table 37 - Range of motion

## Linearity Test

Not performed

## Transfer functions

Last set of measurements:

Seismic/BSC-ISI/H2/ITMY/Data/Transfer_Functions/Measurements/Undamped/

- LHO_ISI_BSC8_Data_L2L_700mHz_10Hz_ST1_ST2_20120201-041506.mat

- LHO_ISI_BSC8_Data_L2L_10Hz_100Hz_ST1_ST2_20120201-004803.mat

- LHO_ISI_BSC8_Data_L2L_100Hz_500Hz_ST1_ST2_20120131-230304.mat

- LHO_ISI_BSC8_Data_L2L_500Hz_1000Hz_ST1_ST2_20120131-213833.mat

Transfer functions figures are located in the SVN

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/H2/ITMY/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LHO_ISI_BSC8_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2012_01_31.fig

- LHO_ISI_BSC8_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2012_01_31.fig

- LHO_ISI_BSC8_TF_L2L_Raw_from_ST1_ACT_to_ST1_T240_2012_01_31.fig

- LHO_ISI_BSC8_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2012_01_31.fig

- LHO_ISI_BSC8_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2012_01_31.fig

## State of health of the T240 on February 6, 2012

The state of health was realized on February 6, 2012.

Then snapshots of the SOF are presented below:

Corner 1

Corner 2

Corner 3

## Validation

After reviewing the testing results on February 3, 2012, the ISI passed the ISI-phase II tests. The Dome was put on.

Conclusion

The ISI-BSC8 was moved from the Staging building to the LVEA test stand on June 30, 2011.

This document presents the series of tests (Phase II) performed on the ISI-BSC8 (ITMY) in the LVEA. Some tests were done with and without the QUAD and the folding miror. Tests started on July 8, 2011. Damping loops were closed for the first time on July 22, 2011 and Isolation loops on August 4, 2011.

The cartridge installation occurred on Friday 9 December 2011 (LHO Alog 1887). TMDs and vibrations absorbers were installed mid January 2012.

The dome of the chamber was put on February 2, 2012.The first aLIGO BSC-ISI built in the LHO