LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1100856-V2 	                 LIGO                   December 17th, 2012

L1 BSC 1 BSC-ISI, Pre-integration Testing report,

Phase II

E1100856-v2

Adrien Le Roux, Michael Puma, Céline Ramet

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

Introduction	4

1.	Phase II-a	5

1.	Hardware changes	5

1.	CPS – E1100369	5

2.	GS13 – E1100740	5

3.	L4C – E1100740	5

4.	T240 – E1100740	5

5.	Cables – E1100822	5

6.	Misc	5

2.	Electronic Inventory	6

3.	Models Modifications	6

4.	Mass distribution	7

1.	Seismic	7

2.	Suspension	7

3.	Misc	7

4.	Total	7

5.	Basic functionalities just after installing the BSC-ISI on the test stand	8

1.	Pressure sensors	8

2.	Spectra	8

3.	Actuators-cables resistance	9

4.	Offsets CPS Unlocked vs locked	10

5.	Offset local drive	10

6.	Offset Cartesian drive	11

7.	Range of motion	12

8.	Linearity test	12

6.	Transfer functions and Comparison with measurements done in the staging building.	14

1.	At the end station	14

2.	Comparisons with measurements in the staging building –not done yet	17

7.	Conclusion Phase II-a	19

2.	Phase II-b	20

1.	Hardware changes	20

1.	CPS – E1100369	20

2.	GS13 – E1100740	20

3.	L4C – E1100740	20

4.	T240 – E1100740	20

5.	Cables – E1100822	20

6.	Misc	20

2.	Electronic Inventory	20

3.	Models Modifications	20

4.	Mass distribution –not updated yet	21

1.	Seismic	21

2.	Suspension	22

3.	Misc	22

4.	Total	23

5.	Basic functionalities just after installing the BSC-ISI on the test stand	24

5.	Pressure sensors	24

6.	Spectra	24

7.	Actuators-cables resistance	25

8.	Offsets CPS Unlocked vs locked	26

9.	Offset local drive	26

10.	Offset Cartesian drive	27

11.	Range of motion	28

12.	Linearity test	28

6.	Transfer functions and Comparison with measurements done in the staging building.	30

13.	At the end station	30

14.	Comparisons with measurements on test stand-not done yet	33

7.	Conclusion Phase II-b	38

Conclusion	39

Introduction

The BSC-ISI testing is performed in three phases:

1) BSC-ISI, Pre-integration Testing, Phase I (post-assembly, in the staging building)

2) BSC-ISI, Pre-integration Testing, Phase II: Final tests done before insertion in the chamber
3) BSC-ISI, Integration Phase Testing: Procedure and results related to the commissioning in the chamber.

The ISI-BSC2 was moved from the Staging building to the LVEA on July 17th, 2012. Then the ISI was placed on the test stand a few days later.

This document presents results from the series of tests (Phase II) performed on the ISI-BSC2 (BS) in the corner station. The tests were done with the ITMY quadruple suspension installed in its final version.

First tests started on October 24th, 2012. The first testing phase (II-a validation before cartridge installation) was completed on October 31st 2012.

All results are posted on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/

The following types of documents can be found in the SVN:

- Data location

- Figures location

# Phase II-a

## Hardware changes

### CPS – E1100369

No change.

### GS13 – E1100740

GS13 have not been replaced since phase I testing in the staging building.

### L4C – E1100740

L4C have not been replaced since phase I testing in the staging building.

### T240 – E1100740

T240s have not been replaced since phase I testing in the staging building.

### Cables – E1100822

No cable have been replaced since phase I testing in the staging building.

### Misc

No hardware changes since phase I testing in the staging building.

## Electronic Inventory

This table reports the electronic equipment used in the LVEA.

| Hardware | Ligo reference | S/N |
| --- | --- | --- |
| Interface Chassis Pod 1 | D1002432 | S1201326 |
| Interface Chassis Pod 2 | D1002432 | S1201325 |
| Interface Chassis Pod 3 | D1002432 | S1201324 |
| Anti-aliasing Chassis Pod 1 | D1002693 | S1203001 |
| Anti-aliasing Chassis Pod 2 | D1002693 | S1203000 |
| Anti-aliasing Chassis Pod 3 | D1002693 | S1202999 |
| Binary Input Chassis 1 | D1001726 | S1101295 |
| Binary Input Chassis 2 | D1001726 | S1101296 |
| Binary Output Chassis | D1001728 | S1101327 |
| T240 Interface Pod 1 | D1002694 | S1201402 |
| T240 Interface Pod 2 | D1002694 | S1201391 |
| T240 Interface Pod 3 | D1002694 | S1201386 |
| Anti-image Chassis | D1000305 | S1203007 |
| Coil driver Pod 1 | D0902744 | S1103328 |
| Coil driver Pod 2 | D0902744 | S1103322 |
| Coil driver Pod 3 | D092744 | S1103311 |
| Expansion chassis | L1seibsc1 | S1001147 |

Table 2 - Electronic inventory

## Models Modifications

No model modifications were done between the beginning and the end of Phase 2a testing.

## Mass distribution

This final mass distribution will be presented once all elements will be installed on the ISI (during phase II-b). These elements are the vibration absorbers on stage 1 and the QUAD structure.

### Seismic

Stage 1

| Stage 1 | HighBay (lbs) | LVEA (lbs) | LVEA (kgs) |
| --- | --- | --- | --- |
| Corner 1 | 56.90 | 59.00 | 26.76 |
| Corner 2 | 32.50 | 43.72 | 19.83 |
| Corner 3 | 43.00 | 43.72 | 19.83 |
| Total | 132.40 | 146.44 | 66.42 |

Stage 2

The total of masses on Stage 2 is 1453.65 lbs (=663.90 kgs).

### Suspension

The quad structure was weighed to be:

|  | Weight (lbs) | Weight (kgs) |
| --- | --- | --- |
| Upper structure | 266 | 120.66 |
| Lower structure | 531 | 240.86 |
| Total | 797 | 361.5128 |

### Misc

20 dog clamps at 1.26 lbs each create an extra load of 25.2 lbs (=11.47 kgs)

### Total

Nominal mass hanging on stage 0-1 blades (without stage 2): 912Kg – 2010lb

Nominal mass hanging on stage 1-2 blades: 2830Kg – 6239lb

Nominal payload on stage 1: 109Kg – 240lb

Nominal payload on stage 2: 1185Kg – 2612lb

|  |  |  | Staging Bldg | LVEA |  |
| --- | --- | --- | --- | --- | --- |
|  | Plan | 6/6/2012 | Detail | Overall | Difference |
| Stage 1 (kgs) | 108.86 | 60.06 | 66.42 | 66.42 | 6.37 |
| Stage 2 (kgs) | Masses | 1183.42 | 1071.48 | 663.90 | 1036.88 | -34.60 |
|  | Suspension | N/A | N/A | 361.51 |  |  |
|  | Miscellaneous | N/A | N/A | 11.47 |  |  |
| Total (kgs) | 1292.28 | 1131.53 | 1103.30 | 1103.30 | -28.23 |

Stage 1 difference in weight passes requirement (10 kgs) but Stage 2 fails because we are lighter by 34.6 kgs and the requirement is to be within 25 kgs.

Because of the relatively high uncertainty of the weight of the temporary masses, we decide to waive that test.

Test result:                                              Passed:                Failed:         .   W Waived:   X_

## Basic functionalities just after installing the BSC-ISI on the test stand

### Pressure sensors

All pressure sensors are working.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Static_Tests/
LLO_ISI_ITMY_Pressure_Sensors_Check_Calibrated_2012_10_24_163944.mat

|  | Pressure (KPa) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-P | 100.65 | 100.38 | 99.69 |
| ST1-L4C-D | 0.00 | -0.43 | 0.76 |
| ST1-GS13-P | 100.53 | 100.59 | 99.83 |
| ST1-GS13-D | -0.56 | -0.16 | -0.08 |
| ST1-T240-P | 155.14 | 155.22 | 155.29 |

Table 3 - Geophones Pressure sensors

Note/comment about this test: N/A.

Test result:                                              Passed:    X             Failed:         .   W Waived:   _

### Spectra

Spectra of the instrument can be found in the SVN at:

seismic/BSC-ISI/L1/ITMY/Data/Spectra/Undamped/

L1_ISI_ITMY_ASD_m_LOC_CPS_T240_L4C_GS13_2012_10_28_175304.mat

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Figures/Spectra/Undamped/
L1_ISI_ITMY_ASD_m_LOC_CPS_T240_L4C_GS13_2012_10_28_175304.fig

Figure 1 - Spectra inboard instruments - ISI Unlocked

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Actuators-cables resistance

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Static_Tests/

- L1_ISI_ITMY_Actuators_Resistance_20121026T175017.mat

|  | Stage 1 | Stage 2 |
| --- | --- | --- |
|  | H1 | H2 | H3 | V1 | V2 | V3 | H1 | H2 | H3 | V1 | V2 | V3 |
| Script | 8.10 | 8.18 | 7.86 | 8.07 | 7.87 | 7.64 | 11.74 | 12.06 | 11.57 | 11.83 | 11.77 | 11.48 |

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offsets CPS Unlocked vs locked

The table is not perfectly balanced but it is considered sufficiently good to perform the series of test before the cartridge installation. A fine balancing will be done during phase II-b.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Static_Tests/

- L1_ISI_ITMY_CPS_Read_Back_ISI_Locked_2012_10_26_142944.mat

- L1_ISI_ITMY_CPS_Read_Back_ISI_Unlocked_2012_10_26_170205.mat

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | -872.99 | 46.92 | -104.30 | 5.47 | -768.70 | -0.92 |
| ST1 - H2 | 1234.01 | 30.98 | 98.13 | 5.64 | 1135.88 | 1.35 |
| ST1 - H3 | -1121.20 | 66.05 | 386.94 | 6.78 | -1508.13 | -1.80 |
| ST1 - V1 | -213.83 | 47.66 | 476.14 | 6.05 | -689.97 | -0.82 |
| ST1 - V2 | 639.41 | 58.47 | -143.73 | 4.69 | 783.13 | 0.93 |
| ST1 - V3 | -805.52 | 50.14 | -78.90 | 6.17 | -726.62 | -0.87 |
| ST2 - H1 | -656.49 | 70.03 | 1058.86 | 28.82 | -1715.36 | -0.51 |
| ST2 - H2 | 1568.78 | 69.74 | 244.71 | 20.47 | 1324.07 | 0.39 |
| ST2 - H3 | -829.97 | 64.96 | 154.99 | 17.47 | -984.96 | -0.29 |
| ST2 - V1 | -755.05 | 121.50 | -1091.32 | 29.12 | 336.28 | 0.10 |
| ST2 - V2 | -1091.14 | 146.86 | -256.65 | 30.01 | -834.49 | -0.25 |
| ST2 - V3 | 1049.88 | 142.61 | 601.49 | 23.19 | 448.39 | 0.13 |

Table 4 - Locked vs Unlocked Position

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offset local drive

Note: Due to longer cables, offsets measured by CPS for a 7000 count drive are slightly lower than offsets measured in the staging building.

Results of this test can be found in the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Static_Tests/

- L1_ISI_ITMY_Offset_Local_Drive_20121026.mat

|  | Sensors |
| --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 3435.81 | 1393.05 | 1379.44 | 27.70 | -8.86 | 13.07 |
|  | ST1 - H2 | -1375.31 | -3451.56 | -1398.02 | 27.61 | -39.66 | 3.51 |
|  | ST1 - H3 | 1367.46 | 1373.53 | 3391.12 | 16.09 | -22.26 | 33.75 |
|  | ST1 - V1 | 40.80 | -128.55 | 78.11 | 2652.58 | -453.54 | -425.14 |
|  | ST1 - V2 | -50.27 | -18.66 | 144.13 | 457.43 | -2671.08 | 458.83 |
|  | ST1 - V3 | -115.91 | 59.43 | 30.40 | -444.71 | -449.29 | 2740.44 |

Table 5 - Static Tests – Local to Local - Stage 1

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2026.36 | 348.28 | 294.00 | 32.87 | -8.28 | 31.22 |
|  | ST2 - H2 | 326.31 | 2025.60 | 304.39 | -7.17 | -14.01 | 26.68 |
|  | ST2 - H3 | 103.66 | -180.78 | 65.77 | 312.23 | 9.60 | 2406.21 |
|  | ST2 - V1 | 107.06 | 94.45 | -101.37 | 2374.45 | 289.06 | 47.35 |
|  | ST2 - V2 | -157.91 | 110.59 | 101.14 | 25.24 | 2446.20 | 332.52 |
|  | ST2 - V3 | 308.04 | 306.90 | 2059.75 | 25.48 | -6.26 | 3.50 |

Table 6 - Static Tests – Local to Local - Stage 2

It is apparent from this test that St1 corner 2 actuator cables are plugged in the opposite direction and that St 2 corner 3 actuator cables are flipped. Those issues will be corrected after the cartridge install.

Test result:                                              Passed:                 Failed:         .   W Waived:   _X_

### Offset Cartesian drive

The test was not performed because it only tests the matrices, which can be changed at any time.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

### Range of motion

The range of motion of the table is measured by pushing on the table in a direction collinear to the CPS. The Static tests results can be found on the SVN at:

seismic/BSC-ISI/L1/ITMY/Data/Static_Tests/

- L1_ISI_ITMY_Range_Of_Motion_20121030.mat

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -15056.15 | -900.00 | 13842.44 | 28898.59 | 34.40 |
| ST1 - H2 | 15961.27 | 1181.00 | -13604.83 | -29566.10 | -35.20 |
| ST1 - H3 | -15733.94 | -1177.00 | 13387.61 | 29121.54 | 34.67 |
| ST1 - V1 | -11414.23 | -46.00 | 11275.96 | 22690.19 | 27.01 |
| ST1 - V2 | 12239.13 | 765.00 | -10707.36 | -22946.49 | -27.32 |
| ST1 - V3 | -12298.51 | -504.00 | 11242.83 | 23541.35 | 28.03 |
| ST2 - H1 | -9290.43 | -600.00 | 8076.13 | 17366.56 | 5.17 |
| ST2 - H2 | -7154.60 | 1528.00 | 10171.62 | 17326.21 | 5.16 |
| ST2 - H3 | -9645.89 | 1437.00 | 7952.85 | 17598.75 | 5.24 |
| ST2 - V1 | -10543.42 | -410.00 | 9702.18 | 20245.60 | 6.03 |
| ST2 - V2 | -11129.58 | -655.00 | 9764.19 | 20893.77 | 6.22 |
| ST2 - V3 | -8843.81 | -834.00 | 11589.89 | 20433.70 | 6.08 |

Table 7 - Range of motion - Actuator drive in the LVEA

Note that this table was corrected after the fact to adjust based on the miscabling of Stage 2 corner 3. (Stage 2 H3 excitation at 1035644646/1035644683 and V3 at 1035644849/1035644880)

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Linearity test

The data of the linearity test can be found on the SVN at:

seismic/BSC-ISI/L1/ITMY/Data/Linearity_Test/

- L1_ISI_ITMY_Linearity_test_20121031.mat

The figures of the linearity test can be found on the SVN at:

seismic\BSC-ISI\H2\BS\Data\Figures\Linearity_Test\

- L1_ISI_ITMY_Linearity_test_20121031.fig

Figure 2 - Linearity test –L1 - BS – In LVEA

|  |  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- | --- |
| Stage 1 | ST1 - H1 | 0.491925 | -895.194 | 0.490326 | 0.326181 |
|  | ST1 - H2 | 0.493962 | -1178.07 |  | 0.741565 |
|  | ST1 - H3 | 0.48509 | -1138.05 |  | -1.06775 |
|  | ST1 - V1 | 0.377912 | -88.2464 | 0.384375 | -1.68135 |
|  | ST1 - V2 | 0.382807 | -745.375 |  | -0.40799 |
|  | ST1 - V3 | 0.392406 | -604.221 |  | 2.08934 |
| Stage 2 | ST2 - H1 | 0.290144 | -616.071 | 0.291044 | -0.30894 |
|  | ST2 - H2 | 0.289504 | 1554.751 |  | -0.52891 |
|  | ST2 - H3 | 0.293482 | -826.434 |  | 0.837849 |
|  | ST2 - V1 | 0.338293 | -448.857 | 0.34323 | -1.43833 |
|  | ST2 - V2 | 0.350097 | -821.769 |  | 2.000706 |
|  | ST2 - V3 | 0.341299 | 1257.176 |  | -0.56237 |

Table 8 - Slope – Offset Linearity test

Note that this table was corrected after the fact to adjust based on the mis-cabling of Stage 2 corner 3 and stage 1 corner 2.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

## Transfer functions and Comparison with measurements done in the staging building.

### At the end station

The parameters for the measurements in the LVEA are slightly different from those in the staging building. We chose to have weaker excitation but longer averages in an effort to reduce risk of the attached suspension.

At this point, only the tuned mass dampers on the spring are installed (No vibration absorbers, no Viton under the keel masses, No vibration absorbers on the BS)

Measurements data can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMY/Data/Transfer_Functions/Measurements/Undamped:

- LLO_ISI_BSC1_Data_L2L_10mHz_100mHz_ST1_ST2_20121029-214002.mat

- LLO_ISI_BSC1_Data_L2L_100mHz_700mHz_ST1_ST2_20121029-161711.mat

- LLO_ISI_BSC1_Data_L2L_700mHz_10Hz_ST1_ST2_20121027-000402.mat

- LLO_ISI_BSC1_Data_L2L_10Hz_100Hz_ST1_ST2_20121026-212711.mat

- LLO_ISI_BSC1_Data_L2L_100Hz_500Hz_ST1_ST2_20121026-200720.mat

- LLO_ISI_BSC1_Data_L2L_500Hz_1000Hz_ST1_ST2_20121026-190253.mat

Data after processing can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMY/Data/Transfer_Functions/Simulations/Undamped

- LLO_ISI_BSC1_TF_L2L_Raw_10mHz_1000Hz_2012_10_29.mat

The transfer functions can be found in the SVN at:

seismic/BSC-ISI/L1/ITMY/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_BSC1_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2012_11_07.fig

- LLO_ISI_BSC1_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2012_11_07.fig

- LLO_ISI_BSC1_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2012_11_07.fig

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2012_11_07.fig

Note 1:  The transfer functions are measured from the Output filters bank excitation point to the input (IN1) of the input filters bank. The transfer functions presented below are raw transfer functions without any electronic compensation.

Note 2: The L4Cs are out of phase (should be -90 before 1Hz).  A minus sign is added in the calibration filters that convert count to nm/s.

Note 3: On the ST1-ACT-H to ST1-CPS-H transfer functions, we can see the first resonances of the LVEA test stand at 21.6 Hz, 23.8Hz and 31Hz (matches within a few Hz our BSC 2 results).

Note 4: Even without vibration absorbers on the Quad structure, the first resonance visible on Stage 2 is at 154 Hz.

Figure 3 - TF ST1 ACT to ST1 CPS

Figure 4 - TF ST1 ACT to ST1 L4C

Figure 5 - TF ST1 ACT to ST1 T240

Figure 6 - TF ST2 ACT to ST2 CPS

Figure 7 - TF ST2 ACT to ST2 GS13

### Comparisons with measurements in the staging building –not done yet

The script used to compare transfer function can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMY/Scripts/Misc

- Comparison_TF_L2L_LLO_ISI_BSC2.m

The figure that shows the comparison between the transfer functions of the staging building and the LVEA are located in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMY/Data/Figures/Transfer_Functions/Measurements/Comparison/

- COMP_LVEA_L1_ISI_ITMY_ST1_CPS_2012_10_31_.fig

- COMP_LVEA_L1_ISI_ITMY_ST1_L4C_2012_10_31_.fig

- COMP_LVEA_L1_ISI_ITMY_ST2_CPS_2012_10_31_.fig

- COMP_LVEA_L1_ISI_ITMY_ST2_GS13_2012_10_31_.fig

Main differences are:

- The DC gains (cables resistance is different due to the length difference)

- Resonances of the rigid body modes [1; 10]Hz (different payload) It is especially visible from 1 to 2 Hz on all sensors.

- ST1-CPS resonances different in the staging building and EY (Test stand short legs vs long legs).

- Resonance at 20.5 Hz on stage 2 sensors (4th vertical mode of the quad? –cf https://lhocds.ligo-wa.caltech.edu/wiki/Resonances)

- Similar at high frequencies

Figure 8 - Transfer functions comparison - ST1 ACT to ST1 CPS

Figure 10 - Transfer functions comparison - ST1 ACT to ST1 L4C

Figure 12 - Transfer functions comparison - ST2 ACT to ST2 CPS

Figure 15 - Transfer functions comparison – ST2 ACT to ST2 GS13

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

## Conclusion Phase II-a

All results appear satisfying, apart from a few test waived (will be done during phase II-b):

- Added mass on Stage 2 is more 36 kgs lighter than it used to be during testing (requirement is 25 kgs). We think the uncertainty on the loaded payload is such that this is not worrisome.

- Static tests in the Cartesian basis (redundant with the static test in the local basis)

- Linearity test: some corners have slightly different slopes but nothing strikingly alarming. We have seen in the past (HAM-ISI phase 2 testing) stronger slopes different, blamed on different cable lengths inside the LVEA.

- St 1 corner 2 actuators are wired in the opposite direction and St 2 corner 3 actuator cables are flipped. This can be corrected after the cartridge install.

Test result:                                              Passed:   X              Failed:         .   W Waived:   _

# Phase II-b

## Hardware changes

### CPS – E1100369

No change.

### GS13 – E1100740

GS13s have not been replaced since phase II-a testing on the test stand.

### L4C – E1100740

L4Cs have not been replaced since phase II-a testing on the test stand.

### T240 – E1100740

T240s have not been replaced since phase II-a testing on the test stand.

### Cables – E1100822

2 T240 extension cables were swapped in order to allow connecting those cables to the vacuum feedthrus.

### Misc

No hardware changes since II-a testing on the test stand.

## Electronic Inventory

No change since phase II-a testing on the test stand.

## Models Modifications

The model was rebuilt 3 times between mid October and mid November:

- October 25th 2012: simple rebuild

- November 12th 2012 latest change from MIT/Hanford

- November 13th 2012:terminate IPC

## Mass distribution –not updated yet

This final mass distribution will be presented once all elements will be installed on the ISI (during phase II-b). These elements are the vibration absorbers on stage 1 and the QUAD structure.

### Seismic

Stage 1

| 12/17/2012 | D0902612 | D0902613 | D0902616 | D1001760 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 1 | 2 | 3 | 4 | 5 |  |  |  |
|  | 12 | 15 | 3.5 | 1 | 2 | 5 | 0.5 | 15.86 | Lbs | kgs |
| C1-1 |  |  |  |  |  |  |  |  | 0 | 0 |
| C1-2 |  | 1 |  |  |  |  |  |  | 15 | 6.81 |
| C1-3 |  |  |  |  |  | 2 |  | 2 | 41.72 | 18.94088 |
| C2-1 |  |  |  |  |  |  |  |  | 0 | 0 |
| C2-2 |  |  |  |  |  |  |  |  | 0 | 0 |
| C2-3 |  |  | 1 | 1 |  | 1 |  | 2 | 41.22 | 18.71388 |
| C3-1 |  |  |  |  |  |  |  |  | 0 | 0 |
| C3-2 |  |  |  |  |  |  |  |  | 0 | 0 |
| C3-3 |  |  | 1 |  |  | 2 |  | 2 | 45.22 | 20.52988 |
| Stage 1 | 0 | 1 | 2 | 1 | 0 | 5 | 0 | 6 | 143.16 | 64.99464 |

| Stage 1 | HighBay (lbs) | LVEA (lbs) | LVEA (kgs) | After Cartridge (lbs) | After Cartridge (kgs) |
| --- | --- | --- | --- | --- | --- |
| Corner 1 | 56.90 | 59.00 | 26.76 | 56.72 |  |
| Corner 2 | 32.50 | 43.72 | 19.83 | 41.22 | 18.71 |
| Corner 3 | 43.00 | 43.72 | 19.83 | 45.22 | 20.53 |
| Total | 132.40 | 146.44 | 66.42 | 143.16 | 64.99 |

Stage 2

The total of masses on Stage 2 is 1453.65 lbs (=663.90 kgs).

| 12/17/2012 | D1003136 | D1003161 | D071200 |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 0 | 1 | 2 | 3 | 4 | 5 | 6 |  |  |
|  | 50 | 47.62 | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| Keel | 12 |  |  |  |  |  |  |  |  | 600 | 272.40 |
| Optical |  | 16 |  |  |  |  |  |  |  | 761.92 | 345.91 |
| E-1 |  |  |  | 2 |  | 1 |  |  |  | 6.7 | 3.04 |
| E-2 |  |  | 3 | 1 |  |  |  | 2 | 1 | 61.3 | 27.83 |
| E-3 |  |  | 1 | 2 |  | 3 |  |  | 1 | 43.5 | 19.75 |
| F1 |  |  |  |  |  |  |  |  |  | 0 | 0.00 |
| F2 |  |  |  |  |  |  |  |  |  | 0 | 0.00 |
| F3 |  |  |  |  |  |  |  |  |  | 0 | 0.00 |
| Stage 2 | 12 | 16 | 4 | 5 | 0 | 4 | 0 | 2 | 2 | 1473.42 | 668.93 |

| Stage 2 | HighBay (lbs) | LVEA (lbs) | LVEA (kgs) | After Cartridge (lbs) | After Cartridge (kgs) |
| --- | --- | --- | --- | --- | --- |
| Total | 0 | 1453.65 | 663.9 | 1473.42 | 668.93 |

### Suspension

The quad structure was weighed to be:

|  | Weight (lbs) | Weight (kgs) |
| --- | --- | --- |
| Upper structure | 266 | 120.66 |
| Lower structure | 531 | 240.86 |
| Total | 797 | 361.5128 |

### Misc

20 dog clamps at 1.26 lbs each create an extra load of 25.2 lbs (=11.47 kgs)

### Total

Nominal mass hanging on stage 0-1 blades (without stage 2): 912Kg – 2010lb

Nominal mass hanging on stage 1-2 blades: 2830Kg – 6239lb

Nominal payload on stage 1: 109Kg – 240lb

Nominal payload on stage 2: 1185Kg – 2612lb

|  |  |  | Assembly | LVEA | After Cartridge | Difference with Assembly |
| --- | --- | --- | --- | --- | --- | --- |
|  | Plan | 6/6/2012 | Detail | Overall | Detail | Overall |  |
| Stage 1 (kgs) | 108.86 | 60.06 | 66.42 | 66.42 | 64.99 | 64.99 | +4.93 |
| Stage 2 (kgs) | Masses | 1183.42 | 1071.48 | 663.90 | 1036.88 | 668.93 | 1042.14 | -29.34 |
|  | Suspension | N/A | N/A | 361.51 |  | 361.51 |  |  |
|  | Miscellaneous | N/A | N/A | 11.47 |  | 11.47 |  |  |
| Total (kgs) | 1292.28 | 1131.53 | 1103.30 | 1103.30 | 1107.13 | 1107.13 |  |

Stage 1 difference in weight passes requirement (10 kgs), whereas stage 2 fails as the difference exceeds the 25 kg change.

Because of the relatively high uncertainty of the weight of the temporary masses, we decide to waive that test.

Test result:                                              Passed:                Failed:         .   W Waived:   X_

## Basic functionalities just after installing the BSC-ISI on the test stand

### Pressure sensors

All pressure sensors are working.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Static_Tests/
LLO_ISI_ITMY_Pressure_Sensors_Check_Calibrated_2012_12_16_180617.mat

|  | Pressure (KPa) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-P | 100.31 | 100.00 | 99.36 |
| ST1-L4C-D | -0.18 | -0.47 | 0.81 |
| ST1-GS13-P | 100.34 | 100.44 | 99.67 |
| ST1-GS13-D | -0.55 | -0.19 | -0.13 |
| ST1-T240-P | 154.28 | 154.38 | 154.50 |

Table 3 - Geophones Pressure sensors

Note/comment about this test: N/A.

Test result:                                              Passed:    X             Failed:         .   W Waived:   _

### Spectra

Spectra of the instrument can be found in the SVN at:

seismic/BSC-ISI/L1/ITMY/Data/Spectra/Undamped/

L1_ISI_ITMY_ASD_m_LOC_CPS_T240_L4C_GS13_2012_12_16_222822.mat

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Figures/Spectra/Undamped/
L1_ISI_ITMY_ASD_m_LOC_CPS_T240_L4C_GS13_2012_12_16_222822.fig

Figure 1 - Spectra inboard instruments - ISI Unlocked

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Actuators-cables resistance

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Static_Tests/

- L1_ISI_ITMY_Actuators_Resistance_20121026T175017.mat

|  | Stage 1 | Stage 2 |
| --- | --- | --- |
|  | H1 | H2 | H3 | V1 | V2 | V3 | H1 | H2 | H3 | V1 | V2 | V3 |
| Script |  |  |  |  |  |  |  |  |  |  |  |  |

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offsets CPS Unlocked vs locked

The table is not perfectly balanced but it is considered sufficiently good to perform the series of test before the cartridge installation. A fine balancing will be done during phase II-b.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Static_Tests/

- L1_ISI_ITMY_CPS_Read_Back_ISI_Locked_2012_12_17_094105.mat

- L1_ISI_ITMY_CPS_Read_Back_ISI_Unlocked_2012_12_05_171334.mat

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | -120.098 | 4.818902 | -844.736 | 27.99514 | 724.6376 | 0.862664 |
| ST1 - H2 | -416.744 | 5.614173 | -1234.19 | 32.54445 | 817.4503 | 0.973155 |
| ST1 - H3 | -179.702 | 8.446006 | -975.728 | 30.68611 | 796.0263 | 0.94765 |
| ST1 - V1 | 604.8375 | 4.357051 | 1087.314 | 55.51869 | -482.476 | -0.57438 |
| ST1 - V2 | 175.6026 | 7.864287 | -940.162 | 33.23999 | 1115.764 | 1.328291 |
| ST1 - V3 | 891.1182 | 7.266689 | -187.756 | 41.18756 | 1078.875 | 1.284375 |
| ST2 - H1 | -104.458 | 8.561111 | -888.362 | 41.0035 | 783.904 | 0.233305 |
| ST2 - H2 | -811.872 | 6.992354 | 683.6736 | 48.2142 | -1495.55 | -0.4451 |
| ST2 - H3 | -1557.67 | 9.110095 | 3242.06 | 38.5595 | -4799.73 | -1.42849 |
| ST2 - V1 | 2178.387 | 14.96533 | 1006.21 | 117.4369 | 1172.177 | 0.348862 |
| ST2 - V2 | 1081.196 | 15.60391 | -2935.81 | 84.84448 | 4017.011 | 1.195539 |
| ST2 - V3 | 825.0956 | 16.02009 | 306.4185 | 91.47976 | 518.6771 | 0.154368 |

Table 4 - Locked vs Unlocked Position

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offset local drive

Note: Due to longer cables, offsets measured by CPS for a 7000 count drive are slightly lower than offsets measured in the staging building.

Results of this test can be found in the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Static_Tests/

- L1_ISI_ITMY_Offset_Local_Drive_20121205.mat

|  | Sensors |
| --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 3782.413 | 1537.945 | 1503.76 | 8.15604 | -16.9929 | 31.74798 |
|  | ST1 - H2 | -1517.71 | -3791.08 | -1528.08 | -2.59626 | -48.2496 | 18.19501 |
|  | ST1 - H3 | 1498.112 | 1511.087 | 3722.614 | -3.01758 | -17.4626 | 34.44715 |
|  | ST1 - V1 | 49.55296 | -113.163 | 84.20534 | 2892.656 | -508.902 | -437.901 |
|  | ST1 - V2 | -61.3174 | -23.8398 | 133.4533 | 487.9887 | -2926.83 | 477.7768 |
|  | ST1 - V3 | -118.392 | 37.1654 | 15.01116 | -461.584 | -474.996 | 2879.417 |

Table 5 - Static Tests – Local to Local - Stage 1

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2168.75 | 342.1649 | 310.843 | 41.63178 | 9.7978 | 11.4344 |
|  | ST2 - H2 | 360.5707 | 2134.287 | 333.9448 | 7.63996 | -39.5812 | 59.91491 |
|  | ST2 - H3 | 326.1505 | 321.9442 | 2142.294 | 69.3909 | -8.9448 | 5.23701 |
|  | ST2 - V1 | 113.7671 | 69.66098 | -93.753 | 2515.065 | 339.811 | 70.92607 |
|  | ST2 - V2 | -153.323 | 102.8306 | 87.9056 | 71.23128 | 2584.927 | 348.6969 |
|  | ST2 - V3 | 120.6902 | -170.311 | 50.2812 | 355.5377 | 38.24 | 2472.573 |

Table 6 - Static Tests – Local to Local - Stage 2

It is apparent from this test that St1 corner 2 actuator cables are plugged in the opposite direction. This issue was corrected before closing the dome.

Test result:                                              Passed:                 Failed:         .   W Waived:   _X_

### Offset Cartesian drive

The test was not performed because it only tests the matrices, which can be changed at any time.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

### Range of motion

The range of motion of the table is measured by pushing on the table in a direction collinear to the CPS. The Static tests results can be found on the SVN at:

seismic/BSC-ISI/L1/ITMY/Data/Static_Tests/

- L1_ISI_ITMY_Range_Of_Motion_20121205.mat

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -15550.6 | -863 | 15413.05 | 30963.67 | 36.86152 |
| ST1 - H2 | 15044.14 | -1227 | -16048.5 | -31092.7 | -37.0151 |
| ST1 - H3 | -15765.2 | -960 | 15052.46 | 30817.62 | 36.68765 |
| ST1 - V1 | -11339.2 | 1083 | 13504.74 | 24843.9 | 29.57607 |
| ST1 - V2 | 11640.03 | -978 | -13519.4 | -25159.4 | -29.9517 |
| ST1 - V3 | -12600.7 | -112 | 12221.77 | 24822.44 | 29.55052 |
| ST2 - H1 | -10147.6 | -862 | 8432.486 | 18580.04 | 5.529774 |
| ST2 - H2 | -8666.24 | 623 | 9840.774 | 18507.01 | 5.508039 |
| ST2 - H3 | -5994.58 | 3261 | 12492.22 | 18486.79 | 5.502022 |
| ST2 - V1 | -9962 | 999 | 11926.23 | 21888.23 | 6.514353 |
| ST2 - V2 | -14100.2 | -2988 | 8075.137 | 22175.34 | 6.599803 |
| ST2 - V3 | -10159.5 | 477 | 10989.4 | 21148.93 | 6.294323 |

Table 7 - Range of motion - Actuator drive in the LVEA

It is apparent from this test that St1 corner 2 actuator cables are plugged in the opposite direction. This issue was corrected before closing the dome.

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Linearity test

The data of the linearity test can be found on the SVN at:

seismic/BSC-ISI/L1/ITMY/Data/Linearity_Test/

- L1_ISI_ITMY_Linearity_test_20121206.mat

The figures of the linearity test can be found on the SVN at:

seismic\BSC-ISI\H2\BS\Data\Figures\Linearity_Test\

- L1_ISI_ITMY_Linearity_test_20121206.fig

Figure 2 - Linearity test –L1 - BS – In LVEA

|  |  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- | --- |
| Stage 1 | ST1 - H1 | 0.54184 | -906.475 | 0.538735 | 0.576286 |
|  | ST1 - H2 | 0.542533 | -1281.66 |  | 0.704927 |
|  | ST1 - H3 | 0.531833 | -1016.26 |  | -1.28121 |
|  | ST1 - V1 | 0.412395 | 1107.798 | 0.413989 | -0.38512 |
|  | ST1 - V2 | 0.417982 | -948.447 |  | 0.964468 |
|  | ST1 - V3 | 0.411591 | -110.754 |  | -0.57935 |
| Stage 2 | ST2 - H1 | 0.3081 | -725.737 | 0.306914 | 0.386534 |
|  | ST2 - H2 | 0.306895 | 775.5541 |  | -0.00608 |
|  | ST2 - H3 | 0.305746 | 3395.5 |  | -0.38045 |
|  | ST2 - V1 | 0.362764 | 952.3927 | 0.361424 | 0.37093 |
|  | ST2 - V2 | 0.371469 | -3034 |  | 2.779329 |
|  | ST2 - V3 | 0.350038 | 404.1597 |  | -3.15026 |

Table 8 - Slope – Offset Linearity test

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

## Transfer functions and Comparison with measurements done in the staging building.

### At the end station

The parameters for the measurements in the LVEA are slightly different from those in the staging building. We chose to have weaker excitation but longer averages in an effort to reduce risk of the attached suspension.

At this point, only the tuned mass dampers on the spring are installed (No vibration absorbers, no Viton under the keel masses, No vibration absorbers on the BS)

Measurements data can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMY/Data/Transfer_Functions/Measurements/Undamped:

- LLO_ISI_BSC1_Data_L2L_10mHz_100mHz_ST1_ST2_20121029-214002.mat

- LLO_ISI_BSC1_Data_L2L_100mHz_700mHz_ST1_ST2_20121029-161711.mat

- LLO_ISI_BSC1_Data_L2L_700mHz_10Hz_ST1_ST2_20121027-000402.mat

- LLO_ISI_BSC1_Data_L2L_10Hz_100Hz_ST1_ST2_20121026-212711.mat

- LLO_ISI_BSC1_Data_L2L_100Hz_500Hz_ST1_ST2_20121026-200720.mat

- LLO_ISI_BSC1_Data_L2L_500Hz_1000Hz_ST1_ST2_20121026-190253.mat

Data after processing can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMY/Data/Transfer_Functions/Simulations/Undamped

- LLO_ISI_BSC1_TF_L2L_Raw_10mHz_1000Hz_2012_10_29.mat

The transfer functions can be found in the SVN at:

seismic/BSC-ISI/L1/ITMY/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_BSC1_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2012_12_16.fig

- LLO_ISI_BSC1_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2012_12_16.fig

- LLO_ISI_BSC1_TF_L2L_Raw_from_ST1_ACT_to_ST1_T240_2012_12_16.fig

- LLO_ISI_BSC1_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2012_12_16.fig

- LLO_ISI_BSC2_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2012_12_16.fig

Note 1:  The transfer functions are measured from the Output filters bank excitation point to the input (IN1) of the input filters bank. The transfer functions presented below are raw transfer functions without any electronic compensation.

Note 2: The L4Cs are out of phase (should be -90 before 1Hz).  A minus sign is added in the calibration filters that convert count to nm/s.

Note 3: On the ST1-ACT-H to ST1-CPS-H transfer functions, we can see the first resonances of the LVEA test stand at 21.6 Hz, 23.8Hz and 31Hz (matches within a few Hz our BSC 2 results).

Note 4: Even without vibration absorbers on the BS structure, the first resonance visible on Stage 2 is at 154 Hz.

Figure 3 - TF ST1 ACT to ST1 CPS

Figure 4 - TF ST1 ACT to ST1 L4C

Figure 5 - TF ST1 ACT to ST1 T240

Figure 6 - TF ST2 ACT to ST2 CPS

Figure 7 - TF ST2 ACT to ST2 GS13

### Comparisons with measurements on test stand-not done yet

The script used to compare transfer function can be found in the SVN at:

seismic/BSC-ISI/L1/ITMY/Scripts/Control_Scripts

- Comparison_TF_C2C_LHO_ISI_BSC2.m

The figure that shows the comparison between the transfer functions of the staging building and the LVEA are located in the SVN at:

seismic/BSC-ISI/L1/ITMY/Data/Figures/Transfer_Functions/Comparisons/L2L/

- LLO_ISI_BSC1_vs_LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_H_to_ST1_CPS_H_20120512_vs_20121107.fig

- LLO_ISI_BSC1_vs_LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_H_to_ST1_L4C_H_20120512_vs_20121107.fig

- LLO_ISI_BSC1_vs_LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_V_to_ST1_CPS_V_20120512_vs_20121107.fig

- LLO_ISI_BSC1_vs_LHO_ISI_BSC8_Comparison_TF_L2L_ST1_ACT_V_to_ST1_L4C_V_20120512_vs_20121107.fig

- LLO_ISI_BSC1_vs_LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_H_to_ST2_CPS_H_20120512_vs_20121107.fig

- LLO_ISI_BSC1_vs_LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_H_to_ST2_GS13_H_20120512_vs_20121107.fig

- LLO_ISI_BSC1_vs_LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_V_to_ST2_CPS_V_20120512_vs_20121107.fig

- LLO_ISI_BSC1_vs_LHO_ISI_BSC8_Comparison_TF_L2L_ST2_ACT_V_to_ST2_GS13_V_20120512_vs_20121107.fig

Main differences are:

- The DC gains (cables resistance is different due to the length difference)

- Resonances of the rigid body modes [1; 10]Hz (different payload) It is especially visible from 1 to 2 Hz on all sensors.

- ST1-CPS resonances different on the test stand and on the HEPI piers

- Resonance at 20.5 Hz on stage 2 sensors (4th vertical mode of the quad? –cf https://lhocds.ligo-wa.caltech.edu/wiki/Resonances) that had been visible during test stand measurement with the quad locked is somewhat visible here, even though not as clearly

- High frequency behavior different due to data rate change between test stand test and cartridge test

Figure 8 - Transfer functions comparison - ST1 ACT to ST1 CPS H

Figure 8 - Transfer functions comparison - ST1 ACT to ST1 CPS V

Figure 10 - Transfer functions comparison - ST1 ACT to ST1 L4C H

Figure 10 - Transfer functions comparison - ST1 ACT to ST1 L4C V

Figure 12 - Transfer functions comparison - ST2 ACT to ST2 CPS H

Figure 12 - Transfer functions comparison - ST2 ACT to ST2 CPS V

Figure 15 - Transfer functions comparison – ST2 ACT to ST2 GS13 H

Figure 15 - Transfer functions comparison – ST2 ACT to ST2 GS13 V

Note that St 2 H3 and V3 transfer function from 11/07/12 were removed from those past 4 plots as we had some cabling issues later resolved.

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

## Conclusion Phase II-b

All results appear satisfying, apart from a few test waived (will be done during phase II-b):

- Static tests in the Cartesian basis (redundant with the static test in the local basis)

- Linearity test: some corners have slightly different slopes but nothing strikingly alarming. We have seen in the past (HAM-ISI phase 2 testing) stronger slopes different, blamed on different cable lengths inside the LVEA.

- St 1 corner 2 actuators are wired in the opposite direction in some of the static tests. This was corrected before closing the dome.

Test result:                                              Passed:   X              Failed:         .   W Waived:   _

Conclusion

The ISI-BSC1 was moved from the Staging building to the LVEA test stand on July 17th 2012.

This document presents series of tests (Phase II) performed on the ISI-BSC1 (ITMY) after the ITMY quad suspension (in final configuration) was mated with the ISI and the ISI rebalanced.

Phase 2-a tests started in on October 24th 2012. Testing was completed on November 7th , the cartridge install tool place November 12th 2012.

Phase 2b testing was conducted between December 5th and December 17th 2012. The dome was approved to be closed and closed on December 17th 2012.