LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1100857-V8 	                 LIGO                   March 19th, 2013

L1 BSC 3 BSC-ISI, Pre-integration Testing report,

Phase II

E1100857-v8

Adrien Le Roux, Michael Puma, Céline Ramet

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

Introduction	3

1.	Phase II-a	4

1.	Hardware changes	4

1.	CPS – E1100369	4

2.	GS13 – E1100740	4

3.	L4C – E1100740	4

4.	T240 – E1100740	4

5.	Cables – E1100822	4

6.	Misc	4

2.	Electronic Inventory	5

3.	Models Modifications	5

4.	Mass distribution	6

1.	Seismic	6

2.	Suspension	6

3.	Misc	6

4.	Total	7

5.	Basic functionalities just after installing the BSC-ISI on the test stand	7

1.	Pressure sensors	7

2.	Spectra	8

3.	Actuators-cables resistance	8

4.	Offsets CPS Unlocked vs locked	9

5.	Offset local drive	9

6.	Offset Cartesian drive	10

7.	Range of motion	11

8.	Linearity test	11

6.	Transfer functions and Comparison with measurements done in the staging building.	13

1.	At the end station	13

2.	Comparisons with measurements in the staging building –not done yet	16

7.	Conclusion Phase II-a	22

Introduction

The BSC-ISI testing is performed in three phases:

1) BSC-ISI, Pre-integration Testing, Phase I (post-assembly, in the staging building)

2) BSC-ISI, Pre-integration Testing, Phase II: Final tests done before insertion in the chamber
3) BSC-ISI, Integration Phase Testing: Procedure and results related to the commissioning in the chamber.

The ISI-BSC3 was moved from the Staging building to the LVEA in December, 2012. Then the ISI was placed on the test stand a few days later.

This document presents results from the series of tests (Phase II) performed on the ISI-BSC3 (ITMX) in the corner station. The tests were done with the ITMX quadruple suspension installed in its final version.

First tests started on January 31th, 2012. The first testing phase (II-a validation before cartridge installation) was completed on February 4th 2012.

All results are posted on the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMX/

The following types of documents can be found in the SVN:

- Data location

- Figures location

# Phase II-a

## Hardware changes

### CPS – E1100369

Stage 1 H3 CPS SN 13466 has been changed with 13531, because it was giving a constant signal even when we were moving the probe.

### GS13 – E1100740

GS13 have not been replaced since phase I testing in the staging building.

### L4C – E1100740

L4C have not been replaced since phase I testing in the staging building.

### T240 – E1100740

T240s have not been replaced since phase I testing in the staging building.

### Cables – E1100822

One cable has been replaced on this BSC since its Phase I Testing:

- D1100152 SN S1107246 extension cable for L4C V1 was giving a wrong pressure readout and has been replaced with SN S1107249. The issue is still there but we know it is not coming from the pod, but from the connection cable-feedthru. This won’t prevent us from testing and this issue can be solved once the BSC-ISI is in chamber.

### Misc

No hardware changes since phase I testing in the staging building.

## Electronic Inventory

This table reports the electronic equipment used in the LVEA.

| Hardware | Ligo reference | S/N |
| --- | --- | --- |
| Interface Chassis Pod 1 | D1002432 | S1201323 |
| Interface Chassis Pod 2 | D1002432 | S1201322 |
| Interface Chassis Pod 3 | D1002432 | S1201327 |
| Anti-aliasing Chassis Pod 1 | D1002693 | S1203102 |
| Anti-aliasing Chassis Pod 2 | D1002693 | S1203101 |
| Anti-aliasing Chassis Pod 3 | D1002693 | S1203100 |
| Binary Input Chassis 1 | D1001726 | S1101299 |
| Binary Input Chassis 2 | D1001726 | S1101298 |
| Binary Output Chassis | D1001728 | S1101328 |
| T240 Interface Pod 1 | D1002694 | S1201892 |
| T240 Interface Pod 2 | D1002694 | S1104429 |
| T240 Interface Pod 3 | D1002694 | S1104428 |
| Anti-image Chassis | D1000305 | S1203269 |
| Coil driver Pod 1 | D0902744 | S1103564 |
| Coil driver Pod 2 | D0902744 | S1103315 |
| Coil driver Pod 3 | D092744 | S1103329 |
| Expansion chassis | L1seibsc3 | S1100946 |

Table 1 - Electronic inventory

## Models Modifications

No model modifications were done between the beginning and the end of Phase 2a testing.

## Mass distribution

This final mass distribution will be presented once all elements will be installed on the ISI (during phase II-b). These elements are the vibration absorbers on stage 1 and the QUAD structure.

### Seismic

Stage 1

| Stage 1 | HighBay (lbs) | LVEA (lbs) | LVEA (kgs) |
| --- | --- | --- | --- |
| Corner 1 | 31.50 | 31.50 | 14.29 |
| Corner 2 | 15 | 15 | 6.80 |
| Corner 3 | 40.50 | 42.50 | 19.28 |
| Total | 87.0 | 89.0 | 40.37 |

Stage 2

The seismic masses on BSC 3 are;

- On the Keel Plate 1020.5 lbs (=462.9kgs)

- On the Optical Table 637.14 lbs (=289kgs)

The total of masses on Stage 2 is 1657.64 lbs (=751.9 kgs).

### Suspension

The quad structure was weighed to be:

|  | Weight (lbs) | Weight (kgs) |
| --- | --- | --- |
| Upper structure | 266 | 120.66 |
| Lower structure | 531 | 240.86 |
| Total | 797 | 361.5128 |

### Misc

- 44 lbs = quad sleeve [Betsy’s alog 3621]

- 2 lbs = quad sleeve wedges [estimate]

- 5 lbs = cabling [estimate]

- 2 lbs cable brackets [estimate]

- 28 lbs = Vibration absorbers, quantity 4 [from E1000337-v2: 12.7 kg = 28 lbs]

- 3 lbs = Ring heater + brackets + cables [estimate]

Total Weight for Suspension: 906 lbs = 412 kg

### Total

Nominal mass hanging on stage 0-1 blades (without stage 2): 912Kg – 2010lb

Nominal mass hanging on stage 1-2 blades: 2830Kg – 6239lb

Nominal payload on stage 1: 109Kg – 240lb

Nominal payload on stage 2: 1185Kg – 2612lb

|  |  |  | Staging Bldg | LVEA |  |
| --- | --- | --- | --- | --- | --- |
|  | Plan | 10/01/2012 | Detail | Overall | Difference |
| Stage 1 (kgs) | 108.86 | 39.46 | 40.37 | 40.37 | 0.91 |
| Stage 2 (kgs) | Masses | 1183.42 | 1161.33 | 751.9 | 1164.29 | 2.96 |
|  | Suspension | N/A | N/A | 361.51 |  |  |
|  | Miscellaneous | N/A | N/A | 50.88 |  |  |
| Total (kgs) | 1292.28 | 1200.79 | 1204.66 | 1204.66 | 3.87 |

Both Stages difference in weight passes requirement (10 kgs) and are very similar to what we had in the Staging Building, our results are very consistent.

Test result:                                              Passed:   X             Failed:         .   W Waived:   _

## Basic functionalities just after installing the BSC-ISI on the test stand

### Pressure sensors

All pressure sensors are working except the L4C Corner 1. After trouble shooting the issue, it turns out that the pressure sensor in Corner 1 L4C is working and that the wrong value is coming from its extension cable D1100152 SN S1107246 that we replaced with SN S1107249. We also changed the fake feedthru part to connect this extension to a real feedthru; the pressure readout was good for a few minutes and went bad again. We’re suspecting a bad connection of this cable extension and the feedthru. This problem is not holding the testing back though, we will change the extension again once BSC 4 is in chamber.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMX/Data/Static_Tests/
L1_ISI_ITMX_Pressure_Sensors_Check_Calibrated_2013_01_31_120446.mat

|  | Pressure (KPa) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-P | 99.5 | 100.1 | 99.5 |
| ST1-L4C-D | -9.6 | -1.2 | -0.2 |
| ST1-GS13-P | 100.0 | 100.1 | 100.0 |
| ST1-GS13-D | -0.7 | 0.6 | -0.3 |
| ST1-T240-P | 154.1 | 153.3 | 153.8 |

Table 2 - Geophones Pressure sensors

Note/comment about this test: All the pressure sensors are good but we can see one value that doesn’t meet the requirements (ST1 L4C D, which is L4C V1) but as it has been said earlier, after trouble shooting the problem, we now know it is not coming from the sensor itself but from the connection to the feed through. This issue will be solved after the install and doesn’t prevent us from testing this Unit.

Test result:                                              Passed:                 Failed:         .   W Waived:   X_

### Spectra

Spectra of the instrument can be found in the SVN at:

seismic/BSC-ISI/L1/ITMX/Data/Spectra/Undamped/

L1_ISI_ITMX_ASD_m_LOC_CPS_T240_L4C_GS13_2020_02_01_3 1:1:.mat

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Figures/Spectra/Undamped/
L1_ISI_ITMX_ASD_m_LOC_CPS_T240_L4C_GS13_2020_02_01_3 1:1:.fig

Figure 1 - Spectra inboard instruments - ISI Unlocked

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Actuators-cables resistance

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMX/Data/Static_Tests/

- L1_ISI_ITMX_Actuators_Resistance_20130201T115630.mat

|  | Stage 1 | Stage 2 |
| --- | --- | --- |
|  | H1 | H2 | H3 | V1 | V2 | V3 | H1 | H2 | H3 | V1 | V2 | V3 |
| Script | 7.9788 | 7.9443 | 7.7395 | 7.7443 | 7.8884 | 7.8606 | 11.672 | 11.561 | 11.877 | 11.481 | 12 | 11.941 |

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offsets CPS Unlocked vs locked

The table is not perfectly balanced but it is considered sufficiently good to perform the series of test before the cartridge installation. A fine balancing will be done during phase II-b.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMX/Data/Static_Tests/

- L1_ISI_ITMX_CPS_Read_Back_ISI_Locked_2013_02_01_120424.mat

- L1_ISI_ITMX_CPS_Read_Back_ISI_Unlocked_2013_01_31_120424.mat

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | 3.66 | 3.89 | 146.74 | 26.35 | -143.08 | -0.17 |
| ST1 - H2 | 300.20 | 5.12 | -1151.25 | 37.70 | 1451.45 | 1.73 |
| ST1 - H3 | 649.07 | 3.61 | -1670.27 | 61.84 | 2319.34 | 2.76 |
| ST1 - V1 | -133.38 | 5.10 | 43.61 | 46.30 | -176.99 | -0.21 |
| ST1 - V2 | -75.85 | 7.75 | -298.18 | 39.55 | 222.33 | 0.26 |
| ST1 - V3 | -151.65 | 5.09 | 65.77 | 39.17 | -217.42 | -0.26 |
| ST2 - H1 | 526.33 | 12.79 | 1785.12 | 65.38 | -1258.79 | -0.38 |
| ST2 - H2 | -680.52 | 7.65 | -700.63 | 42.51 | 20.11 | 0.01 |
| ST2 - H3 | 978.38 | 9.69 | -58.97 | 89.19 | 1037.35 | 0.31 |
| ST2 - V1 | 1244.30 | 8.56 | 2663.75 | 105.37 | -1419.45 | -0.42 |
| ST2 - V2 | -356.76 | 8.82 | 1190.84 | 95.83 | -1547.60 | -0.46 |
| ST2 - V3 | 1574.90 | 8.00 | 2061.60 | 100.89 | -486.70 | -0.15 |

Table 3 - Locked vs Unlocked Position

Note: The two highlighted values don’t meet the requirement by a few hundred counts. We’re not too worried about it, since the final masses (vibration absorbers, all the Optical Table Masses and the Ballast masses) are not on yet and the CPS will have to be re-adjusted after the cartridge Install, so this will be an easy fix and doesn’t prevent us from testing this Unit.

Test result:                                              Passed:                 Failed:         .   W Waived:   X__

### Offset local drive

Note: Due to longer cables, offsets measured by CPS for a 7000 count drive are slightly lower than offsets measured in the staging building.

Results of this test can be found in the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMX/Data/Static_Tests/

- L1_ISI_ITMX_Offset_Local_Drive_20130131.mat

|  | Sensors |
| --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 3418.12 | 1385.70 | 1372.85 | 21.18 | -2.03 | 1.68 |
|  | ST1 - H2 | 1381.29 | 3450.95 | 1381.11 | 6.85 | 12.59 | 4.49 |
|  | ST1 - H3 | 1366.62 | 1379.50 | 3407.40 | -6.16 | 11.54 | 38.85 |
|  | ST1 - V1 | 60.78 | -108.18 | 81.39 | 2717.36 | -447.05 | -433.51 |
|  | ST1 - V2 | 67.97 | 23.86 | -113.63 | -414.86 | 2653.26 | -442.18 |
|  | ST1 - V3 | -122.85 | 80.11 | 54.03 | -434.84 | -417.44 | 2875.05 |

Table 4 - Static Tests – Local to Local - Stage 1

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2069.686 | 333.1783 | 322.7549 | -13.5116 | -1.85548 | 22.3928 |
|  | ST2 - H2 | 307.8634 | 2027.177 | 314.0907 | -6.1246 | -4.9305 | -14.043 |
|  | ST2 - H3 | 298.0452 | 318.4077 | 1995.944 | -16.0376 | -2.83834 | 33.9722 |
|  | ST2 - V1 | 58.575 | 108.586 | -160.55 | 2348.07 | 293.3302 | 29.8138 |
|  | ST2 - V2 | -151.905 | 54.3535 | 101.2956 | 19.5372 | 2318.535 | 343.1502 |
|  | ST2 - V3 | 100.4344 | -145.447 | 64.46388 | 286.7038 | 52.38362 | 2330.52 |

Table 5 - Static Tests – Local to Local - Stage 2

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offset Cartesian drive

The test was not performed because it only tests the matrices, which can be changed at any time.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

### Range of motion

The range of motion of the table is measured by pushing on the table in a direction collinear to the CPS. The Static tests results can be found on the SVN at:

seismic/BSC-ISI/L1/ITMX/Data/Static_Tests/

- L1_ISI_ITMX_Range_Of_Motion_20130131.mat

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -14425.16 | NaN | 14882.31 | 29307.48 | 34.89 |
| ST1 - H2 | -15814.71 | -1040.00 | 13758.24 | 29572.95 | 35.21 |
| ST1 - H3 | -14264.10 | -1638.00 | 12983.27 | 27247.37 | 32.44 |
| ST1 - V1 | -11631.68 | 16.00 | 11681.96 | 23313.64 | 27.75 |
| ST1 - V2 | -11736.11 | -330.00 | 11052.50 | 22788.61 | 27.13 |
| ST1 - V3 | -12329.28 | -22.00 | 12271.50 | 24600.78 | 29.29 |
| ST2 - H1 | -7057.95 | 1802.00 | 10638.40 | 17696.35 | 5.27 |
| ST2 - H2 | -9282.06 | -646.00 | 7984.24 | 17266.30 | 5.14 |
| ST2 - H3 | -8586.21 | -87.00 | 8388.57 | 16974.78 | 5.05 |
| ST2 - V1 | -7555.45 | 2540.00 | 12592.38 | 20147.83 | 6.00 |
| ST2 - V2 | -9040.98 | 963.00 | 10938.88 | 19979.86 | 5.95 |
| ST2 - V3 | -8041.29 | 1867.00 | 11736.88 | 19778.17 | 5.89 |

Table 6 - Range of motion - Actuator drive in the LVEA

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Linearity test

The data of the linearity test can be found on the SVN at:

seismic/BSC-ISI/L1/ITMX/Data/Linearity_Test/

- L1_ISI_ITMX_Linearity_test_20130201.mat

The figures of the linearity test can be found on the SVN at:

seismic\BSC-ISI\H2\BS\Data\Figures\Linearity_Test\

- L1_ISI_ITMX_Linearity_test_20130201.fig

Figure 2 - Linearity test –L1 - BS – In LVEA

|  |  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- | --- |
| Stage 1 | ST1 - H1 | 0.4885 | 182.5163 | 0.4894 | 0.1907 |
|  | ST1 - H2 | 0.4935 | -1113.6000 |  | -0.8309 |
|  | ST1 - H3 | 0.4863 | -1637.0000 |  | 0.6402 |
|  | ST1 - V1 | 0.3879 | 106.6968 | 0.3926 | 1.1971 |
|  | ST1 - V2 | 0.3795 | -202.7173 |  | 3.3367 |
|  | ST1 - V3 | 0.4104 | 98.5823 |  | -4.5339 |
| Stage 2 | ST2 - H1 | 0.2955 | 1779.0000 | 0.2899 | -1.9317 |
|  | ST2 - H2 | 0.2891 | -691.9026 |  | 0.2760 |
|  | ST2 - H3 | 0.2851 | -59.7437 |  | 1.6557 |
|  | ST2 - V1 | 0.3372 | 2653.6000 | 0.3335 | -1.0993 |
|  | ST2 - V2 | 0.3333 | 1185.9000 |  | 0.0700 |
|  | ST2 - V3 | 0.3301 | 2048.9000 |  | 1.0294 |

Table 7 - Slope – Offset Linearity test

Note all the results seem very coherent. On Stage 1 V2 and V3 actuators seem a little off compared to the average slope (we don’t meet the requirement of 2.5%).

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

## Transfer functions and Comparison with measurements done in the staging building.

### At the end station

The parameters for the measurements in the LVEA are slightly different from those in the staging building. We chose to have weaker excitation but longer averages in an effort to reduce risk for the attached suspension.

At this point, only the tuned mass dampers on the spring are installed (No vibration absorbers, no Viton under the keel masses)

Measurements data can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMX/Data/Transfer_Functions/Measurements/Undamped:

- LLO_ISI_BSC3_Data_L2L_10mHz_100mHz_ST1_ST2_20130204-105939.mat

- LLO_ISI_BSC3_Data_L2L_100mHz_700mHz_ST1_ST2_20130201-150630.mat

- LLO_ISI_BSC3_Data_L2L_700mHz_10Hz_ST1_ST2_20130203-003300.mat

- LLO_ISI_BSC3_Data_L2L_10Hz_100Hz_ST1_ST2_20130202-215607.mat

- LLO_ISI_BSC3_Data_L2L_100Hz_500Hz_ST1_ST2_20130202-203613.mat

- LLO_ISI_BSC3_Data_L2L_500Hz_1000Hz_ST1_ST2_20130202-193145.mat

Data after processing can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMX/Data/Transfer_Functions/Simulations/Undamped

- LLO_ISI_BSC3_TF_L2L_Raw_10mHz_1000Hz_2013_02_04.mat

The transfer functions can be found in the SVN at:

seismic/BSC-ISI/L1/ITMX/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_BSC3_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2013_02_04.fig

- LLO_ISI_BSC3_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2013_02_04.fig

- LLO_ISI_BSC3_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2013_02_04.fig

- LLO_ISI_BSC3_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2013_02_04.fig

Note 1:  The transfer functions are measured from the Output filters bank excitation point to the input (IN1) of the input filters bank. The transfer functions presented below are raw transfer functions without any electronic compensation.

Note 2: The L4Cs are out of phase (should be -90 before 1Hz). A minus sign is added in the calibration filters that convert count to nm/s.

Note 3: On the ST1-ACT-H to ST1-CPS-H transfer functions, we can see the first resonances of the LVEA test stand at 21.1 Hz, 23.4Hz and 33.3Hz (matches within a few Hz our BSC 3 results).

Note 4: During the measurements, the Compensation filter for Stage 1 H1 Actuator was not turned on, that explains why its Magnitude is not as high as the others.

Note5: We had to retake both 10mHz to 100mHz and 100mHz to700mHz TF

Figure 3 - TF ST1 ACT to ST1 CPS

Figure 4 - TF ST1 ACT to ST1 L4C

Figure 5 - TF ST1 ACT to ST1 T240

Figure 6 - TF ST2 ACT to ST2 CPS

Figure 7 - TF ST2 ACT to ST2 GS13

### Comparisons with measurements in the staging building

The script used to compare transfer function can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMX/Scripts/Control_Scripts

- Comparison_TF_C2C_LHO_ISI_BSC3.m

The figure that shows the comparison between the transfer functions of the staging building and the LVEA are located in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMY/Data/Figures/Transfer_Functions/Comparisons/L2L

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST1_ACT_H_to_ST1_CPS_H_20130204_vs_20121216.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST1_ACT_V_to_ST1_CPS_V_20130204_vs_20121216.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST1_ACT_H_to_ST1_L4C_H_20130204_vs_20121216.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST1_ACT_V_to_ST1_L4C_V_20130204_vs_20121216.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST2_ACT_H_to_ST2_CPS_H_20130204_vs_20121216.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST2_ACT_V_to_ST2_CPS_V_20130204_vs_20121216.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST2_ACT_H_to_ST2_GS13_H_20130204_vs_20121216.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST2_ACT_V_to_ST2_GS13_V_20130204_vs_20121216.fig

Main differences are:

- The DC gains (cables resistance is different due to the length difference)

- Resonances of the rigid body modes [1; 10]Hz (different payload) It is especially visible from 1 to 2 Hz on all sensors.

- ST1-CPS resonances different in the staging building and EY (Test stand short legs vs long legs).

- Resonance at 20.5 Hz on stage 2 sensors (4th vertical mode of the quad? –cf https://lhocds.ligo-wa.caltech.edu/wiki/Resonances)

- Similar at high frequencies

Figure 8 - Transfer functions comparison - ST1 ACT to ST1 CPS

Note: we can notice that Stage 1 ACT V3 to Stage CPS V3 resonances were at different frequencies on BSC 1 compared to BSC 3.

Figure 9 - Transfer functions comparison - ST1 ACT to ST1 L4C

Figure 10 - Transfer functions comparison - ST2 ACT to ST2 CPS

Note: we can notice that Stage 1 ACT H1 to Stage CPS H1 resonances are at different frequencies on BSC 3 compared to BSC 1.

Figure 11 - Transfer functions comparison – ST2 ACT to ST2 GS13

Test result:                                              Passed:                 Failed:         .   W Waived:   X_

## Conclusion Phase II-a

All results appear satisfying; apart from a few tests waived (will be done during phase II-b):

- Pressure Sensor Test failed. This issue has been found, and we know it is not coming from the L4C V1 pressure sensor itself but from a bad connection at the feedthru or a bad extension cable, so we will solve this issue after the cartridge install.

- Offset CPS Unlocked vs Locked: the difference between the averages Locked-Unlocked on Stage 1 H3 is a little bit too big to meet the requirements but nothing alarming and the final weights are not on so the ISI will have to be rebalanced and the CPS re-centered anyway.

- Static tests in the Cartesian basis (redundant with the static test in the local basis)

- Linearity test: Stage 1 V2 and V3 actuators seem a little off compared to their average slope but nothing preventing us from Install.

- During the Transfer Functions, the Compensation filter for Stage 1 H1 Actuator was not turned on, but we will do another set of Transfer Functions after the Cartridge Install.

Test result:                                              Passed:   X              Failed:         .   W Waived:   _

# Phase II-a

## Hardware changes

### CPS – E1100369

No change.

### GS13 – E1100740

GS13 have not been replaced since phase I testing in the staging building.

### L4C – E1100740

L4C have not been replaced since phase I testing in the staging building.

### T240 – E1100740

T240s have not been replaced since phase I testing in the staging building.

### Cables – E1100822

Cables have not been changed since the testing in the LVEA. The issue we were having during the testing on the Test stand in the LVEA, on the Vertical L4C pressure sensor in corner 1 has disappeared which confirms our hypothesis that it was coming from the connection at the fake feed through. We will keep an eye on it though.

### Misc

No hardware changes since phase I testing in the staging building.

## Electronic Inventory

This table reports the electronic equipment used in the LVEA.

| Hardware | Ligo reference | S/N |
| --- | --- | --- |
| Interface Chassis Pod 1 | D1002432 | S1201323 |
| Interface Chassis Pod 2 | D1002432 | S1201322 |
| Interface Chassis Pod 3 | D1002432 | S1201327 |
| Anti-aliasing Chassis Pod 1 | D1002693 | S1203102 |
| Anti-aliasing Chassis Pod 2 | D1002693 | S1203101 |
| Anti-aliasing Chassis Pod 3 | D1002693 | S1203100 |
| Binary Input Chassis 1 | D1001726 | S1101299 |
| Binary Input Chassis 2 | D1001726 | S1101298 |
| Binary Output Chassis | D1001728 | S1101328 |
| T240 Interface Pod 1 | D1002694 | S1201892 |
| T240 Interface Pod 2 | D1002694 | S1104429 |
| T240 Interface Pod 3 | D1002694 | S1104428 |
| Anti-image Chassis | D1000305 | S1203269 |
| Coil driver Pod 1 | D0902744 | S1103564 |
| Coil driver Pod 2 | D0902744 | S1103315 |
| Coil driver Pod 3 | D092744 | S1103329 |
| Expansion chassis | L1seibsc3 | S1100946 |

Table 8 - Electronic inventory

## Models Modifications

No model modifications were done between the beginning and the end of Phase 2a testing.

## Mass distribution

This final mass distribution will be presented once all elements will be installed on the ISI (during phase II-b). These elements are the vibration absorbers on stage 1 and the QUAD structure.

### Seismic

Stage 1

| Stage 1 | HighBay (lbs) | LVEA (lbs) | LVEA (kgs) | In Chamber (lbs) | In Chamber (kgs) |
| --- | --- | --- | --- | --- | --- |
| Corner 1 | 31.50 | 31.50 | 14.29 | 27.86 | 12.6 |
| Corner 2 | 15 | 15 | 6.80 | 10 | 4.5 |
| Corner 3 | 40.50 | 42.50 | 19.28 | 43.72 | 19.8 |
| Total | 87.0 | 89.0 | 40.37 | 81.58 | 37.0 |

Stage 2

The seismic masses on BSC 3 are;

- On the Keel Plate 650 lbs (=294.8kgs)

- On the Optical Table 714.3 lbs (=324kgs)

- On the Side 335.7 lbs (152.3 kgs)

The total of masses on Stage 2 is 1700.0 lbs (=771.1 kgs).

Note: We had to add 2 D071200 Type 01 and 1 D071200 Type 00 on Corner 1 to re-balance the ISI because after the Install, SUS removed a light part of the Quad, so we had to make up for it.

### Suspension

The quad structure was weighed to be:

|  | Weight (lbs) | Weight (kgs) |
| --- | --- | --- |
| Upper structure | 266 | 120.66 |
| Lower structure | 531 | 240.86 |
| Total | 797 | 361.5128 |

### Misc

- 44 lbs = quad sleeve [Betsy’s alog 3621]

- 2 lbs = quad sleeve wedges [estimate]

- 5 lbs = cabling [estimate]

- 2 lbs cable brackets [estimate]

- 28 lbs = Vibration absorbers, quantity 4 [from E1000337-v2: 12.7 kg = 28 lbs]

- 3 lbs = Ring heater + brackets + cables [estimate]

Total Weight for Suspension: 906 lbs = 412 kg

### Total

Nominal mass hanging on stage 0-1 blades (without stage 2): 912Kg – 2010lb

Nominal mass hanging on stage 1-2 blades: 2830Kg – 6239lb

Nominal payload on stage 1: 109Kg – 240lb

Nominal payload on stage 2: 1185Kg – 2612lb

|  |  |  | Staging Bldg | LVEA | In chamber |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Plan | 10/1/2012 | Detail | Overall | Detail | Overall | Difference |
| Stage 1 (kgs) | 108.86 | 39.46 | 40.37 | 40.37 | 37 | 37 | -2.5 |
| Stage 2 (kgs) | Masses | 1183.42 | 1161.33 | 751.9 | 1164.29 | 771.1 | 1183.49 | 22.2 |
|  | Suspension | N/A | N/A | 361.51 |  | 361.51 |  |  |
|  | Miscellaneous | N/A | N/A | 50.88 |  | 50.88 |  |  |
| Total (kgs) | 1292.28 | 1200.79 | 1204.66 | 1204.66 | 1220.49 | 1220.49 | 19.7 |

Even if Stage 2 doesn’t meet the requirements in weight (10 kgs) we are very similar to what we had in the Staging Building overall, our results are very consistent (within 20kgs since the Staging building).

Test result:                                              Passed:                Failed:         .   W Waived:   X_

## Basic functionalities just after installing the BSC-ISI on the test stand

### Pressure sensors

All pressure sensors are working.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMX/Data/Static_Tests/
L1_ISI_ITMX_Pressure_Sensors_Check_Calibrated_2020_03_11_3 3:0:.mat

|  | Pressure (KPa) |
| --- | --- |
| Sensors | Corner 1 | Corner 2 | Corner 3 |
| ST1-L4C-P | 99.5 | 100.2 | 99.5 |
| ST1-L4C-D | 0.5 | -1.2 | -0.2 |
| ST1-GS13-P | 100.2 | 100.2 | 100.1 |
| ST1-GS13-D | -0.7 | 0.5 | -0.3 |
| ST1-T240-P | 154.1 | 153.3 | 153.8 |

Table 9 - Geophones Pressure sensors

Note/comment about this test: The issue we were having during the testing on the Test stand in the LVEA, on the Vertical L4C pressure sensor in corner 1 has disappeared which confirms our hypothesis that it was coming from the connection at the fake feed through. We will keep an eye on it though.

Test result:                                              Passed:   X             Failed:         .   W Waived:    _

### Spectra

Spectra of the instrument can be found in the SVN at:

seismic/BSC-ISI/L1/ITMX/Data/Spectra/Undamped/

L1_ISI_ITMX_ASD_m_LOC_CPS_T240_L4C_GS13_2020_02_14_3 3:9:.mat

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMY/Data/Figures/Spectra/Undamped/
L1_ISI_ITMX_ASD_m_LOC_CPS_T240_L4C_GS13_2020_02_14_3 3:9:.fig

Figure 12 - Spectra inboard instruments - ISI Unlocked

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Actuators-cables resistance

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMX/Data/Static_Tests/

- L1_ISI_ITMX_Actuators_Resistance_20130214T151112.mat

|  | Stage 1 | Stage 2 |
| --- | --- | --- |
|  | H1 | H2 | H3 | V1 | V2 | V3 | H1 | H2 | H3 | V1 | V2 | V3 |
| Script | 7.9788 | 7.9443 | 7.7395 | 7.7443 | 7.8884 | 7.8606 | 11.6722 | 11.5613 | 11.8766 | 11.4809 | 12 | 11.9405 |

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offsets CPS Unlocked vs locked

The table is not perfectly balanced but it is considered sufficiently good to perform the series of test before the cartridge installation. A fine balancing will be done during phase II-b.

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMX/Data/Static_Tests/

- L1_ISI_ITMX_CPS_Read_Back_ISI_UnLocked_2020_03_08_3 1:0:.mat

- L1_ISI_ITMX_CPS_Read_Back_ISI_UnLocked_2020_03_08_3 0:4:.mat

|  | Table locked | Table unlocked | Difference locked - unlocked |
| --- | --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | mil |
| ST1 - H1 | 140.88 | 20.99 | 326.45 | 29.79 | 185.57 | 0.22 |
| ST1 - H2 | 50.88 | 14.90 | 486.32 | 36.96 | 435.44 | 0.52 |
| ST1 - H3 | -278.31 | 21.85 | 56.87 | 34.71 | 335.18 | 0.40 |
| ST1 - V1 | 52.36 | 140.80 | -680.13 | 54.88 | -732.48 | -0.87 |
| ST1 - V2 | 134.45 | 47.42 | 811.29 | 47.08 | 676.84 | 0.81 |
| ST1 - V3 | -21.40 | 75.70 | 1093.31 | 48.72 | 1114.71 | 1.33 |
| ST2 - H1 | -493.17 | 18.99 | 2101.34 | 51.40 | 2594.51 | 0.77 |
| ST2 - H2 | -194.96 | 54.73 | 2339.62 | 62.17 | 2534.58 | 0.75 |
| ST2 - H3 | 60.12 | 102.21 | -357.82 | 50.16 | -417.95 | -0.12 |
| ST2 - V1 | 227.21 | 48.02 | 2485.00 | 155.09 | 2257.79 | 0.67 |
| ST2 - V2 | 9.09 | 228.98 | -280.23 | 143.65 | -289.32 | -0.09 |
| ST2 - V3 | 631.75 | 350.19 | 1472.40 | 145.34 | 840.66 | 0.25 |

Table 10 - Locked vs Unlocked Position

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offset local drive

Note: Due to longer cables, offsets measured by CPS for a 7000 count drive are slightly lower than offsets measured in the staging building.

Results of this test can be found in the SVN at:

https://svn.ligo.caltech.edu/svn/seismic/BSC-ISI/L1/ITMX/Data/Static_Tests/

- L1_ISI_ITMX_Offset_Local_Drive_20130214.mat

|  | Sensors |
| --- | --- |
|  |  | ST1 - H1 | ST1 - H2 | ST1 - H3 | ST1 - V1 | ST1 - V2 | ST1 - V3 |
| Actuators | ST1 - H1 | 3754.76 | 1509.15 | 1503.04 | -4.53 | 1.66 | -0.25 |
|  | ST1 - H2 | 1518.88 | 3820.09 | 1510.33 | -0.46 | -3.75 | 5.47 |
|  | ST1 - H3 | 1506.59 | 1514.34 | 3759.80 | -10.28 | -0.42 | 29.26 |
|  | ST1 - V1 | 51.97 | -124.37 | 77.66 | 2912.47 | -494.18 | -470.74 |
|  | ST1 - V2 | 86.60 | 23.43 | -142.77 | -466.15 | 2937.36 | -488.69 |
|  | ST1 - V3 | -125.67 | 78.28 | 49.22 | -485.40 | -475.54 | 3169.11 |

Table 11 - Static Tests – Local to Local - Stage 1

|  |  | Sensors |
| --- | --- | --- |
|  |  | ST2 - H1 | ST2 - H2 | ST2 - H3 | ST2 - V1 | ST2 - V2 | ST2 - V3 |
| Actuators | ST2 - H1 | 2183.13 | 339.84 | 319.90 | -22.40 | 20.98 | -5.96 |
|  | ST2 - H2 | 320.63 | 2160.34 | 346.04 | 23.05 | -25.05 | -118.62 |
|  | ST2 - H3 | 319.41 | 341.88 | 2108.08 | -24.62 | 27.88 | 31.68 |
|  | ST2 - V1 | 73.89 | 114.21 | -179.55 | 2452.69 | 330.60 | 36.92 |
|  | ST2 - V2 | -171.65 | 50.12 | 116.75 | 51.13 | 2460.69 | 201.05 |
|  | ST2 - V3 | 114.48 | -193.83 | 108.23 | 380.07 | -7.44 | 2232.55 |

Table 12 - Static Tests – Local to Local - Stage 2

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Offset Cartesian drive

The test was not performed because it only tests the matrices, which can be changed at any time.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

### Range of motion

The range of motion of the table is measured by pushing on the table in a direction collinear to the CPS. The Static tests results can be found on the SVN at:

seismic/BSC-ISI/L1/ITMX/Data/Static_Tests/

- L1_ISI_ITMX_Range_Of_Motion_20130214.mat

| Sensor readout (counts) | Negative drive | no drive | Positive drive | Amplitude count | mil |
| --- | --- | --- | --- | --- | --- |
| ST1 - H1 | -15677.43 | 354.00 | 16027.83 | 31705.26 | 37.74 |
| ST1 - H2 | -15607.16 | 731.00 | 16334.96 | 31942.12 | 38.03 |
| ST1 - H3 | -15649.00 | 427.00 | 15950.87 | 31599.87 | 37.62 |
| ST1 - V1 | -13891.84 | -1399.00 | 11089.97 | 24981.82 | 29.74 |
| ST1 - V2 | -11988.38 | 624.00 | 13210.79 | 25199.18 | 30.00 |
| ST1 - V3 | -12912.43 | 679.00 | 14205.68 | 27118.11 | 32.28 |
| ST2 - H1 | -7245.15 | 2097.00 | 11419.01 | 18664.16 | 5.55 |
| ST2 - H2 | -6685.11 | 2597.00 | 11831.86 | 18516.97 | 5.51 |
| ST2 - H3 | -9165.07 | -194.00 | 8836.26 | 18001.33 | 5.36 |
| ST2 - V1 | -9240.17 | 1261.00 | 11808.06 | 21048.23 | 6.26 |
| ST2 - V2 | -12981.14 | -2308.00 | 8315.57 | 21296.71 | 6.34 |
| ST2 - V3 | -11624.82 | -1024.00 | 9151.12 | 20775.94 | 6.18 |

Table 13 - Range of motion - Actuator drive in the LVEA

Test result:                                              Passed:   X              Failed:         .   W Waived:   __

### Linearity test

The data of the linearity test can be found on the SVN at:

seismic/BSC-ISI/L1/ITMX/Data/Linearity_Test/

- L1_ISI_ITMX_Linearity_test_20130215.mat

The figures of the linearity test can be found on the SVN at:

seismic\BSC-ISI\L1\ITMX\Data\Figures\Linearity_Test\

- L1_ISI_ITMX_Linearity_test_20130215.fig

Figure 13 - Linearity test –L1 - BS – In LVEA

|  |  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- | --- |
| Stage 1 | ST1 - H1 | 0.5396 | 385.6500 | 0.5405 | -0.1616 |
|  | ST1 - H2 | 0.5442 | 679.7500 |  | 0.6895 |
|  | ST1 - H3 | 0.5376 | 383.7700 |  | -0.5279 |
|  | ST1 - V1 | 0.4166 | -1277.3000 | 0.4293 | -2.9684 |
|  | ST1 - V2 | 0.4196 | 736.8200 |  | -2.2487 |
|  | ST1 - V3 | 0.4517 | 883.9200 |  | 5.2171 |
| Stage 2 | ST2 - H1 | 0.3122 | 2147.8000 | 0.3116 | 0.1786 |
|  | ST2 - H2 | 0.3093 | 2611.3000 |  | -0.7423 |
|  | ST2 - H3 | 0.3134 | -256.5200 |  | 0.5637 |
|  | ST2 - V1 | 0.3520 | 1469.3000 | 0.3508 | 0.3678 |
|  | ST2 - V2 | 0.3563 | -2144.8000 |  | 1.5709 |
|  | ST2 - V3 | 0.3440 | -954.9900 |  | -1.9387 |

Table 14 - Slope – Offset Linearity test

Note all the results seem very coherent. On Stage 1 V1, V2 and V3 actuators seem a little off compared to the average slope (we don’t meet the requirement of 2.5%). We had similar results on the test stand outside the chamber.

On Stage 2 CPS we can see that the data are noisy, it’s probably due to people working in the LVEA. If we have more time we’ll try to retake this test when the LVEA is quieter.

Test result:                                              Passed:   _              Failed:         .   W Waived:   X_

## Transfer functions and Comparison with measurements done in the staging building.

### At the end station

The parameters for the measurements in the LVEA are slightly different from those in the staging building. We chose to have weaker excitation but longer averages in an effort to reduce risk for the attached suspension.

At this point, only the tuned mass dampers on the spring are installed (No vibration absorbers, no Viton under the keel masses)

Measurements data can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMX/Data/Transfer_Functions/Measurements/Undamped:

- LLO_ISI_BSC3_Data_L2L_10mHz_100mHz_ST1_ST2_20130307-024232.mat

- LLO_ISI_BSC3_Data_L2L_100mHz_700mHz_ST1_ST2_20130307-003939.mat

- LLO_ISI_BSC3_Data_L2L_700mHz_10Hz_ST1_ST2_20130306-142057.mat

- LLO_ISI_BSC3_Data_L2L_10Hz_100Hz_ST1_ST2_20130305-183045.mat

- LLO_ISI_BSC3_Data_L2L_100Hz_500Hz_ST1_ST2_20130305-171054.mat

- LLO_ISI_BSC3_Data_L2L_500Hz_1000Hz_ST1_ST2_20130305-160627.mat

Data after processing can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMX/Data/Transfer_Functions/Simulations/Undamped

- LLO_ISI_BSC3_TF_L2L_Raw_10mHz_1000Hz_2013_03_07.mat

The transfer functions can be found in the SVN at:

seismic/BSC-ISI/L1/ITMX/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_BSC3_TF_L2L_Raw_from_ST1_ACT_to_ST1_CPS_2013_03_07.fig

- LLO_ISI_BSC3_TF_L2L_Raw_from_ST1_ACT_to_ST1_L4C_2013_03_07.fig

- LLO_ISI_BSC3_TF_L2L_Raw_from_ST2_ACT_to_ST2_CPS_2013_03_07.fig

- LLO_ISI_BSC3_TF_L2L_Raw_from_ST2_ACT_to_ST2_GS13_2013_03_07.fig

Note 1:  The transfer functions are measured from the Output filters bank excitation point to the input (IN1) of the input filters bank. The transfer functions presented below are raw transfer functions without any electronic compensation.

Note 2: The L4Cs are out of phase (should be -90 before 1Hz). A minus sign is added in the calibration filters that convert count to nm/s.

Note 3: On the ST1-ACT-H to ST1-CPS-H transfer functions, we can see the first resonances of the LVEA test stand at 21.1 Hz, 23.4Hz and 33.3Hz (matches within a few Hz our BSC 3 results).

Note 4: During the measurements, the Compensation filter for Stage 1 H1 Actuator was not turned on, that explains why its Magnitude is not as high as the others.

Note5: We had to retake both 10mHz to 100mHz and 100mHz to700mHz TF

Figure 14 - TF ST1 ACT to ST1 CPS

Figure 15 - TF ST1 ACT to ST1 L4C

Figure 16 - TF ST1 ACT to ST1 T240

Figure 17 - TF ST2 ACT to ST2 CPS

Figure 18 - TF ST2 ACT to ST2 GS13

### Comparisons with measurements in the staging building

The script used to compare transfer function can be found in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMX/Scripts/Control_Scripts

- Comparison_TF_C2C_LHO_ISI_BSC3.m

The figure that shows the comparison between the transfer functions of BSC 1 ITMY and BSC 3 ITMX are located in the SVN at:

SeiSVN/seismic/BSC-ISI/L1/ITMY/Data/Figures/Transfer_Functions/Comparisons/L2L

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST1_ACT_H_to_ST1_CPS_H_20121011_vs_20130307.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST1_ACT_V_to_ST1_CPS_V_20121011_vs_20130307.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST1_ACT_H_to_ST1_L4C_H_20121011_vs_20130307.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST1_ACT_V_to_ST1_L4C_V_20121011_vs_20130307.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST2_ACT_H_to_ST2_CPS_H_20121011_vs_20130307.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST2_ACT_V_to_ST2_CPS_V_20121011_vs_20130307.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST2_ACT_H_to_ST2_GS13_H_20121011_vs_20130307.fig

- LLO_ISI_BSC3_Chamber_vs_LLO_ISI_BSC1_Comparison_TF_L2L_ST2_ACT_V_to_ST2_GS13_V_20121011_vs_20130307.fig

Main differences are:

- The DC gains (cables resistance is different due to the length difference)

- Resonances of the rigid body modes [1; 10]Hz (different payload) It is especially visible from 1 to 2 Hz on all sensors.

- Resonance at 20.5 Hz on stage 2 sensors (4th vertical mode of the quad? –cf https://lhocds.ligo-wa.caltech.edu/wiki/Resonances)

- Similar at high frequencies

Figure 19 - Transfer functions comparison - ST1 ACT to ST1 CPS

Note: we can notice that Stage 1 ACT V3 to Stage CPS V3 resonances were at different frequencies on BSC 1 compared to BSC 3. After further investigation, we found out that this zero at 17 Hz was caused by an untorqued Bracket on Stage 0-1 V1 Actuator:

Figure 20 - Transfer functions comparison - ST1 ACT to ST1 L4C

Figure 21 - Transfer functions comparison - ST2 ACT to ST2 CPS

Note: we can notice that Stage 1 ACT H1 to Stage CPS H1 resonances are at different frequencies on BSC 3 compared to BSC 1.

Figure 22 - Transfer functions comparison – ST2 ACT to ST2 GS13

Test result:                                              Passed:                 Failed:         .   W Waived:   X_

After torqueing the Stage 0-1 V1 Actuator Bracket, we re took the Transfer Function to make sure, we solved our issue and here are the results we had:

Figure 23: Transfer functions - ST1 ACT to ST1 CPS

Figure 24: Comparison of the Transfer functions - ST1 ACT to ST1 CPS between BSC 3 with the Stage 0-1 V1 Actuator Bracket Untorqued (03/07/2013) and Torqued (03/15/2013)

On these figures, we can clearly see that torqueing this bracket solved our issue.

Looking at Figure 21, we can also conclude that Stage 1-2 H1 CPS on LLO BSC1 has the same issue and that next time the dome is opened, we’ll have to go back and make sure everything is torqued in the neighborhood of Stage 1-2 H1 CPS.

## Conclusion Phase II-a

All results appear satisfying; apart from a few tests

- Linearity test: Stage 1 V1, V2 and V3 actuators seem a little off compared to their average slope but nothing too alarming and looking at the other tests, everything seems normal.

- The comparisons of the Transfer Functions between BSC 3 and BSC 1 that has already been approved show us that this BSC 3 is good to go.

- Mass Budget: we saw a 16 kgs difference between the mass budget on the test stand before install and the mass budget in chamber after install (+19 kgs on Stage 2 and –3 kgs on Stage 1). Without being alarming, this change was worth being written down in the Testing Report. Because of that we had to add another D1003136 Ballast Mass on top of the Keel Plate in order to be able to balance the ISI. We decided to start a new stack located as shown below:

Test result:                                              Passed:   X              Failed:         .   W Waived:   _