## proof tests
Viewport materials and dimensions
Revised 26 Dec 2011
[11 empty] | Weibull Parameters | [11 empty] | 2 | reverse pressure (psig)
Type | Supplier | Model Number | Sub-Sys | Flange OD (in) | Optic material | View Dia (in) | window Radius (in) | Optic thick. (in) | E (GPa) | n (--) | S0 (Mpa) | m | F | Inert Strength Si (Mpa) | Limit Stress sf (Mpa) | peak bending stress (MPa) | FEA stress (Mpa) | Factor of Safety | N (--) | B (Mpa^2-s) | proof pressure (atm) | GN2 pressure (psig) | bending stress (Mpa) | Factor of Safety
AOS Commercial 6" | Norcal | ZV800 | various | 9.97 | 7056 glass | 5.6 | 3.05 | 0.375 | 62.8 | 0.21 | [5 empty] | 8.0707 | 7.9 | 2 |  |  | 2 | 14.7 | 1.0981 | 14.7 | FEA slide 18 of file "2008 07 18 SSV v1.6.ppt" in G080409-v1
AOS Commercial 7.8" | MDC | 9722012/450027 | various | 9.97 | Quartz/ Fused Silica | 7.78 | 4.14 | 0.375 | 73.8 | 0.17 | 156.6 | 4.4 | 0.01 | 55.048 | 22.7094 | 14.6847 |  | 1.5465 | 40.5 | 0.00051 | 2.3697 | 20.1352 | 1.9979 | 11.3665
AOS Commercial 4.5" | MDC | 450004 | Oplev | 4.47 | 7056 glass | 2.69 | 1.595 | 0.17 | 62.8 | 0.21 | [5 empty] | 10.7398 |  | 2 |  |  | 2 | 14.7 | 1.4612 | 14.7
AOS High Quality 6" | LIGO | D1100999 | ISC, IO | 9.97 | Fused Silica | 5.24 | 2.669 | 0.75 | 73.8 | 0.17 | 156.6 | 4.4 | 1e-05 | 11.4401 | 5.1002 | 1.5258 | 1.4686 | 3.4729 | 40.5 | 0.00051 | 2.1026 | 16.2081 | 0.20759 | 24.5681
AOS High Quality 6", wedged | LIGO | D1101000 | ISC, IO | 9.97 | Fused Silica | 5.24 | 2.669 | 0.87 | 73.8 | 0.17 | 156.6 | 4.4 | 1e-05 | 11.4401 | 5.1002 | 1.1339 |  | 4.4978 | 40.5 | 0.00051 | 2.0745 | 15.7956 | 0.15428 | 33.0589
AOS High Quality 6", septum | LIGO | D1101535 | ISC, IO | 9.97 | Fused Silica | 5.24 | 2.669 | 0.75 | 73.8 | 0.17 | 156.6 | 4.4 | 1e-05 | 11.4401 | 5.1002 | 1.5258 | 1.4686 | 3.4729 | 40.5 | 0.00051 | 2.1026 | 16.2081 | 0.20759 | 24.5681
AOS High Quality 6", septum, wedged | LIGO | D1101092 | ISC, IO | 9.97 | Fused Silica | 5.24 | 2.669 | 0.87 | 73.8 | 0.17 | 156.6 | 4.4 | 1e-05 | 11.4401 | 5.1002 | 1.1339 |  | 4.4978 | 40.5 | 0.00051 | 2.0745 | 15.7956 | 0.15428 | 33.0589
TCS Dual 3" | LIGO | D1003194 | TCS | 9.97 | Fused Silica | 2.25 | 1.3055 | 0.5 | 73.8 | 0.17 | 156.6 | 4.4 | 1e-05 | 11.4401 | 5.1002 | 0.82137 |  | 6.2093 | 40.5 | 0.00051 | 2.0401 | 15.2891 | 0.11175 | 45.6387
[5 empty] | ZnSe | 2.25 | 1.3055 | 0.5 | 74.3 | 0.31 | 60.6 | 9.6 | 1e-05 | 18.2658 | 7.8636 | 0.85765 | 0.761 | 10.3333 | 39.6 | 0.00067426 | 2.0515 | 15.4568 | 0.11669 | 67.3907
PSL High Power, 6" (non-wedged) | LIGO | D1101670 | IO | 9.97 | Fused Silica | 5.24 | 2.669 | 0.75 | 73.8 | 0.17 | 156.6 | 4.4 | 1e-05 | 11.4401 | 5.1002 | 1.5258 | 1.4686 | 3.4729 | 40.5 | 0.00051 | 2.1026 | 16.2081 | 0.20759 | 24.5681
PSL High Power, 6", wedged | LIGO | D1101714 | IO | 9.97 | Fused Silica | 5.24 | 2.669 | 0.87 | 73.8 | 0.17 | 156.6 | 4.4 | 1e-05 | 11.4401 | 5.1002 | 1.1339 |  | 4.4978 | 40.5 | 0.00051 | 2.0745 | 15.7956 | 0.15428 | 33.0589
Notes:
1) The lifetime assumed for these proof tests is 20 years.
2) The TCS Dual 3" viewport proof pressure should be the highest of the individual proof pressures of each window. This will ensure a minimum of 20 years of life in sevice.
3) The outer radius of the glass window is assumed to be .25" larger than the clear aperture, based on measurements for one design.
[5 empty] | SiO2 |  | ZnSe
[5 empty] | 4.4 | m | 9.6 | --
[5 empty] | 156.6 | S0 | 60.6 | MPa
[5 empty] | 1e-05 | F | 1e-05 | --
[5 empty] | 11.4401 | Si | 18.2658 | MPa
[5 empty] | 40.5 | N | 39.6 | --
[5 empty] | 630720000 | tf | 630720000 | sec
[5 empty] | 0.00051 | B | 0.00067426 | MPa^2 s
[5 empty] | 5.1002 | sf | 7.8636 | MPa
[5 empty] | 7485964.4592 | A | 1090 | (MPa m^.5)^-N
[5 empty] | 0.722 | Kic | 0.9 | MPa m^.5
| 657 | P | lb |  | 1.95 | Y | 1.95 | --
| 74 | l | in
| 10000000 | E | psi
| 5.2 | I | in^4
| 0.10666 | y | in
| 0.066665 | y | in
0.0625
0.0047244

## circular plate
Flat Circular Plate of constant thickness, with uniform pressure & simple support
Roark's Formulas for Stress and Strain, 6th ed., Table 24, case 10a with ro=0
value | units | value | units | symbol | description
2.8 | in |  |  | a | radius to center of simple support (either o-ring or flex support)
14.7 | psi |  |  | q | applied pressure load
0.375 | in |  |  | t | thickness
9110000 | psi | 62.8112 | GPa | E | modulus of elasticity | 9108369.9294
0.21 | - |  |  | ν | Poisson's ratio
41881.1379 | in-lb |  |  | D | plate bending stiffness
-0.0014515 | in |  |  | yc | center deflection
986.5229 | psi |  |  | sc | surface stress at center
Flat Circular Plate of constant thickness, with a uniform annular line load and simple support
Roark's Formulas for Stress and Strain, 6th ed., Table 24, case 9a
value | units | value | units | symbol | description | Notes
2.855 | in |  |  | a | radius to center of outer support (clamping ring)
2.7855 | in |  |  | r0 | radius to center of inner support (o-ring)
0.75 | in |  |  | t | thickness
287 | lb/in |  |  | w | unit line load (force per unit of circumferential length) |  | 5023.0202
10700000 | psi | 73.7739 |  | E | modulus of elasticity | Fused Silica
0.17 |  |  |  | ν | Poisson's ratio | Fused Silica
387366.7748 |  |  |  | D | plate bending stiffness
2.3748e-06 |  |  |  | L3 | loading constant
0.023803 |  |  |  | L9 | loading constant
-0.00017534 |  |  |  | yc | deflection at the center of the plate | compare to FEA value of -0.00010 in
19.5035 |  |  |  | Mc | moment at the center of the plate
208.0373 |  |  |  | σc | tensile stress at center | compare to FEA value of 180 psi
8702.2644 | psi | 60000000 | Pa |  | static fatigue stress limit?
4 |  |  |  | FS
2175.5661 |  |  |  |  | allowable stress

## lifetime fused silica
3153600000 | sec | 100 | yrs | tf | lifetime
0.72 | MPa m^.5 |  |  | Kic | critical stress intensity
38 | -- |  |  | N | crack growth parameter
31.1
1.7725 | -- |  |  | Y | crack geometry parameter
35000000 | m/s |  |  | A | crack growth parameter
6.911e-05 | MPa^2 sec |  |  | B | crack growth parameter, from crack propagation
8.6e-06 | MPa^2 sec |  |  | B | from static fatigue data
151 | MPa |  |  | Si | inert strength
157
50.6803 | MPa | 7350.5626 | psi | sa | applied stress, from crack propagation
38.5814 |  | 5595.7659 |  |  | applied stress, from static fatugue data
[7 empty] | 553.0805 |  |  |  |  | 367.5537
0.0001 |  |  |  | F |  |  | Si (psi) |  |  |  |  | Sa (psi)
4.4 |  |  |  | m |  |  | 1659.2416 |  |  |  |  | 1102.661
156.6 |  |  |  | So
19.3066 |  | 2800.1855 |  | Si |  | 5 | number of pts per decade
0.00051 |  |  |  | B |  | 1.5849 | factor
40.5 |  |  |  | N |  |  | tf (yrs) | [6 empty] | 1 | 1 | 20 | 20 | 50 | 50
630720000 | sec | 20 |  | tf |  |  | tf (sec) | 0.001 | 0.001 | 1 | 1 | 60 | 60 | 31536000 | 31536000 | 630720000 | 630720000 | 1576800000 | 1576800000
8.3876 | MPa | 1216.524 | psi | Sa |  |  | N | 40.5 | 31.1 | 40.5 | 31.1 | 40.5 | 31.1 | 40.5 | 31.1 | 40.5 | 31.1 | 40.5 | 31.1
[7 empty] | B | 0.00051 | 8.6e-06 | 0.00051 | 8.6e-06 | 0.00051 | 8.6e-06 | 0.00051 | 8.6e-06 | 0.00051 | 8.6e-06 | 0.00051 | 8.6e-06
5.5158 |  | 800 |  |  |  | F | Si (MPa) | Sa (MPa) | Sa (MPa) | Sa (MPa) | Sa (MPa) | Sa (MPa) | Sa (MPa) | Sa (MPa) | Sa (MPa) | Sa (MPa) | Sa (MPa) | Sa (MPa) | Sa (MPa)
[6 empty] | 1e-06 | 6.7788 | 6.0658 | 5.1439 | 5.1147 | 4.1193 | 4.6229 | 3.6112 | 3.3393 | 2.3643 | 3.1012 | 2.1472 | 3.0319 | 2.0849
|  | 220 |  |  |  | 1.5849e-06 | 7.5268 | 6.7004 | 5.6731 | 5.6497 | 4.5432 | 5.1065 | 3.9828 | 3.6887 | 2.6076 | 3.4257 | 2.3681 | 3.349 | 2.2994
|  | 660 |  |  |  | 2.5119e-06 | 8.3573 | 7.4013 | 6.2568 | 6.2407 | 5.0106 | 5.6407 | 4.3925 | 4.0746 | 2.8759 | 3.784 | 2.6118 | 3.6994 | 2.536
|  | 4.5505 |  |  |  | 3.9811e-06 | 9.2794 | 8.1756 | 6.9006 | 6.8936 | 5.5261 | 6.2308 | 4.8445 | 4.5008 | 3.1718 | 4.1799 | 2.8805 | 4.0864 | 2.7969
[6 empty] | 6.3096e-06 | 10.3032 | 9.0309 | 7.6105 | 7.6148 | 6.0947 | 6.8826 | 5.3429 | 4.9716 | 3.4981 | 4.6172 | 3.1769 | 4.5139 | 3.0846
[6 empty] | 1e-05 | 11.4401 | 9.9756 | 8.3936 | 8.4114 | 6.7218 | 7.6026 | 5.8926 | 5.4917 | 3.858 | 5.1002 | 3.5037 | 4.9861 | 3.402
12.9621 | MPa | 1880 | psi | Sa at 1 atm |  | 1.5849e-05 | 12.7023 | 11.0192 | 9.2572 | 9.2913 | 7.4134 | 8.3979 | 6.4989 | 6.0663 | 4.255 | 5.6337 | 3.8642 | 5.5077 | 3.752
2.36 | atm |  |  | SP, proof load |  | 2.5119e-05 | 14.1039 | 12.172 | 10.2097 | 10.2633 | 8.1761 | 9.2764 | 7.1676 | 6.7009 | 4.6928 | 6.2231 | 4.2618 | 6.0839 | 4.1381
30.5907 | MPa | 4436.8 | psi | Sp, proof stress |  | 3.9811e-05 | 15.6601 | 13.4453 | 11.2601 | 11.337 | 9.0174 | 10.2469 | 7.9051 | 7.4019 | 5.1756 | 6.8741 | 4.7003 | 6.7203 | 4.5639
690762317.4553 | sec | 21.9039 | yrs | Tmin, after Sp |  | 6.3096e-05 | 17.388 | 14.8519 | 12.4187 | 12.523 | 9.9452 | 11.3189 | 8.7184 | 8.1762 | 5.7081 | 7.5933 | 5.1839 | 7.4234 | 5.0334
[6 empty] | 0.0001 | 19.3066 | 16.4057 | 13.6965 | 13.8331 | 10.9685 | 12.503 | 9.6155 | 9.0316 | 6.2954 | 8.3876 | 5.7173 | 8.2 | 5.5513
[6 empty] | 0.00015849 | 21.437 | 18.122 | 15.1058 | 15.2803 | 12.0971 | 13.8111 | 10.6049 | 9.9765 | 6.9432 | 9.2651 | 6.3056 | 9.0579 | 6.1225
[6 empty] | 0.00025119 | 23.8025 | 20.018 | 16.6601 | 16.879 | 13.3418 | 15.256 | 11.6961 | 11.0202 | 7.6577 | 10.2345 | 6.9544 | 10.0055 | 6.7525
[6 empty] | 0.00039811 | 26.4292 | 22.1125 | 18.3746 | 18.645 | 14.7148 | 16.8523 | 12.8997 | 12.1733 | 8.4457 | 11.3053 | 7.6701 | 11.0524 | 7.4474
[6 empty] | 0.00063096 | 29.3461 | 24.4263 | 20.2656 | 20.5961 | 16.2292 | 18.6157 | 14.2272 | 13.4471 | 9.3149 | 12.4883 | 8.4595 | 12.209 | 8.2139
[6 empty] | 0.001 | 32.5854 | 26.9827 | 22.3515 | 22.7516 | 17.8997 | 20.5639 | 15.6917 | 14.8544 | 10.2737 | 13.7953 | 9.3302 | 13.4867 | 9.0593
[6 empty] | 0.0015849 | 36.1832 | 29.8073 | 24.6528 | 25.1332 | 19.7425 | 22.7166 | 17.3072 | 16.4094 | 11.3314 | 15.2394 | 10.2908 | 14.8985 | 9.992
[6 empty] | 0.0025119 | 40.1797 | 32.9288 | 27.192 | 27.7653 | 21.776 | 25.0956 | 19.0898 | 18.1278 | 12.4985 | 16.8353 | 11.3507 | 16.4587 | 11.0212
[6 empty] | 0.0039811 | 44.6205 | 36.3794 | 29.9944 | 30.6747 | 24.0202 | 27.7253 | 21.0572 | 20.0274 | 13.7866 | 18.5995 | 12.5206 | 18.1834 | 12.1571
[6 empty] | 0.0063096 | 49.5569 | 40.1952 | 33.0887 | 33.8923 | 26.4982 | 30.6334 | 23.2296 | 22.1281 | 15.2089 | 20.5504 | 13.8122 | 20.0907 | 13.4112
[6 empty] | 0.01 | 55.048 | 44.418 | 36.5075 | 37.4528 | 29.2361 | 33.8516 | 25.6297 | 24.4528 | 16.7803 | 22.7094 | 15.2393 | 22.2013 | 14.7969
[6 empty] | 0.015849 | 61.1629 | 49.096 | 40.289 | 41.3973 | 32.2644 | 37.4168 | 28.2845 | 27.0281 | 18.5184 | 25.1011 | 16.8179 | 24.5395 | 16.3296
[6 empty] | 0.025119 | 67.9842 | 54.2873 | 44.4789 | 45.7746 | 35.6197 | 41.3732 | 31.2259 | 29.8861 | 20.4443 | 27.7552 | 18.5668 | 27.1343 | 18.0278
[6 empty] | 0.039811 | 75.615 | 60.0644 | 49.134 | 50.6457 | 39.3477 | 45.776 | 34.494 | 33.0664 | 22.584 | 30.7088 | 20.51 | 30.0218 | 19.9146
[6 empty] | 0.063096 | 84.1906 | 66.5225 | 54.3297 | 56.0911 | 43.5085 | 50.6978 | 38.1416 | 36.6217 | 24.9721 | 34.0106 | 22.6789 | 33.2498 | 22.0204
[6 empty] | 0.1 | 93.9016 | 73.7967 | 60.1725 | 62.2247 | 48.1876 | 56.2416 | 42.2435 | 40.6263 | 27.6577 | 37.7296 | 25.1178 | 36.8856 | 24.3886