| AUTHOR: | DATE | CHECKED: |
| --- | --- | --- |
| Michael Smith | 1-30-12 | Peter Fritschel |
| Lisa C. Austin | 3-11-12 | Mike Smith |

This specification is for anti-reflection coated, laser-quality, wedged window. It will be used at the input of the OFI to correct the input beam angle

Applicable Documents

Requirements

Physical Configuration

| Diameter | 1.00 +/- 0.01 |
| --- | --- |
| Thickness | 0.375 +/- 0.010 (at thickest point) |
| Wedge | 1 degree +/- 6 arc minute Thick edge of optic to be marked with a scribe line across the barrel |
| Bevel | 45 deg. bevel on both edges, width approx. 0.03 or as appropriate to prevent chipping |
| Clear Aperture | Exceeds 85% of diameter |
| Markings | Markings: Part and serial number, as defined below, shall be scribed on the barrel |

Substrate Material

Fused silica, Corning 7980 or equivalent, grade 0A or better

Part and Serial Number

The Part/Serial number shall be of the format:
D1200434-v1 		YY 		is incremental, starting at 01

Surfaces 1 and 2, measured over the clear aperture

Surface Quality: Surfaces 1 and 2

Laser grade polish: 10-5 scratch-dig surface quality

Transmitted wavefront error:

/20 p-v or less, at 633 nm

Surface roughness, Surfaces 1 and 2:

< 1 Angstrom RMS

Anti-Reflection Coatings:

R < 0.05% for 0-5 deg. AOI, s- & p-polarization, 1064 nm

Applied to surfaces S1 and S2.

Durability per MIL-C-675C, Coating Adhesion and Durability, or current compatible standard, to be approved by LIGO.

Testing and Documentation

| Specification | Test Method | Frequency of Inspection | Data Delivered |
| --- | --- | --- | --- |
| Physical Dimensions | Visual Inspection | 100% | Certification |
| Surface Quality | Visual Inspection | 100% | Certification |
| Transmitted Wavefront Error | Interferometry | 100% | Certification |
| AR Coatings | Spectrophotometer | Witness sample for each coating run | Spectral scans |