## Sheet1
Sorted by QPD # | [8 empty] | Sorted by score
QPD# | Elem# | Dark Current (nA) | Shunt Resistance (Mohm) | Series Resistance (Ohm) | Junction Capacitance (pF) | Dark Noise (1-10Hz avg) | Dark Noise (200-290Hz avg) | Penalty | QPD# | Elem# | Dark Current (nA) | Shunt Resistance (Mohm) | Series Resistance (Ohm) | Junction Capacitance (pF) | Dark Noise (1-10Hz avg) | Dark Noise (200-290Hz avg) | Penalty
1 | 1 | 4.3964 | 7.099 | 14.4174 | 149.8345 | 0.070069 | 0.087311 | -5 | 2 | 1 | 2.5425 | 11.911 | 22.8181 | 151.9683 | 0.063263 | 0.094382 | 0
1 | 2 | 4.6404 | 7.059 | 14.7511 | 150.1339 | 0.22006 | 0.14601 | -5 |  | 2 | 2.5833 | 11.823 | 23.3658 | 153.481 | 0.052718 | 0.089388
1 | 3 | 4.4453 | 7.067 | 15.3584 | 146.7363 | 0.066117 | 0.094552 | -5 |  | 3 | 2.5744 | 11.938 | 24.1637 | 150.2183 | 0.084382 | 0.083609
1 | 4 | 4.6088 | 7.052 | 15.2291 | 148.307 | 0.15724 | 0.080645 | -5 |  | 4 | 2.6471 | 11.619 | 24.1752 | 151.5971 | 0.050931 | 0.084633
2 | 1 | 2.5425 | 11.911 | 22.8181 | 151.9683 | 0.063263 | 0.094382 | 0 | 6 | 1 | 2.3497 | 13.479 | 22.4025 | 152.5724 | 0.057014 | 0.12455 | 0
2 | 2 | 2.5833 | 11.823 | 23.3658 | 153.481 | 0.052718 | 0.089388 | 0 | 6 | 2 | 2.4477 | 12.233 | 23.2343 | 153.1064 | 0.052003 | 0.11626 | 0
2 | 3 | 2.5744 | 11.938 | 24.1637 | 150.2183 | 0.084382 | 0.083609 | 0 | 6 | 3 | 2.4979 | 12.143 | 24.0217 | 150.8768 | 0.051369 | 0.11967 | 0
2 | 4 | 2.6471 | 11.619 | 24.1752 | 151.5971 | 0.050931 | 0.084633 | 0 | 6 | 4 | 2.5447 | 11.862 | 23.6991 | 152.6348 | 0.053863 | 0.11277 | 0
3 | 1 | [7 empty] | 7 | 1 | 2.7404 | 11.702 | 23.0367 | 148.5073 | 0.15922 | 0.12328 | 0
3 | 2 | [7 empty] | 7 | 2 | 2.5394 | 11.838 | 23.9102 | 149.5335 | 0.083486 | 0.11806 | 0
3 | 3 | [7 empty] | 7 | 3 | 2.501 | 12.105 | 24.617 | 146.973 | 0.056477 | 0.11852 | 0
3 | 4 | [7 empty] | 7 | 4 | 2.502 | 11.995 | 24.1695 | 148.4423 | 0.065496 | 0.13754 | 0
4 | 1 | 3.7648 | 8.529 | 13.247 | 152.1115 | 0.25733 | 0.12438 | -10 | 9 | 1 | 2.7835 | 10.414 | 19.6697 | 153.7882 | 0.06794 | 0.13588 | 0
4 | 2 | 3.9141 | 8.302 | 13.9251 | 153.0628 | 0.069854 | 0.086329 | -10 | 9 | 2 | 3.4317 | 9.836 | 20.6477 | 155.4751 | 0.06173 | 0.081548 | 0
4 | 3 | 4.0194 | 8.153 | 14.5093 | 149.389 | 0.064919 | 0.086336 | -10 | 9 | 3 | 2.9626 | 10.003 | 21.0013 | 151.6738 | 0.066136 | 0.12416 | 0
4 | 4 | 4.1994 | 7.71 | 14.4146 | 151.1375 | 0.1815 | 0.082914 | -10 | 9 | 4 | 3.0338 | 9.76 | 21.0406 | 153.6565 | 0.11143 | 0.12653 | 0
5 | 1 | 5.8964 | 10.963 | 23.1291 | 150.2584 | 0.75743 | 0.12297 | -255 | 22 | 1 | 4.6268 | 6.039 | 16.4697 | 140.751 | 0.098339 | 0.11008 | 0
5 | 2 | 7.6309 | 10.362 | 23.8739 | 151.3545 | 6.4716 | 0.95856 | -255 | 22 | 2 | 4.5227 | 6.165 | 17.0719 | 142.7862 | 0.054616 | 0.10704 | 0
5 | 3 | 6.6126 | 10.297 | 24.4115 | 148.6752 | 2.3518 | 0.41941 | -255 | 22 | 3 | 4.4595 | 6.269 | 17.8693 | 139.0334 | 0.06199 | 0.10502 | 0
5 | 4 | 5.7057 | 9.992 | 24.2674 | 149.9318 | 1.4056 | 0.1074 | -255 | 22 | 4 | 4.4576 | 6.287 | 17.836 | 140.409 | 0.059897 | 0.092577 | 0
6 | 1 | 2.3497 | 13.479 | 22.4025 | 152.5724 | 0.057014 | 0.12455 | 0 | 29 | 1 | 3.8057 | 7.095 | 16.3728 | 141.3435 | 0.047356 | 0.077355 | 0
6 | 2 | 2.4477 | 12.233 | 23.2343 | 153.1064 | 0.052003 | 0.11626 | 0 | 29 | 2 | 3.9432 | 7.171 | 17.076 | 141.9189 | 0.056987 | 0.076159 | 0
6 | 3 | 2.4979 | 12.143 | 24.0217 | 150.8768 | 0.051369 | 0.11967 | 0 | 29 | 3 | 3.9546 | 7.088 | 17.6577 | 139.2942 | 0.055763 | 0.12077 | 0
6 | 4 | 2.5447 | 11.862 | 23.6991 | 152.6348 | 0.053863 | 0.11277 | 0 | 29 | 4 | 4.1233 | 6.947 | 17.7386 | 141.827 | 0.056657 | 0.096091 | 0
7 | 1 | 2.7404 | 11.702 | 23.0367 | 148.5073 | 0.15922 | 0.12328 | 0 | 34 | 1 | 4.0416 | 6.636 | 15.2428 | 138.7633 | 0.047511 | 0.1288 | 0
7 | 2 | 2.5394 | 11.838 | 23.9102 | 149.5335 | 0.083486 | 0.11806 | 0 | 34 | 2 | 4.1906 | 6.422 | 16.2272 | 140.2602 | 0.057384 | 0.1092 | 0
7 | 3 | 2.501 | 12.105 | 24.617 | 146.973 | 0.056477 | 0.11852 | 0 | 34 | 3 | 4.2892 | 6.357 | 17.048 | 136.8665 | 0.052947 | 0.1021 | 0
7 | 4 | 2.502 | 11.995 | 24.1695 | 148.4423 | 0.065496 | 0.13754 | 0 | 34 | 4 | 4.2996 | 6.277 | 16.9417 | 138.9972 | 0.047854 | 0.097609 | 0
8 | 1 | 5.1175 | 6.131 | 13.9585 | 150.0045 | 0.061496 | 0.12341 | -15 | 36 | 1 | 4.3357 | 6.327 | 15.1126 | 141.1717 | 0.099594 | 0.058514 | 0
8 | 2 | 4.8168 | 6.524 | 14.8827 | 150.2332 | 0.28052 | 0.11873 | -15 | 36 | 2 | 4.3629 | 6.256 | 15.8367 | 142.0657 | 0.071837 | 0.043946 | 0
8 | 3 | 4.6381 | 6.836 | 15.1295 | 147.4086 | 0.22311 | 0.19129 | -15 | 36 | 3 | 4.3478 | 6.261 | 16.258 | 139.3132 | 0.048578 | 0.043957 | 0
8 | 4 | 4.5584 | 6.931 | 15.2264 | 148.9462 | 0.083744 | 0.12444 | -15 | 36 | 4 | 4.384 | 6.226 | 16.3118 | 140.0759 | 0.072124 | 0.048525 | 0
9 | 1 | 2.7835 | 10.414 | 19.6697 | 153.7882 | 0.06794 | 0.13588 | 0 | 43 | 1 | 4.7292 | 5.745 | 15.1448 | 137.9639 | 0.11341 | 0.095635 | 0
9 | 2 | 3.4317 | 9.836 | 20.6477 | 155.4751 | 0.06173 | 0.081548 | 0 | 43 | 2 | 4.7779 | 5.729 | 16.0622 | 139.531 | 0.11006 | 0.097291 | 0
9 | 3 | 2.9626 | 10.003 | 21.0013 | 151.6738 | 0.066136 | 0.12416 | 0 | 43 | 3 | 4.7675 | 5.769 | 17.1294 | 136.189 | 0.13315 | 0.092677 | 0
9 | 4 | 3.0338 | 9.76 | 21.0406 | 153.6565 | 0.11143 | 0.12653 | 0 | 43 | 4 | 4.745 | 5.759 | 16.7044 | 137.9662 | 0.125 | 0.08899 | 0
10 | 1 | 6.6134 | 8.761 | 19.6962 | 153.7744 | 3.8494 | 0.46684 | -115 | 45 | 1 | 4.4318 | 6.093 | 15.508 | 137.9489 | 0.055574 | 0.098912 | 0
10 | 2 | 4.0702 | 8.589 | 19.8473 | 154.8591 | 0.52178 | 0.15528 | -115 | 45 | 2 | 4.5088 | 5.956 | 16.2204 | 139.3489 | 0.063164 | 0.099485 | 0
10 | 3 | 3.8966 | 8.614 | 21.0685 | 151.5747 | 0.46484 | 0.16391 | -115 | 45 | 3 | 4.5475 | 5.955 | 16.7741 | 135.9899 | 0.098073 | 0.096312 | 0
10 | 4 | 4.1458 | 8.505 | 21.455 | 153.1675 | 0.76081 | 0.14238 | -115 | 45 | 4 | 4.538 | 5.916 | 16.6003 | 137.7662 | 0.051727 | 0.096069 | 0
11 | 1 | [7 empty] | 53 | 1 | 4.0823 | 6.607 | 15.8573 | 137.1163 | 0.083749 | 0.052914 | 0
11 | 2 | [7 empty] | 53 | 2 | 4.6311 | 6.162 | 16.0343 | 137.6158 | 0.12757 | 0.05076 | 0
11 | 3 | [7 empty] | 53 | 3 | 4.6096 | 5.953 | 16.9344 | 134.6688 | 0.16021 | 0.046954 | 0
11 | 4 | [7 empty] | 53 | 4 | 4.7127 | 5.789 | 16.4421 | 136.2628 | 0.11974 | 0.047679 | 0
12 | 1 | [7 empty] | 58 | 1 | 4.1349 | 6.629 | 15.223 | 140.6505 | 0.078072 | 0.1534 | 0
12 | 2 | [7 empty] | 58 | 2 | 4.2849 | 6.392 | 15.811 | 141.8353 | 0.063876 | 0.15512 | 0
12 | 3 | [7 empty] | 58 | 3 | 4.3675 | 6.277 | 16.155 | 138.1437 | 0.06417 | 0.14889 | 0
12 | 4 | [7 empty] | 58 | 4 | 4.4659 | 6.169 | 16.1718 | 140.5555 | 0.078662 | 0.14289 | 0
13 | 1 | [7 empty] | 76 | 1 | 4.2483 | 6.407 | 14.6811 | 138.2108 | 0.10253 | 0.079871 | 0
13 | 2 | [7 empty] | 76 | 2 | 4.4783 | 6.082 | 15.3733 | 139.3022 | 0.11543 | 0.077495 | 0
13 | 3 | [7 empty] | 76 | 3 | 4.6108 | 5.942 | 16.0342 | 136.0342 | 0.15785 | 0.09159 | 0
13 | 4 | [7 empty] | 76 | 4 | 4.683 | 5.825 | 15.8112 | 138.0157 | 0.084422 | 0.081706 | 0
14 | 1 | [7 empty] | 1 | 1 | 4.3964 | 7.099 | 14.4174 | 149.8345 | 0.070069 | 0.087311 | -5
14 | 2 | [7 empty] | 1 | 2 | 4.6404 | 7.059 | 14.7511 | 150.1339 | 0.22006 | 0.14601 | -5
14 | 3 | [7 empty] | 1 | 3 | 4.4453 | 7.067 | 15.3584 | 146.7363 | 0.066117 | 0.094552 | -5
14 | 4 | [7 empty] | 1 | 4 | 4.6088 | 7.052 | 15.2291 | 148.307 | 0.15724 | 0.080645 | -5
15 | 1 | [7 empty] | 39 | 1 | 4.3616 | 6.479 | 14.9128 | 138.5626 | 0.25256 | 0.14927 | -5
15 | 2 | [7 empty] | 39 | 2 | 4.3383 | 6.299 | 15.6758 | 139.8309 | 0.13543 | 0.14276 | -5
15 | 3 | [7 empty] | 39 | 3 | 4.3642 | 6.24 | 16.1798 | 136.265 | 0.099669 | 0.13511 | -5
15 | 4 | [7 empty] | 39 | 4 | 4.4257 | 6.174 | 16.1966 | 138.3948 | 0.086916 | 0.13372 | -5
16 | 1 | 3.443 | 35.8 | 34.0227 | 135.0245 | 2.3294 | 0.41695 | -265 | 55 | 1 | 4.3581 | 6.379 | 15.904 | 137.5466 | 0.12549 | 0.086451 | -5
16 | 2 | 5.3025 | 34.179 | 34.2454 | 136.0903 | 3.4345 | 0.68954 | -265 | 55 | 2 | 4.5489 | 6.047 | 16.3815 | 136.8164 | 0.29382 | 0.072799 | -5
16 | 3 | 3.2068 | 37.267 | 35.5304 | 133.0175 | 2.6765 | 0.31583 | -265 | 55 | 3 | 4.6357 | 5.911 | 16.9152 | 134.4894 | 0.087812 | 0.072984 | -5
16 | 4 | 2.451 | 39.913 | 34.9018 | 134.7898 | 1.4949 | 0.35127 | -265 | 55 | 4 | 4.7624 | 5.812 | 16.7028 | 136.0348 | 0.12087 | 0.064085 | -5
17 | 1 | 1.4195 | 43.557 | 30.3862 | 138.6113 | 0.19615 | 0.051494 | -20 | 57 | 1 | 3.9706 | 6.962 | 15.9701 | 140.6681 | 0.056766 | 0.15284 | -5
17 | 2 | 1.5744 | 42.767 | 30.5883 | 139.9965 | 0.44202 | 0.064093 | -20 | 57 | 2 | 4.2632 | 6.54 | 16.6977 | 141.1354 | 0.086741 | 0.14869 | -5
17 | 3 | 1.4428 | 43.997 | 32.1876 | 136.6055 | 0.19468 | 0.054177 | -20 | 57 | 3 | 4.3743 | 6.357 | 16.9776 | 138.239 | 0.080773 | 0.14801 | -5
17 | 4 | 1.5471 | 43.078 | 32.043 | 137.5363 | 0.19962 | 0.073266 | -20 | 57 | 4 | 4.4811 | 6.208 | 17.0336 | 140.393 | 0.21081 | 0.14718 | -5
18 | 1 | 1.476 | 41.677 | 34.8583 | 133.1385 | 0.3781 | 0.063037 | -65 | 4 | 1 | 3.7648 | 8.529 | 13.247 | 152.1115 | 0.25733 | 0.12438 | -10
18 | 2 | 1.6147 | 43.597 | 35.0119 | 134.2182 | 0.41027 | 0.07983 | -65 | 4 | 2 | 3.9141 | 8.302 | 13.9251 | 153.0628 | 0.069854 | 0.086329 | -10
18 | 3 | 1.4186 | 44.842 | 36.8603 | 131.0236 | 0.40293 | 0.088804 | -65 | 4 | 3 | 4.0194 | 8.153 | 14.5093 | 149.389 | 0.064919 | 0.086336 | -10
18 | 4 | 1.6422 | 42.798 | 36.6148 | 132.4966 | 0.63974 | 0.1256 | -65 | 4 | 4 | 4.1994 | 7.71 | 14.4146 | 151.1375 | 0.1815 | 0.082914 | -10
19 | 1 | 3.2137 | 45.029 | 28.997 | 136.7275 | 2.6806 | 0.59035 | -265 | 46 | 1 | 4.2953 | 6.451 | 15.2023 | 137.6717 | 0.1427 | 0.12256 | -10
19 | 2 | 2.6059 | 45.145 | 30.0515 | 138.1055 | 1.4117 | 0.32691 | -265 | 46 | 2 | 4.5012 | 6.181 | 15.7664 | 139.3706 | 0.13811 | 0.11865 | -10
19 | 3 | 3.4361 | 41.101 | 30.8265 | 135.3246 | 1.9771 | 0.35412 | -265 | 46 | 3 | 4.6324 | 6.071 | 16.1368 | 135.9287 | 0.22916 | 0.12478 | -10
19 | 4 | 2.5174 | 44.225 | 30.4923 | 137.6236 | 1.6784 | 0.30912 | -265 | 46 | 4 | 4.7514 | 5.978 | 16.0449 | 137.6437 | 0.24332 | 0.11904 | -10
20 | 1 | 1.398 | 43.131 | 34.4819 | 134.1571 | 0.2454 | 0.13026 | -20 | 50 | 1 | 4.3592 | 6.536 | 15.5271 | 140.9711 | 0.17775 | 0.11911 | -10
20 | 2 | 1.5094 | 42.858 | 34.7685 | 135.8586 | 0.30026 | 0.139 | -20 | 50 | 2 | 4.4929 | 6.408 | 16.3361 | 142.0748 | 0.13971 | 0.12399 | -10
20 | 3 | 2.4225 | 28.368 | 35.6763 | 133.117 | 0.33621 | 0.12674 | -20 | 50 | 3 | 4.5834 | 6.361 | 16.8788 | 138.7704 | 0.18557 | 0.12661 | -10
20 | 4 | 1.4297 | 42.705 | 35.4421 | 135.0175 | 0.31232 | 0.1173 | -20 | 50 | 4 | 4.5888 | 6.272 | 16.8243 | 141.0626 | 0.19309 | 0.11752 | -10
21 | 1 | 5.0656 | 5.4 | 15.0301 | 138.1456 | 1.1384 | 0.086115 | -105 | 8 | 1 | 5.1175 | 6.131 | 13.9585 | 150.0045 | 0.061496 | 0.12341 | -15
21 | 2 | 4.9337 | 5.61 | 15.8176 | 139.795 | 0.57234 | 0.092375 | -105 | 8 | 2 | 4.8168 | 6.524 | 14.8827 | 150.2332 | 0.28052 | 0.11873 | -15
21 | 3 | 4.7915 | 5.769 | 16.4854 | 136.4257 | 0.13682 | 0.088106 | -105 | 8 | 3 | 4.6381 | 6.836 | 15.1295 | 147.4086 | 0.22311 | 0.19129 | -15
21 | 4 | 4.689 | 5.841 | 16.5155 | 138.3545 | 0.29245 | 0.07372 | -105 | 8 | 4 | 4.5584 | 6.931 | 15.2264 | 148.9462 | 0.083744 | 0.12444 | -15
22 | 1 | 4.6268 | 6.039 | 16.4697 | 140.751 | 0.098339 | 0.11008 | 0 | 24 | 1 | 4.2856 | 7.112 | 18.5685 | 135.5126 | 0.21691 | 0.17924 | -15
22 | 2 | 4.5227 | 6.165 | 17.0719 | 142.7862 | 0.054616 | 0.10704 | 0 | 24 | 2 | 4.2916 | 7.084 | 19.1251 | 137.6753 | 0.17149 | 0.17638 | -15
22 | 3 | 4.4595 | 6.269 | 17.8693 | 139.0334 | 0.06199 | 0.10502 | 0 | 24 | 3 | 4.2272 | 7.115 | 19.631 | 133.7378 | 0.24221 | 0.17441 | -15
22 | 4 | 4.4576 | 6.287 | 17.836 | 140.409 | 0.059897 | 0.092577 | 0 | 24 | 4 | 4.3091 | 7.092 | 19.7696 | 135.5542 | 0.43737 | 0.1709 | -15
23 | 1 | 104.886 | 0.52482 | 19.2204 | 134.4593 | 5.4375 | 0.65521 | -600 | 17 | 1 | 1.4195 | 43.557 | 30.3862 | 138.6113 | 0.19615 | 0.051494 | -20
23 | 2 | 100.473 | 0.53754 | 20.0813 | 134.9474 | 5.9365 | 0.58716 | -600 | 17 | 2 | 1.5744 | 42.767 | 30.5883 | 139.9965 | 0.44202 | 0.064093 | -20
23 | 3 | 110.61 | 0.53104 | 20.7149 | 132.2622 | 5.6007 | 0.6133 | -600 | 17 | 3 | 1.4428 | 43.997 | 32.1876 | 136.6055 | 0.19468 | 0.054177 | -20
23 | 4 | 106.631 | 0.51791 | 20.4626 | 134.1588 | 5.574 | 0.64177 | -600 | 17 | 4 | 1.5471 | 43.078 | 32.043 | 137.5363 | 0.19962 | 0.073266 | -20
24 | 1 | 4.2856 | 7.112 | 18.5685 | 135.5126 | 0.21691 | 0.17924 | -15 | 20 | 1 | 1.398 | 43.131 | 34.4819 | 134.1571 | 0.2454 | 0.13026 | -20
24 | 2 | 4.2916 | 7.084 | 19.1251 | 137.6753 | 0.17149 | 0.17638 | -15 | 20 | 2 | 1.5094 | 42.858 | 34.7685 | 135.8586 | 0.30026 | 0.139 | -20
24 | 3 | 4.2272 | 7.115 | 19.631 | 133.7378 | 0.24221 | 0.17441 | -15 | 20 | 3 | 2.4225 | 28.368 | 35.6763 | 133.117 | 0.33621 | 0.12674 | -20
24 | 4 | 4.3091 | 7.092 | 19.7696 | 135.5542 | 0.43737 | 0.1709 | -15 | 20 | 4 | 1.4297 | 42.705 | 35.4421 | 135.0175 | 0.31232 | 0.1173 | -20
25 | 1 | 8.8528 | 3.841 | 18.3242 | 136.5791 | 1.2096 | 0.19883 | -210 | 38 | 1 | 4.1613 | 8.199 | 19.3987 | 135.0844 | 0.25068 | 0.14109 | -20
25 | 2 | 9.3754 | 3.681 | 19.1377 | 137.6749 | 1.3275 | 0.20741 | -210 | 38 | 2 | 4.1224 | 8.092 | 20.1132 | 136.6354 | 0.28734 | 0.14164 | -20
25 | 3 | 9.6491 | 3.649 | 19.9622 | 134.0737 | 1.6802 | 0.1777 | -210 | 38 | 3 | 3.9931 | 8.133 | 20.5466 | 133.445 | 0.22737 | 0.13801 | -20
25 | 4 | 9.7724 | 3.602 | 19.6668 | 136.7933 | 1.3834 | 0.17446 | -210 | 38 | 4 | 4.1111 | 8.028 | 20.6968 | 135.0281 | 0.31013 | 0.1278 | -20
26 | 1 | 5.1704 | 7.4 | 19.9567 | 134.3881 | 0.40553 | 0.18878 | -80 | 44 | 1 | 4.328 | 7.62 | 17.9525 | 138.5812 | 0.25744 | 0.064456 | -20
26 | 2 | 4.9955 | 7.323 | 20.7742 | 136.1258 | 0.32579 | 0.19295 | -80 | 44 | 2 | 4.2053 | 7.363 | 19.5068 | 139.036 | 0.23376 | 0.081928 | -20
26 | 3 | 5.2008 | 7.319 | 21.3462 | 131.9464 | 0.44362 | 0.18962 | -80 | 44 | 3 | 4.5634 | 7.186 | 19.983 | 136.2842 | 0.3141 | 0.073425 | -20
26 | 4 | 5.5289 | 7.262 | 21.0774 | 134.332 | 0.91062 | 0.17835 | -80 | 44 | 4 | 4.683 | 7.064 | 19.8751 | 138.727 | 0.37517 | 0.074151 | -20
27 | 1 | 52.4652 | 1.264 | 18.6371 | 135.0774 | 10.7771 | 0.92376 | -420 | 51 | 1 | 3.6975 | 10.121 | 20.9647 | 132.917 | 0.28239 | 0.13544 | -20
27 | 2 | 53.5538 | 1.245 | 19.0932 | 136.1872 | 9.9415 | 0.96719 | -420 | 51 | 2 | 3.8048 | 9.156 | 20.8586 | 134.9159 | 0.28882 | 0.13475 | -20
27 | 3 | 54.7801 | 1.235 | 19.4993 | 132.7102 | 12.3277 | 1.0313 | -420 | 51 | 3 | 4.1153 | 8.727 | 21.4992 | 131.1407 | 0.30108 | 0.11965 | -20
27 | 4 | 53.7946 | 1.235 | 19.709 | 134.0705 | 11.6977 | 0.97275 | -420 | 51 | 4 | 4.3477 | 8.305 | 21.5718 | 132.9817 | 0.27489 | 0.081275 | -20
28 | 1 | 4.8063 | 5.725 | 15.1532 | 136.6662 | 0.13977 | 0.16901 | -55 | 54 | 1 | 4.4005 | 8.641 | 21.7986 | 134.2117 | 0.4271 | 0.066125 | -20
28 | 2 | 4.836 | 5.717 | 15.7726 | 138.2628 | 0.1208 | 0.16715 | -55 | 54 | 2 | 4.4849 | 8.259 | 21.8343 | 134.7858 | 0.32557 | 0.065177 | -20
28 | 3 | 4.8055 | 5.734 | 15.9905 | 135.474 | 0.098446 | 0.15748 | -55 | 54 | 3 | 4.5748 | 8.149 | 22.5552 | 131.1404 | 0.27593 | 0.053753 | -20
28 | 4 | 7.3712 | 5.694 | 16.2442 | 136.9433 | 1.2541 | 0.20915 | -55 | 54 | 4 | 4.6361 | 7.985 | 22.315 | 132.7387 | 0.49101 | 0.069997 | -20
29 | 1 | 3.8057 | 7.095 | 16.3728 | 141.3435 | 0.047356 | 0.077355 | 0 | 56 | 1 | 3.6812 | 8.891 | 20.7197 | 135.4234 | 0.27369 | 0.13609 | -20
29 | 2 | 3.9432 | 7.171 | 17.076 | 141.9189 | 0.056987 | 0.076159 | 0 | 56 | 2 | 3.9608 | 8.395 | 20.9511 | 136.1513 | 0.28828 | 0.13457 | -20
29 | 3 | 3.9546 | 7.088 | 17.6577 | 139.2942 | 0.055763 | 0.12077 | 0 | 56 | 3 | 4.083 | 8.237 | 21.9806 | 132.5639 | 0.34804 | 0.12097 | -20
29 | 4 | 4.1233 | 6.947 | 17.7386 | 141.827 | 0.056657 | 0.096091 | 0 | 56 | 4 | 4.1661 | 8.022 | 21.2761 | 134.0929 | 0.42166 | 0.13136 | -20
30 | 1 | 5.4868 | 6.304 | 16.6592 | 140.8387 | 0.41513 | 0.1505 | -25 | 63 | 1 | 4.1193 | 7.468 | 14.7098 | 135.9938 | 0.26413 | 0.049805 | -20
30 | 2 | 5.0834 | 6.168 | 16.5149 | 140.777 | 0.43433 | 0.25708 | -25 | 63 | 2 | 4.5037 | 6.657 | 15.5757 | 136.7661 | 0.24555 | 0.054207 | -20
30 | 3 | 5.0898 | 6.101 | 16.9946 | 137.3995 | 0.30387 | 0.16517 | -25 | 63 | 3 | 5.1031 | 6.232 | 16.1916 | 134.0473 | 0.32439 | 0.049953 | -20
30 | 4 | 5.0101 | 6.022 | 16.9261 | 139.7337 | 0.27957 | 0.1513 | -25 | 63 | 4 | 5.1496 | 5.929 | 15.9425 | 135.229 | 0.2165 | 0.047817 | -20
31 | 1 | 5.4866 | 7.186 | 15.5689 | 131.3756 | 2.5639 | 0.20361 | -220 | 68 | 1 | 3.6384 | 9.112 | 20.1902 | 135.3932 | 0.21254 | 0.053575 | -20
31 | 2 | 5.9163 | 6.586 | 16.2766 | 133.0441 | 0.89315 | 0.1878 | -220 | 68 | 2 | 4.1316 | 8.688 | 20.9773 | 137.0335 | 0.34255 | 0.066539 | -20
31 | 3 | 8.3318 | 6.267 | 17.3021 | 129.3452 | 1.1746 | 0.26325 | -220 | 68 | 3 | 4.4471 | 8.548 | 21.8351 | 133.4975 | 0.49484 | 0.068611 | -20
31 | 4 | 6.4477 | 6.044 | 17.0102 | 131.473 | 1.1182 | 0.21189 | -220 | 68 | 4 | 4.0205 | 8.333 | 21.4758 | 135.0656 | 0.4098 | 0.061784 | -20
32 | 1 | 6.9038 | 5.94 | 15.2105 | 137.8106 | 0.61494 | 0.12863 | -110 | 73 | 1 | 4.805 | 6.112 | 14.2394 | 140.2068 | 0.25846 | 0.051337 | -20
32 | 2 | 6.2741 | 5.789 | 15.9675 | 138.7305 | 0.51905 | 0.13404 | -110 | 73 | 2 | 4.8188 | 5.889 | 14.964 | 140.2616 | 0.18677 | 0.046877 | -20
32 | 3 | 6.5598 | 5.728 | 16.4556 | 135.4726 | 0.55661 | 0.1214 | -110 | 73 | 3 | 4.9441 | 5.803 | 15.474 | 137.3414 | 0.20757 | 0.046304 | -20
32 | 4 | 7.3975 | 5.672 | 16.5206 | 137.8898 | 0.78761 | 0.13181 | -110 | 73 | 4 | 5.1512 | 5.713 | 15.3533 | 139.0028 | 0.26927 | 0.049606 | -20
33 | 1 | 17.8367 | 2.707 | 15.0923 | 142.9027 | 28.5309 | 1.3066 | -420 | 75 | 1 | 5.5826 | 5.81 | 14.7931 | 137.3327 | 0.31636 | 0.068775 | -20
33 | 2 | 18.9232 | 2.615 | 15.8969 | 142.7054 | 25.9452 | 1.2978 | -420 | 75 | 2 | 5.6538 | 5.637 | 15.6247 | 137.5287 | 0.33279 | 0.073566 | -20
33 | 3 | 18.7826 | 2.549 | 16.1686 | 141.2023 | 22.1601 | 1.2789 | -420 | 75 | 3 | 5.8325 | 5.553 | 16.4179 | 133.9771 | 0.3479 | 0.06855 | -20
33 | 4 | 19.1805 | 2.528 | 16.144 | 142.0302 | 22.8827 | 1.6687 | -420 | 75 | 4 | 6.0911 | 5.461 | 16.2948 | 136.2947 | 0.38618 | 0.079298 | -20
34 | 1 | 4.0416 | 6.636 | 15.2428 | 138.7633 | 0.047511 | 0.1288 | 0 | 30 | 1 | 5.4868 | 6.304 | 16.6592 | 140.8387 | 0.41513 | 0.1505 | -25
34 | 2 | 4.1906 | 6.422 | 16.2272 | 140.2602 | 0.057384 | 0.1092 | 0 | 30 | 2 | 5.0834 | 6.168 | 16.5149 | 140.777 | 0.43433 | 0.25708 | -25
34 | 3 | 4.2892 | 6.357 | 17.048 | 136.8665 | 0.052947 | 0.1021 | 0 | 30 | 3 | 5.0898 | 6.101 | 16.9946 | 137.3995 | 0.30387 | 0.16517 | -25
34 | 4 | 4.2996 | 6.277 | 16.9417 | 138.9972 | 0.047854 | 0.097609 | 0 | 30 | 4 | 5.0101 | 6.022 | 16.9261 | 139.7337 | 0.27957 | 0.1513 | -25
35 | 1 | 24.7129 | 2.053 | 18.5002 | 137.881 | 10.0424 | 0.70603 | -420 | 52 | 1 | 4.5763 | 6.233 | 15.0751 | 133.5736 | 0.1961 | 0.19304 | -25
35 | 2 | 25.8295 | 2.006 | 18.9138 | 138.3494 | 11.4574 | 0.82512 | -420 | 52 | 2 | 4.9303 | 6.001 | 15.7038 | 134.9386 | 0.29919 | 0.19632 | -25
35 | 3 | 26.0365 | 1.977 | 19.4429 | 135.6102 | 12.7158 | 0.82719 | -420 | 52 | 3 | 4.7876 | 5.906 | 16.2294 | 131.4823 | 0.31429 | 0.15414 | -25
35 | 4 | 25.8404 | 1.977 | 19.4773 | 136.2931 | 12.0441 | 0.80717 | -420 | 52 | 4 | 4.9122 | 5.81 | 16.2141 | 133.2142 | 0.16532 | 0.11297 | -25
36 | 1 | 4.3357 | 6.327 | 15.1126 | 141.1717 | 0.099594 | 0.058514 | 0 | 47 | 1 | 4.461 | 7.387 | 17.2914 | 137.8292 | 0.27702 | 0.1847 | -35
36 | 2 | 4.3629 | 6.256 | 15.8367 | 142.0657 | 0.071837 | 0.043946 | 0 | 47 | 2 | 4.699 | 7.01 | 18.0508 | 139.2402 | 0.29976 | 0.18373 | -35
36 | 3 | 4.3478 | 6.261 | 16.258 | 139.3132 | 0.048578 | 0.043957 | 0 | 47 | 3 | 4.8947 | 6.854 | 18.532 | 135.9515 | 0.29203 | 0.1816 | -35
36 | 4 | 4.384 | 6.226 | 16.3118 | 140.0759 | 0.072124 | 0.048525 | 0 | 47 | 4 | 5.4852 | 6.718 | 18.6891 | 138.2191 | 0.43372 | 0.17986 | -35
37 | 1 | 7.8293 | 5.462 | 16.6566 | 136.4061 | 1.2269 | 0.2823 | -120 | 64 | 1 | 3.7505 | 7.547 | 15.4255 | 140.247 | 0.10175 | 0.10201 | -50
37 | 2 | 6.2735 | 5.671 | 17.5517 | 138.0005 | 0.54635 | 0.17752 | -120 | 64 | 2 | 5.0124 | 6.849 | 16.147 | 141.2571 | 0.56052 | 0.14943 | -50
37 | 3 | 5.1651 | 5.828 | 17.8275 | 134.2455 | 0.24103 | 0.14896 | -120 | 64 | 3 | 4.4288 | 6.468 | 16.5524 | 138.3338 | 0.17999 | 0.089225 | -50
37 | 4 | 7.3313 | 5.886 | 17.9432 | 135.8938 | 1.3172 | 0.46034 | -120 | 64 | 4 | 4.6204 | 6.191 | 16.3856 | 140.6044 | 0.15862 | 0.065673 | -50
38 | 1 | 4.1613 | 8.199 | 19.3987 | 135.0844 | 0.25068 | 0.14109 | -20 | 80 | 1 | 3.9555 | 6.753 | 14.5603 | 136.8074 | 0.075366 | 0.16735 | -50
38 | 2 | 4.1224 | 8.092 | 20.1132 | 136.6354 | 0.28734 | 0.14164 | -20 | 80 | 2 | 4.2582 | 6.297 | 15.1378 | 138.2429 | 0.092188 | 0.11045 | -50
38 | 3 | 3.9931 | 8.133 | 20.5466 | 133.445 | 0.22737 | 0.13801 | -20 | 80 | 3 | 4.4427 | 6.065 | 15.8577 | 134.9259 | 0.84258 | 0.10751 | -50
38 | 4 | 4.1111 | 8.028 | 20.6968 | 135.0281 | 0.31013 | 0.1278 | -20 | 80 | 4 | 4.5706 | 5.9 | 15.7672 | 136.2188 | 0.054267 | 0.11388 | -50
39 | 1 | 4.3616 | 6.479 | 14.9128 | 138.5626 | 0.25256 | 0.14927 | -5 | 28 | 1 | 4.8063 | 5.725 | 15.1532 | 136.6662 | 0.13977 | 0.16901 | -55
39 | 2 | 4.3383 | 6.299 | 15.6758 | 139.8309 | 0.13543 | 0.14276 | -5 | 28 | 2 | 4.836 | 5.717 | 15.7726 | 138.2628 | 0.1208 | 0.16715 | -55
39 | 3 | 4.3642 | 6.24 | 16.1798 | 136.265 | 0.099669 | 0.13511 | -5 | 28 | 3 | 4.8055 | 5.734 | 15.9905 | 135.474 | 0.098446 | 0.15748 | -55
39 | 4 | 4.4257 | 6.174 | 16.1966 | 138.3948 | 0.086916 | 0.13372 | -5 | 28 | 4 | 7.3712 | 5.694 | 16.2442 | 136.9433 | 1.2541 | 0.20915 | -55
40 | 1 | 6.0903 | 5.948 | 14.6648 | 138.3103 | 1.2692 | 0.1693 | -205 | 66 | 1 | 3.3666 | 8.667 | 18.2721 | 135.7426 | 0.11995 | 0.047502 | -55
40 | 2 | 6.7014 | 5.805 | 15.2707 | 139.3038 | 0.92774 | 0.19373 | -205 | 66 | 2 | 3.7425 | 7.907 | 18.6979 | 137.1581 | 0.3192 | 0.078922 | -55
40 | 3 | 6.4497 | 5.738 | 15.9797 | 135.7404 | 0.66131 | 0.17007 | -205 | 66 | 3 | 3.8249 | 7.549 | 19.4347 | 133.6664 | 0.10182 | 0.05751 | -55
40 | 4 | 6.6641 | 5.675 | 15.9008 | 138.0984 | 1.5561 | 0.16188 | -205 | 66 | 4 | 3.9944 | 7.234 | 19.3377 | 134.9237 | 2.5897 | 0.066221 | -55
41 | 1 | 31.1935 | 1.665 | 15.4732 | 143.5999 | 14.0179 | 1.2012 | -420 | 18 | 1 | 1.476 | 41.677 | 34.8583 | 133.1385 | 0.3781 | 0.063037 | -65
41 | 2 | 33.0587 | 1.589 | 16.3793 | 143.6402 | 15.06 | 1.3583 | -420 | 18 | 2 | 1.6147 | 43.597 | 35.0119 | 134.2182 | 0.41027 | 0.07983 | -65
41 | 3 | 34.1801 | 1.546 | 16.876 | 141.4274 | 13.6898 | 1.2979 | -420 | 18 | 3 | 1.4186 | 44.842 | 36.8603 | 131.0236 | 0.40293 | 0.088804 | -65
41 | 4 | 34.5848 | 1.521 | 17.0127 | 142.7072 | 13.7172 | 1.2531 | -420 | 18 | 4 | 1.6422 | 42.798 | 36.6148 | 132.4966 | 0.63974 | 0.1256 | -65
42 | 1 | 10.6657 | 4.204 | 18.2912 | 138.4385 | 10.725 | 0.52716 | -225 | 62 | 1 | 6.1511 | 6.279 | 15.2205 | 135.5888 | 0.51916 | 0.17436 | -65
42 | 2 | 10.842 | 4.103 | 19.1062 | 140.4393 | 6.7298 | 0.033971 | -225 | 62 | 2 | 6.5086 | 5.94 | 15.6917 | 137.7057 | 0.55621 | 0.17496 | -65
42 | 3 | 71.3285 | 3.974 | 19.2264 | 136.6253 | 0.0013553 | 0.00055236 | -225 | 62 | 3 | 6.5876 | 5.761 | 16.3065 | 133.9438 | 0.4984 | 0.16223 | -65
42 | 4 | 11.3854 | 3.983 | 19.1359 | 138.0956 | 0.0013281 | 0.00055707 | -225 | 62 | 4 | 6.763 | 5.599 | 16.3874 | 135.2198 | 0.71377 | 0.1693 | -65
43 | 1 | 4.7292 | 5.745 | 15.1448 | 137.9639 | 0.11341 | 0.095635 | 0 | 70 | 1 | 6.3457 | 6.466 | 15.5627 | 134.6586 | 0.48311 | 0.15769 | -65
43 | 2 | 4.7779 | 5.729 | 16.0622 | 139.531 | 0.11006 | 0.097291 | 0 | 70 | 2 | 6.5838 | 5.956 | 16.0457 | 136.9131 | 0.47355 | 0.15601 | -65
43 | 3 | 4.7675 | 5.769 | 17.1294 | 136.189 | 0.13315 | 0.092677 | 0 | 70 | 3 | 6.8334 | 5.694 | 16.5534 | 133.516 | 0.60152 | 0.15189 | -65
43 | 4 | 4.745 | 5.759 | 16.7044 | 137.9662 | 0.125 | 0.08899 | 0 | 70 | 4 | 7.1662 | 5.504 | 16.4696 | 134.4972 | 0.53499 | 0.14577 | -65
44 | 1 | 4.328 | 7.62 | 17.9525 | 138.5812 | 0.25744 | 0.064456 | -20 | 74 | 1 | 4.9454 | 6.587 | 15.6743 | 139.4919 | 0.36181 | 0.063451 | -65
44 | 2 | 4.2053 | 7.363 | 19.5068 | 139.036 | 0.23376 | 0.081928 | -20 | 74 | 2 | 5.8553 | 6.273 | 16.2733 | 140.5164 | 1.2248 | 0.093917 | -65
44 | 3 | 4.5634 | 7.186 | 19.983 | 136.2842 | 0.3141 | 0.073425 | -20 | 74 | 3 | 5.0311 | 6.142 | 16.804 | 137.1043 | 0.30254 | 0.060743 | -65
44 | 4 | 4.683 | 7.064 | 19.8751 | 138.727 | 0.37517 | 0.074151 | -20 | 74 | 4 | 5.2191 | 6.013 | 16.8859 | 138.5953 | 0.34105 | 0.066814 | -65
45 | 1 | 4.4318 | 6.093 | 15.508 | 137.9489 | 0.055574 | 0.098912 | 0 | 26 | 1 | 5.1704 | 7.4 | 19.9567 | 134.3881 | 0.40553 | 0.18878 | -80
45 | 2 | 4.5088 | 5.956 | 16.2204 | 139.3489 | 0.063164 | 0.099485 | 0 | 26 | 2 | 4.9955 | 7.323 | 20.7742 | 136.1258 | 0.32579 | 0.19295 | -80
45 | 3 | 4.5475 | 5.955 | 16.7741 | 135.9899 | 0.098073 | 0.096312 | 0 | 26 | 3 | 5.2008 | 7.319 | 21.3462 | 131.9464 | 0.44362 | 0.18962 | -80
45 | 4 | 4.538 | 5.916 | 16.6003 | 137.7662 | 0.051727 | 0.096069 | 0 | 26 | 4 | 5.5289 | 7.262 | 21.0774 | 134.332 | 0.91062 | 0.17835 | -80
46 | 1 | 4.2953 | 6.451 | 15.2023 | 137.6717 | 0.1427 | 0.12256 | -10 | 21 | 1 | 5.0656 | 5.4 | 15.0301 | 138.1456 | 1.1384 | 0.086115 | -105
46 | 2 | 4.5012 | 6.181 | 15.7664 | 139.3706 | 0.13811 | 0.11865 | -10 | 21 | 2 | 4.9337 | 5.61 | 15.8176 | 139.795 | 0.57234 | 0.092375 | -105
46 | 3 | 4.6324 | 6.071 | 16.1368 | 135.9287 | 0.22916 | 0.12478 | -10 | 21 | 3 | 4.7915 | 5.769 | 16.4854 | 136.4257 | 0.13682 | 0.088106 | -105
46 | 4 | 4.7514 | 5.978 | 16.0449 | 137.6437 | 0.24332 | 0.11904 | -10 | 21 | 4 | 4.689 | 5.841 | 16.5155 | 138.3545 | 0.29245 | 0.07372 | -105
47 | 1 | 4.461 | 7.387 | 17.2914 | 137.8292 | 0.27702 | 0.1847 | -35 | 32 | 1 | 6.9038 | 5.94 | 15.2105 | 137.8106 | 0.61494 | 0.12863 | -110
47 | 2 | 4.699 | 7.01 | 18.0508 | 139.2402 | 0.29976 | 0.18373 | -35 | 32 | 2 | 6.2741 | 5.789 | 15.9675 | 138.7305 | 0.51905 | 0.13404 | -110
47 | 3 | 4.8947 | 6.854 | 18.532 | 135.9515 | 0.29203 | 0.1816 | -35 | 32 | 3 | 6.5598 | 5.728 | 16.4556 | 135.4726 | 0.55661 | 0.1214 | -110
47 | 4 | 5.4852 | 6.718 | 18.6891 | 138.2191 | 0.43372 | 0.17986 | -35 | 32 | 4 | 7.3975 | 5.672 | 16.5206 | 137.8898 | 0.78761 | 0.13181 | -110
48 | 1 | 7.4528 | 5.452 | 21.0686 | 134.8884 | 0.97664 | 0.1785 | -210 | 72 | 1 | 6.4166 | 5.608 | 15.5826 | 135.3885 | 0.53342 | 0.15178 | -110
48 | 2 | 9.4607 | 5.201 | 21.7403 | 136.6733 | 1.867 | 0.46916 | -210 | 72 | 2 | 6.6362 | 5.639 | 16.111 | 137.251 | 0.5198 | 0.14384 | -110
48 | 3 | 9.4489 | 5.118 | 22.0618 | 132.6916 | 1.7566 | 0.43536 | -210 | 72 | 3 | 6.5246 | 5.677 | 16.4193 | 133.8186 | 0.72032 | 0.14365 | -110
48 | 4 | 8.0145 | 5.075 | 22.0656 | 134.7525 | 0.98832 | 0.17806 | -210 | 72 | 4 | 6.4386 | 5.67 | 16.5582 | 135.0145 | 1.2372 | 0.14096 | -110
49 | 1 | 13.7778 | 3.083 | 19 | 135.7228 | 2.2641 | 0.18902 | -240 | 79 | 1 | 5.9353 | 6.181 | 15.2847 | 134.752 | 0.45654 | 0.16493 | -110
49 | 2 | 15.0417 | 2.892 | 19.6218 | 136.4361 | 2.3796 | 0.31337 | -240 | 79 | 2 | 6.0295 | 5.887 | 15.8756 | 135.6469 | 0.88676 | 0.16973 | -110
49 | 3 | 14.4729 | 2.773 | 20.404 | 133.4906 | 2.3412 | 0.18614 | -240 | 79 | 3 | 6.3624 | 5.755 | 16.4741 | 132.469 | 0.43876 | 0.15439 | -110
49 | 4 | 15.4296 | 2.693 | 20.3135 | 134.1614 | 2.3943 | 0.24154 | -240 | 79 | 4 | 6.8387 | 5.656 | 16.1881 | 134.606 | 0.56573 | 0.15862 | -110
50 | 1 | 4.3592 | 6.536 | 15.5271 | 140.9711 | 0.17775 | 0.11911 | -10 | 10 | 1 | 6.6134 | 8.761 | 19.6962 | 153.7744 | 3.8494 | 0.46684 | -115
50 | 2 | 4.4929 | 6.408 | 16.3361 | 142.0748 | 0.13971 | 0.12399 | -10 | 10 | 2 | 4.0702 | 8.589 | 19.8473 | 154.8591 | 0.52178 | 0.15528 | -115
50 | 3 | 4.5834 | 6.361 | 16.8788 | 138.7704 | 0.18557 | 0.12661 | -10 | 10 | 3 | 3.8966 | 8.614 | 21.0685 | 151.5747 | 0.46484 | 0.16391 | -115
50 | 4 | 4.5888 | 6.272 | 16.8243 | 141.0626 | 0.19309 | 0.11752 | -10 | 10 | 4 | 4.1458 | 8.505 | 21.455 | 153.1675 | 0.76081 | 0.14238 | -115
51 | 1 | 3.6975 | 10.121 | 20.9647 | 132.917 | 0.28239 | 0.13544 | -20 | 65 | 1 | 5.0494 | 8.414 | 17.2703 | 139.3918 | 1.399 | 0.23017 | -115
51 | 2 | 3.8048 | 9.156 | 20.8586 | 134.9159 | 0.28882 | 0.13475 | -20 | 65 | 2 | 4.3911 | 7.577 | 17.7231 | 140.7551 | 0.66485 | 0.33642 | -115
51 | 3 | 4.1153 | 8.727 | 21.4992 | 131.1407 | 0.30108 | 0.11965 | -20 | 65 | 3 | 4.3195 | 7.253 | 18.2865 | 137.232 | 0.23011 | 0.094989 | -115
51 | 4 | 4.3477 | 8.305 | 21.5718 | 132.9817 | 0.27489 | 0.081275 | -20 | 65 | 4 | 4.339 | 7 | 18.3131 | 138.269 | 0.17092 | 0.11894 | -115
52 | 1 | 4.5763 | 6.233 | 15.0751 | 133.5736 | 0.1961 | 0.19304 | -25 | 71 | 1 | 3.8782 | 7.697 | 14.9992 | 138.3575 | 0.23091 | 0.1122 | -115
52 | 2 | 4.9303 | 6.001 | 15.7038 | 134.9386 | 0.29919 | 0.19632 | -25 | 71 | 2 | 6.6327 | 6.764 | 15.5469 | 139.8627 | 1.1482 | 0.18825 | -115
52 | 3 | 4.7876 | 5.906 | 16.2294 | 131.4823 | 0.31429 | 0.15414 | -25 | 71 | 3 | 7.5863 | 6.258 | 16.0612 | 136.3899 | 0.66338 | 0.12723 | -115
52 | 4 | 4.9122 | 5.81 | 16.2141 | 133.2142 | 0.16532 | 0.11297 | -25 | 71 | 4 | 5.0339 | 5.928 | 16.0728 | 138.2489 | 0.2275 | 0.091336 | -115
53 | 1 | 4.0823 | 6.607 | 15.8573 | 137.1163 | 0.083749 | 0.052914 | 0 | 37 | 1 | 7.8293 | 5.462 | 16.6566 | 136.4061 | 1.2269 | 0.2823 | -120
53 | 2 | 4.6311 | 6.162 | 16.0343 | 137.6158 | 0.12757 | 0.05076 | 0 | 37 | 2 | 6.2735 | 5.671 | 17.5517 | 138.0005 | 0.54635 | 0.17752 | -120
53 | 3 | 4.6096 | 5.953 | 16.9344 | 134.6688 | 0.16021 | 0.046954 | 0 | 37 | 3 | 5.1651 | 5.828 | 17.8275 | 134.2455 | 0.24103 | 0.14896 | -120
53 | 4 | 4.7127 | 5.789 | 16.4421 | 136.2628 | 0.11974 | 0.047679 | 0 | 37 | 4 | 7.3313 | 5.886 | 17.9432 | 135.8938 | 1.3172 | 0.46034 | -120
54 | 1 | 4.4005 | 8.641 | 21.7986 | 134.2117 | 0.4271 | 0.066125 | -20 | 77 | 1 | 5.4823 | 6.109 | 16.3373 | 130.5481 | 0.49054 | 0.1235 | -120
54 | 2 | 4.4849 | 8.259 | 21.8343 | 134.7858 | 0.32557 | 0.065177 | -20 | 77 | 2 | 7.2097 | 6.079 | 16.6445 | 132.3669 | 1.1997 | 0.22414 | -120
54 | 3 | 4.5748 | 8.149 | 22.5552 | 131.1404 | 0.27593 | 0.053753 | -20 | 77 | 3 | 6.3728 | 6.124 | 17.2263 | 129.221 | 0.97751 | 0.19652 | -120
54 | 4 | 4.6361 | 7.985 | 22.315 | 132.7387 | 0.49101 | 0.069997 | -20 | 77 | 4 | 5.5284 | 6.098 | 17.1169 | 130.6257 | 0.35951 | 0.11329 | -120
55 | 1 | 4.3581 | 6.379 | 15.904 | 137.5466 | 0.12549 | 0.086451 | -5 | 40 | 1 | 6.0903 | 5.948 | 14.6648 | 138.3103 | 1.2692 | 0.1693 | -205
55 | 2 | 4.5489 | 6.047 | 16.3815 | 136.8164 | 0.29382 | 0.072799 | -5 | 40 | 2 | 6.7014 | 5.805 | 15.2707 | 139.3038 | 0.92774 | 0.19373 | -205
55 | 3 | 4.6357 | 5.911 | 16.9152 | 134.4894 | 0.087812 | 0.072984 | -5 | 40 | 3 | 6.4497 | 5.738 | 15.9797 | 135.7404 | 0.66131 | 0.17007 | -205
55 | 4 | 4.7624 | 5.812 | 16.7028 | 136.0348 | 0.12087 | 0.064085 | -5 | 40 | 4 | 6.6641 | 5.675 | 15.9008 | 138.0984 | 1.5561 | 0.16188 | -205
56 | 1 | 3.6812 | 8.891 | 20.7197 | 135.4234 | 0.27369 | 0.13609 | -20 | 25 | 1 | 8.8528 | 3.841 | 18.3242 | 136.5791 | 1.2096 | 0.19883 | -210
56 | 2 | 3.9608 | 8.395 | 20.9511 | 136.1513 | 0.28828 | 0.13457 | -20 | 25 | 2 | 9.3754 | 3.681 | 19.1377 | 137.6749 | 1.3275 | 0.20741 | -210
56 | 3 | 4.083 | 8.237 | 21.9806 | 132.5639 | 0.34804 | 0.12097 | -20 | 25 | 3 | 9.6491 | 3.649 | 19.9622 | 134.0737 | 1.6802 | 0.1777 | -210
56 | 4 | 4.1661 | 8.022 | 21.2761 | 134.0929 | 0.42166 | 0.13136 | -20 | 25 | 4 | 9.7724 | 3.602 | 19.6668 | 136.7933 | 1.3834 | 0.17446 | -210
57 | 1 | 3.9706 | 6.962 | 15.9701 | 140.6681 | 0.056766 | 0.15284 | -5 | 48 | 1 | 7.4528 | 5.452 | 21.0686 | 134.8884 | 0.97664 | 0.1785 | -210
57 | 2 | 4.2632 | 6.54 | 16.6977 | 141.1354 | 0.086741 | 0.14869 | -5 | 48 | 2 | 9.4607 | 5.201 | 21.7403 | 136.6733 | 1.867 | 0.46916 | -210
57 | 3 | 4.3743 | 6.357 | 16.9776 | 138.239 | 0.080773 | 0.14801 | -5 | 48 | 3 | 9.4489 | 5.118 | 22.0618 | 132.6916 | 1.7566 | 0.43536 | -210
57 | 4 | 4.4811 | 6.208 | 17.0336 | 140.393 | 0.21081 | 0.14718 | -5 | 48 | 4 | 8.0145 | 5.075 | 22.0656 | 134.7525 | 0.98832 | 0.17806 | -210
58 | 1 | 4.1349 | 6.629 | 15.223 | 140.6505 | 0.078072 | 0.1534 | 0 | 31 | 1 | 5.4866 | 7.186 | 15.5689 | 131.3756 | 2.5639 | 0.20361 | -220
58 | 2 | 4.2849 | 6.392 | 15.811 | 141.8353 | 0.063876 | 0.15512 | 0 | 31 | 2 | 5.9163 | 6.586 | 16.2766 | 133.0441 | 0.89315 | 0.1878 | -220
58 | 3 | 4.3675 | 6.277 | 16.155 | 138.1437 | 0.06417 | 0.14889 | 0 | 31 | 3 | 8.3318 | 6.267 | 17.3021 | 129.3452 | 1.1746 | 0.26325 | -220
58 | 4 | 4.4659 | 6.169 | 16.1718 | 140.5555 | 0.078662 | 0.14289 | 0 | 31 | 4 | 6.4477 | 6.044 | 17.0102 | 131.473 | 1.1182 | 0.21189 | -220
59 | 1 | 45.4283 | 1.425 | 16.052 | 140.788 | 20.0502 | 1.9895 | -420 | 42 | 1 | 10.6657 | 4.204 | 18.2912 | 138.4385 | 10.725 | 0.52716 | -225
59 | 2 | 43.0272 | 1.323 | 16.8155 | 141.2752 | 28.5008 | 3.3122 | -420 | 42 | 2 | 10.842 | 4.103 | 19.1062 | 140.4393 | 6.7298 | 0.033971 | -225
59 | 3 | 45.4846 | 1.266 | 17.3631 | 139.5389 | 31.0685 | 1.997 | -420 | 42 | 3 | 71.3285 | 3.974 | 19.2264 | 136.6253 | 0.0013553 | 0.00055236 | -225
59 | 4 | 46.6555 | 1.235 | 17.3874 | 140.0931 | 24.7329 | 1.7463 | -420 | 42 | 4 | 11.3854 | 3.983 | 19.1359 | 138.0956 | 0.0013281 | 0.00055707 | -225
60 | 1 | 79.8456 | 0.68523 | 18.3477 | 136.0034 | 18.2453 | 1.8081 | -420 | 49 | 1 | 13.7778 | 3.083 | 19 | 135.7228 | 2.2641 | 0.18902 | -240
60 | 2 | 90.3259 | 0.62108 | 19.0987 | 136.9436 | 22.9096 | 2.2349 | -420 | 49 | 2 | 15.0417 | 2.892 | 19.6218 | 136.4361 | 2.3796 | 0.31337 | -240
60 | 3 | 93.9439 | 0.59902 | 19.4715 | 134.1994 | 25.4307 | 2.2114 | -420 | 49 | 3 | 14.4729 | 2.773 | 20.404 | 133.4906 | 2.3412 | 0.18614 | -240
60 | 4 | 92.9587 | 0.59667 | 19.556 | 134.7638 | 24.4731 | 2.2219 | -420 | 49 | 4 | 15.4296 | 2.693 | 20.3135 | 134.1614 | 2.3943 | 0.24154 | -240
61 | 1 | 42.1228 | 1.481 | 14.7278 | 135.6401 | 4.5554 | 0.48634 | -330 | 5 | 1 | 5.8964 | 10.963 | 23.1291 | 150.2584 | 0.75743 | 0.12297 | -255
61 | 2 | 44.9723 | 1.399 | 15.5255 | 135.5834 | 4.7504 | 0.51771 | -330 | 5 | 2 | 7.6309 | 10.362 | 23.8739 | 151.3545 | 6.4716 | 0.95856 | -255
61 | 3 | 47.2817 | 1.35 | 16.0472 | 133.2704 | 5.4953 | 0.58888 | -330 | 5 | 3 | 6.6126 | 10.297 | 24.4115 | 148.6752 | 2.3518 | 0.41941 | -255
61 | 4 | 47.35 | 1.323 | 15.9562 | 133.6879 | 4.9561 | 0.57159 | -330 | 5 | 4 | 5.7057 | 9.992 | 24.2674 | 149.9318 | 1.4056 | 0.1074 | -255
62 | 1 | 6.1511 | 6.279 | 15.2205 | 135.5888 | 0.51916 | 0.17436 | -65 | 16 | 1 | 3.443 | 35.8 | 34.0227 | 135.0245 | 2.3294 | 0.41695 | -265
62 | 2 | 6.5086 | 5.94 | 15.6917 | 137.7057 | 0.55621 | 0.17496 | -65 | 16 | 2 | 5.3025 | 34.179 | 34.2454 | 136.0903 | 3.4345 | 0.68954 | -265
62 | 3 | 6.5876 | 5.761 | 16.3065 | 133.9438 | 0.4984 | 0.16223 | -65 | 16 | 3 | 3.2068 | 37.267 | 35.5304 | 133.0175 | 2.6765 | 0.31583 | -265
62 | 4 | 6.763 | 5.599 | 16.3874 | 135.2198 | 0.71377 | 0.1693 | -65 | 16 | 4 | 2.451 | 39.913 | 34.9018 | 134.7898 | 1.4949 | 0.35127 | -265
63 | 1 | 4.1193 | 7.468 | 14.7098 | 135.9938 | 0.26413 | 0.049805 | -20 | 19 | 1 | 3.2137 | 45.029 | 28.997 | 136.7275 | 2.6806 | 0.59035 | -265
63 | 2 | 4.5037 | 6.657 | 15.5757 | 136.7661 | 0.24555 | 0.054207 | -20 | 19 | 2 | 2.6059 | 45.145 | 30.0515 | 138.1055 | 1.4117 | 0.32691 | -265
63 | 3 | 5.1031 | 6.232 | 16.1916 | 134.0473 | 0.32439 | 0.049953 | -20 | 19 | 3 | 3.4361 | 41.101 | 30.8265 | 135.3246 | 1.9771 | 0.35412 | -265
63 | 4 | 5.1496 | 5.929 | 15.9425 | 135.229 | 0.2165 | 0.047817 | -20 | 19 | 4 | 2.5174 | 44.225 | 30.4923 | 137.6236 | 1.6784 | 0.30912 | -265
64 | 1 | 3.7505 | 7.547 | 15.4255 | 140.247 | 0.10175 | 0.10201 | -50 | 61 | 1 | 42.1228 | 1.481 | 14.7278 | 135.6401 | 4.5554 | 0.48634 | -330
64 | 2 | 5.0124 | 6.849 | 16.147 | 141.2571 | 0.56052 | 0.14943 | -50 | 61 | 2 | 44.9723 | 1.399 | 15.5255 | 135.5834 | 4.7504 | 0.51771 | -330
64 | 3 | 4.4288 | 6.468 | 16.5524 | 138.3338 | 0.17999 | 0.089225 | -50 | 61 | 3 | 47.2817 | 1.35 | 16.0472 | 133.2704 | 5.4953 | 0.58888 | -330
64 | 4 | 4.6204 | 6.191 | 16.3856 | 140.6044 | 0.15862 | 0.065673 | -50 | 61 | 4 | 47.35 | 1.323 | 15.9562 | 133.6879 | 4.9561 | 0.57159 | -330
65 | 1 | 5.0494 | 8.414 | 17.2703 | 139.3918 | 1.399 | 0.23017 | -115 | 67 | 1 | 17.3539 | 2.523 | 16.299 | 139.3857 | 3.5347 | 0.32706 | -330
65 | 2 | 4.3911 | 7.577 | 17.7231 | 140.7551 | 0.66485 | 0.33642 | -115 | 67 | 2 | 18.4829 | 2.375 | 16.8969 | 139.179 | 4.0895 | 0.34038 | -330
65 | 3 | 4.3195 | 7.253 | 18.2865 | 137.232 | 0.23011 | 0.094989 | -115 | 67 | 3 | 19.9211 | 2.275 | 17.6854 | 137.2257 | 58.3998 | 1.1749 | -330
65 | 4 | 4.339 | 7 | 18.3131 | 138.269 | 0.17092 | 0.11894 | -115 | 67 | 4 | 19.9273 | 2.225 | 17.5417 | 137.9622 | 15.3038 | 0.73489 | -330
66 | 1 | 3.3666 | 8.667 | 18.2721 | 135.7426 | 0.11995 | 0.047502 | -55 | 27 | 1 | 52.4652 | 1.264 | 18.6371 | 135.0774 | 10.7771 | 0.92376 | -420
66 | 2 | 3.7425 | 7.907 | 18.6979 | 137.1581 | 0.3192 | 0.078922 | -55 | 27 | 2 | 53.5538 | 1.245 | 19.0932 | 136.1872 | 9.9415 | 0.96719 | -420
66 | 3 | 3.8249 | 7.549 | 19.4347 | 133.6664 | 0.10182 | 0.05751 | -55 | 27 | 3 | 54.7801 | 1.235 | 19.4993 | 132.7102 | 12.3277 | 1.0313 | -420
66 | 4 | 3.9944 | 7.234 | 19.3377 | 134.9237 | 2.5897 | 0.066221 | -55 | 27 | 4 | 53.7946 | 1.235 | 19.709 | 134.0705 | 11.6977 | 0.97275 | -420
67 | 1 | 17.3539 | 2.523 | 16.299 | 139.3857 | 3.5347 | 0.32706 | -330 | 33 | 1 | 17.8367 | 2.707 | 15.0923 | 142.9027 | 28.5309 | 1.3066 | -420
67 | 2 | 18.4829 | 2.375 | 16.8969 | 139.179 | 4.0895 | 0.34038 | -330 | 33 | 2 | 18.9232 | 2.615 | 15.8969 | 142.7054 | 25.9452 | 1.2978 | -420
67 | 3 | 19.9211 | 2.275 | 17.6854 | 137.2257 | 58.3998 | 1.1749 | -330 | 33 | 3 | 18.7826 | 2.549 | 16.1686 | 141.2023 | 22.1601 | 1.2789 | -420
67 | 4 | 19.9273 | 2.225 | 17.5417 | 137.9622 | 15.3038 | 0.73489 | -330 | 33 | 4 | 19.1805 | 2.528 | 16.144 | 142.0302 | 22.8827 | 1.6687 | -420
68 | 1 | 3.6384 | 9.112 | 20.1902 | 135.3932 | 0.21254 | 0.053575 | -20 | 35 | 1 | 24.7129 | 2.053 | 18.5002 | 137.881 | 10.0424 | 0.70603 | -420
68 | 2 | 4.1316 | 8.688 | 20.9773 | 137.0335 | 0.34255 | 0.066539 | -20 | 35 | 2 | 25.8295 | 2.006 | 18.9138 | 138.3494 | 11.4574 | 0.82512 | -420
68 | 3 | 4.4471 | 8.548 | 21.8351 | 133.4975 | 0.49484 | 0.068611 | -20 | 35 | 3 | 26.0365 | 1.977 | 19.4429 | 135.6102 | 12.7158 | 0.82719 | -420
68 | 4 | 4.0205 | 8.333 | 21.4758 | 135.0656 | 0.4098 | 0.061784 | -20 | 35 | 4 | 25.8404 | 1.977 | 19.4773 | 136.2931 | 12.0441 | 0.80717 | -420
69 | 1 | [7 empty] | 41 | 1 | 31.1935 | 1.665 | 15.4732 | 143.5999 | 14.0179 | 1.2012 | -420
69 | 2 | [7 empty] | 41 | 2 | 33.0587 | 1.589 | 16.3793 | 143.6402 | 15.06 | 1.3583 | -420
69 | 3 | [7 empty] | 41 | 3 | 34.1801 | 1.546 | 16.876 | 141.4274 | 13.6898 | 1.2979 | -420
69 | 4 | [7 empty] | 41 | 4 | 34.5848 | 1.521 | 17.0127 | 142.7072 | 13.7172 | 1.2531 | -420
70 | 1 | 6.3457 | 6.466 | 15.5627 | 134.6586 | 0.48311 | 0.15769 | -65 | 59 | 1 | 45.4283 | 1.425 | 16.052 | 140.788 | 20.0502 | 1.9895 | -420
70 | 2 | 6.5838 | 5.956 | 16.0457 | 136.9131 | 0.47355 | 0.15601 | -65 | 59 | 2 | 43.0272 | 1.323 | 16.8155 | 141.2752 | 28.5008 | 3.3122 | -420
70 | 3 | 6.8334 | 5.694 | 16.5534 | 133.516 | 0.60152 | 0.15189 | -65 | 59 | 3 | 45.4846 | 1.266 | 17.3631 | 139.5389 | 31.0685 | 1.997 | -420
70 | 4 | 7.1662 | 5.504 | 16.4696 | 134.4972 | 0.53499 | 0.14577 | -65 | 59 | 4 | 46.6555 | 1.235 | 17.3874 | 140.0931 | 24.7329 | 1.7463 | -420
71 | 1 | 3.8782 | 7.697 | 14.9992 | 138.3575 | 0.23091 | 0.1122 | -115 | 60 | 1 | 79.8456 | 0.68523 | 18.3477 | 136.0034 | 18.2453 | 1.8081 | -420
71 | 2 | 6.6327 | 6.764 | 15.5469 | 139.8627 | 1.1482 | 0.18825 | -115 | 60 | 2 | 90.3259 | 0.62108 | 19.0987 | 136.9436 | 22.9096 | 2.2349 | -420
71 | 3 | 7.5863 | 6.258 | 16.0612 | 136.3899 | 0.66338 | 0.12723 | -115 | 60 | 3 | 93.9439 | 0.59902 | 19.4715 | 134.1994 | 25.4307 | 2.2114 | -420
71 | 4 | 5.0339 | 5.928 | 16.0728 | 138.2489 | 0.2275 | 0.091336 | -115 | 60 | 4 | 92.9587 | 0.59667 | 19.556 | 134.7638 | 24.4731 | 2.2219 | -420
72 | 1 | 6.4166 | 5.608 | 15.5826 | 135.3885 | 0.53342 | 0.15178 | -110 | 78 | 1 | 69.0331 | 0.92459 | 16.8288 | 139.5139 | 37.2686 | 3.6332 | -420
72 | 2 | 6.6362 | 5.639 | 16.111 | 137.251 | 0.5198 | 0.14384 | -110 | 78 | 2 | 72.283 | 0.87545 | 17.884 | 140.2162 | 61.6989 | 2.9037 | -420
72 | 3 | 6.5246 | 5.677 | 16.4193 | 133.8186 | 0.72032 | 0.14365 | -110 | 78 | 3 | 78.0929 | 0.84312 | 18.5452 | 137.9462 | 47.3977 | 3.4875 | -420
72 | 4 | 6.4386 | 5.67 | 16.5582 | 135.0145 | 1.2372 | 0.14096 | -110 | 78 | 4 | 80.4463 | 0.81261 | 18.0788 | 139.2344 | 55.9581 | 3.6857 | -420
73 | 1 | 4.805 | 6.112 | 14.2394 | 140.2068 | 0.25846 | 0.051337 | -20 | 23 | 1 | 104.886 | 0.52482 | 19.2204 | 134.4593 | 5.4375 | 0.65521 | -600
73 | 2 | 4.8188 | 5.889 | 14.964 | 140.2616 | 0.18677 | 0.046877 | -20 | 23 | 2 | 100.473 | 0.53754 | 20.0813 | 134.9474 | 5.9365 | 0.58716 | -600
73 | 3 | 4.9441 | 5.803 | 15.474 | 137.3414 | 0.20757 | 0.046304 | -20 | 23 | 3 | 110.61 | 0.53104 | 20.7149 | 132.2622 | 5.6007 | 0.6133 | -600
73 | 4 | 5.1512 | 5.713 | 15.3533 | 139.0028 | 0.26927 | 0.049606 | -20 | 23 | 4 | 106.631 | 0.51791 | 20.4626 | 134.1588 | 5.574 | 0.64177 | -600
74 | 1 | 4.9454 | 6.587 | 15.6743 | 139.4919 | 0.36181 | 0.063451 | -65 | 3 | 1
74 | 2 | 5.8553 | 6.273 | 16.2733 | 140.5164 | 1.2248 | 0.093917 | -65 | 3 | 2
74 | 3 | 5.0311 | 6.142 | 16.804 | 137.1043 | 0.30254 | 0.060743 | -65 | 3 | 3
74 | 4 | 5.2191 | 6.013 | 16.8859 | 138.5953 | 0.34105 | 0.066814 | -65 | 3 | 4
75 | 1 | 5.5826 | 5.81 | 14.7931 | 137.3327 | 0.31636 | 0.068775 | -20 | 11 | 1
75 | 2 | 5.6538 | 5.637 | 15.6247 | 137.5287 | 0.33279 | 0.073566 | -20 | 11 | 2
75 | 3 | 5.8325 | 5.553 | 16.4179 | 133.9771 | 0.3479 | 0.06855 | -20 | 11 | 3
75 | 4 | 6.0911 | 5.461 | 16.2948 | 136.2947 | 0.38618 | 0.079298 | -20 | 11 | 4
76 | 1 | 4.2483 | 6.407 | 14.6811 | 138.2108 | 0.10253 | 0.079871 | 0 | 12 | 1
76 | 2 | 4.4783 | 6.082 | 15.3733 | 139.3022 | 0.11543 | 0.077495 | 0 | 12 | 2
76 | 3 | 4.6108 | 5.942 | 16.0342 | 136.0342 | 0.15785 | 0.09159 | 0 | 12 | 3
76 | 4 | 4.683 | 5.825 | 15.8112 | 138.0157 | 0.084422 | 0.081706 | 0 | 12 | 4
77 | 1 | 5.4823 | 6.109 | 16.3373 | 130.5481 | 0.49054 | 0.1235 | -120 | 13 | 1
77 | 2 | 7.2097 | 6.079 | 16.6445 | 132.3669 | 1.1997 | 0.22414 | -120 | 13 | 2
77 | 3 | 6.3728 | 6.124 | 17.2263 | 129.221 | 0.97751 | 0.19652 | -120 | 13 | 3
77 | 4 | 5.5284 | 6.098 | 17.1169 | 130.6257 | 0.35951 | 0.11329 | -120 | 13 | 4
78 | 1 | 69.0331 | 0.92459 | 16.8288 | 139.5139 | 37.2686 | 3.6332 | -420 | 14 | 1
78 | 2 | 72.283 | 0.87545 | 17.884 | 140.2162 | 61.6989 | 2.9037 | -420 | 14 | 2
78 | 3 | 78.0929 | 0.84312 | 18.5452 | 137.9462 | 47.3977 | 3.4875 | -420 | 14 | 3
78 | 4 | 80.4463 | 0.81261 | 18.0788 | 139.2344 | 55.9581 | 3.6857 | -420 | 14 | 4
79 | 1 | 5.9353 | 6.181 | 15.2847 | 134.752 | 0.45654 | 0.16493 | -110 | 15 | 1
79 | 2 | 6.0295 | 5.887 | 15.8756 | 135.6469 | 0.88676 | 0.16973 | -110 | 15 | 2
79 | 3 | 6.3624 | 5.755 | 16.4741 | 132.469 | 0.43876 | 0.15439 | -110 | 15 | 3
79 | 4 | 6.8387 | 5.656 | 16.1881 | 134.606 | 0.56573 | 0.15862 | -110 | 15 | 4
80 | 1 | 3.9555 | 6.753 | 14.5603 | 136.8074 | 0.075366 | 0.16735 | -50 | 69 | 1
80 | 2 | 4.2582 | 6.297 | 15.1378 | 138.2429 | 0.092188 | 0.11045 | -50 |  | 2
80 | 3 | 4.4427 | 6.065 | 15.8577 | 134.9259 | 0.84258 | 0.10751 | -50 |  | 3
80 | 4 | 4.5706 | 5.9 | 15.7672 | 136.2188 | 0.054267 | 0.11388 | -50 |  | 4