## Key
Fully Chacterised Data Key:-
PY = | Sensor & actuator coil Parameters measured at CIT, Yes.
PN = | Sensor & actuator coil Parameters measured at CIT, No.
NY = | Sensor Noise mesaured at LHO, Yes.
NN = | Sensor Noise mesaured at LHO, No.
Key Colour Scheme:-
Green = | Go for installation.
Amber = | Incomplete AOSEM Test Data, proceed with caution.
Red = | No AOSEM test data, stop and replace unit.
OSEM Open-light goal:- |  | 25k - 31k counts

## Quad
Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain | Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain | Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain
L1-ITMX | M0 - BOSEM | Face 1 (F1) | 676 | X | 39.6 |  | 29280 | -14640 | 1.0246 | H1-ITMX | M0 - BOSEM | Face 1 (F1) | 91 |  |  |  | 22745 | -11372.5 | 1.319 | I1-ITMX | M0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) | 519 |  | 39.8 |  | 31270 | -15635 | 0.95939 |  |  | Face 2 (F2) | 494 |  |  |  | 25209 | -12604.5 | 1.1901 |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) | 334 |  |  |  | 27314 | -13657 | 1.0983 |  |  | Face 3 (F3) | 482 |  |  |  | 25497 | -12748.5 | 1.1766 |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | 450 |  | 39.7 |  | 27503 | -13751.5 | 1.0908 |  |  | Left (LF) | 509 |  |  |  | 25276 | -12638 | 1.1869 |  |  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | 524 |  | 39.4 |  | 29711 | -14855.5 | 1.0097 |  |  | Right (RT) | 490 |  |  |  | 24820 | -12410 | 1.2087 |  |  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | 449 |  | 39.4 |  | 28832 | -14416 | 1.0405 |  |  | Side (SD) | 111 |  |  |  | 23964 | -11982 | 1.2519 |  |  | Side (SD) | [5 empty] | 0 | #DIV/0!
| R0 - BOSEM | Face 1 (F1) | 353 |  |  |  | 25870 | -12935 | 1.1596 |  | R0 - BOSEM | Face 1 (F1) | 470 |  |  |  | 25057 | -12528.5 | 1.1973 |  | R0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) | 514 |  | 39.7 |  | 29718 | -14859 | 1.0095 |  |  | Face 2 (F2) | 467 |  |  |  | 30488 | -15244 | 0.98399 |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) | 516 |  | 39.7 |  | 26149 | -13074.5 | 1.1473 |  |  | Face 3 (F3) | 127 |  |  |  | 29015 | -14507.5 | 1.0339 |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | 515 |  | 39.6 |  | 29658 | -14829 | 1.0115 |  |  | Left (LF) | 115 |  |  |  | 29178 | -14589 | 1.0282 |  |  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | 523 |  | 39.3 |  | 27752 | -13876 | 1.081 |  |  | Right (RT) | 664 |  |  |  | 27580 | -13790 | 1.0877 |  |  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | 525 |  | 39.5 |  | 29986 | -14993 | 1.0005 |  |  | Side (SD) | 80 |  |  |  | 27366 | -13683 | 1.0963 |  |  | Side (SD) | [5 empty] | 0 | #DIV/0!
| L1 - BOSEM | Upper Left (UL) |  |  |  |  | 25604 | -12802 | 1.1717 |  | L1 - BOSEM | Upper Left (UL) | 439 |  |  |  | 31509 | -15754.5 | 0.95211 |  | L1 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) |  |  |  |  | 23046 | -11523 | 1.3017 |  |  | Lower Left (LL) | 510 |  |  |  | 29214 | -14607 | 1.0269 |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) |  |  |  |  | 30067 | -15033.5 | 0.99777 |  |  | Upper Right (UR) | 474 |  |  |  | 26506 | -13253 | 1.1318 |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) |  |  |  |  | 29682 | -14841 | 1.0107 |  |  | Lower Right (LR) | 487 |  |  |  | 27979 | -13989.5 | 1.0722 |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
| L2 - AOSEM | Upper Left (UL) |  |  |  |  | 27122 | -13561 | 1.1061 |  | L2 - AOSEM | Upper Left (UL) | 146 |  |  |  | 21877 | -10938.5 | 1.3713 |  | L2 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) |  |  |  |  | 26903 | -13451.5 | 1.1151 |  |  | Lower Left (LL) | 263 |  |  |  | 22829 | -11414.5 | 1.3141 |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) |  |  |  |  | 28398 | -14199 | 1.0564 |  |  | Upper Right (UR) | 112 |  |  |  | 24994 | -12497 | 1.2003 |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) |  |  |  |  | 27045 | -13522.5 | 1.1093 |  |  | Lower Right (LR) | 463 |  |  |  | 22583 | -11291.5 | 1.3284 |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
L1-ITMY | M0 - BOSEM | Face 1 (F1) | 348 |  |  |  | 28880 | -14440 | 1.0388 | H1-ITMY | M0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0! | I1-ITMY | M0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) | 332 |  |  |  | 27438 | -13719 | 1.0934 |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0! |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) | 461 |  |  |  | 25638 | -12819 | 1.1701 |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0! |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | 323 |  |  |  | 27760 | -13880 | 1.0807 |  |  | Left (LF) | [5 empty] | 0 | #DIV/0! |  |  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | 365 |  |  |  | 30382 | -15191 | 0.98743 |  |  | Right (RT) | [5 empty] | 0 | #DIV/0! |  |  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | 322 |  |  |  | 28388 | -14194 | 1.0568 |  |  | Side (SD) | [5 empty] | 0 | #DIV/0! |  |  | Side (SD) | [5 empty] | 0 | #DIV/0!
| R0 - BOSEM | Face 1 (F1) | 316 |  |  |  | 27656 | -13828 | 1.0848 |  | R0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0! |  | R0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) | 319 |  |  |  | 27566 | -13783 | 1.0883 |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0! |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) | 360 |  |  |  | 28958 | -14479 | 1.036 |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0! |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | 364 |  |  |  | 27690 | -13845 | 1.0834 |  |  | Left (LF) | [5 empty] | 0 | #DIV/0! |  |  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | 359 |  |  |  | 27514 | -13757 | 1.0904 |  |  | Right (RT) | [5 empty] | 0 | #DIV/0! |  |  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | 600 | X |  |  | 28626 | -14313 | 1.048 |  |  | Side (SD) | [5 empty] | 0 | #DIV/0! |  |  | Side (SD) | [5 empty] | 0 | #DIV/0!
| L1 - BOSEM | Upper Left (UL) | 647 |  |  |  | 30410 | -15205 | 0.98652 |  | L1 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! |  | L1 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) | 648 |  |  |  | 28688 | -14344 | 1.0457 |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) | 633 | X |  |  | 32401 | -16200.5 | 0.9259 |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) | 649 |  |  |  | 30792 | -15396 | 0.97428 |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
| L2 - AOSEM | Upper Left (UL) | 402 |  |  |  | 24772 | -12386 | 1.211 |  | L2 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! |  | L2 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) | 235 |  |  |  | 25682 | -12841 | 1.1681 |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) | 348 |  |  |  | 27435 | -13717.5 | 1.0935 |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) | 364 |  |  |  | 30120 | -15060 | 0.99602 |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
L1-ETMX | M0 - BOSEM | Face 1 (F1) |  |  |  |  | 24997 | -12498.5 | 1.2001 | H1-ETMX | M0 - BOSEM | Face 1 (F1) | 643 |  |  |  | 24018 | -12009 | 1.2491 | I1-ETMX | M0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) |  |  |  |  | 30063 | -15031.5 | 0.9979 |  |  | Face 2 (F2) | 486 |  |  |  | 29453 | -14726.5 | 1.0186 |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) |  |  |  |  | 24924 | -12462 | 1.2037 |  |  | Face 3 (F3) | 414 |  |  |  | 30661 | -15330.5 | 0.97844 |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) |  |  |  |  | 24635 | -12317.5 | 1.2178 |  |  | Left (LF) | 580 |  |  |  | 24512 | -12256 | 1.2239 |  |  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) |  |  |  |  | 24346 | -12173 | 1.2322 |  |  | Right (RT) | 507 |  |  |  | 24724 | -12362 | 1.2134 |  |  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) |  |  |  |  | 30807 | -15403.5 | 0.9738 |  |  | Side (SD) | 563 |  |  |  | 28519 | -14259.5 | 1.0519 |  |  | Side (SD) | [5 empty] | 0 | #DIV/0!
| R0 - BOSEM | Face 1 (F1) |  |  |  |  | 25878 | -12939 | 1.1593 |  | R0 - BOSEM | Face 1 (F1) | 426 |  |  |  | 28922 | -14461 | 1.0373 |  | R0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) |  |  |  |  | 30505 | -15252.5 | 0.98345 |  |  | Face 2 (F2) | 497 |  |  |  | 23013 | -11506.5 | 1.3036 |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) |  |  |  |  | 28736 | -14368 | 1.044 |  |  | Face 3 (F3) | 504 |  |  |  | 25716 | -12858 | 1.1666 |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) |  |  |  |  | 24796 | -12398 | 1.2099 |  |  | Left (LF) | 468 |  |  |  | 26662 | -13331 | 1.1252 |  |  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) |  |  |  |  | 25315 | -12657.5 | 1.1851 |  |  | Right (RT) | 418 |  |  |  | 24388 | -12194 | 1.2301 |  |  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) |  |  |  |  | 25643 | -12821.5 | 1.1699 |  |  | Side (SD) | 434 |  |  |  | 21961 | -10980.5 | 1.3661 |  |  | Side (SD) | [5 empty] | 0 | #DIV/0!
| L1 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! |  | L1 - BOSEM | Upper Left (UL) |  |  |  |  | 24267 | -12133.5 | 1.2362 |  | L1 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! |  |  | Lower Left (LL) |  |  |  |  | 26538 | -13269 | 1.1305 |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! |  |  | Upper Right (UR) |  |  |  |  | 24545 | -12272.5 | 1.2222 |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! |  |  | Lower Right (LR) |  |  |  |  | 26259 | -13129.5 | 1.1425 |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
| L2 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! |  | L2 - AOSEM | Upper Left (UL) |  |  |  |  | 17935 | -8967.5 | 1.6727 |  | L2 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! |  |  | Lower Left (LL) |  |  |  |  | 18726 | -9363 | 1.6021 |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! |  |  | Upper Right (UR) |  |  |  |  | 25124 | -12562 | 1.1941 |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! |  |  | Lower Right (LR) |  |  |  |  | 25518 | -12759 | 1.1756 |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
L1-ETMY | M0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0! | H1-ETMY | M0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0! | I1-ETMY | M0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) | [5 empty] | 0 | #DIV/0! |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0! |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) | [5 empty] | 0 | #DIV/0! |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0! |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0! |  |  | Left (LF) | [5 empty] | 0 | #DIV/0! |  |  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0! |  |  | Right (RT) | [5 empty] | 0 | #DIV/0! |  |  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0! |  |  | Side (SD) | [5 empty] | 0 | #DIV/0! |  |  | Side (SD) | [5 empty] | 0 | #DIV/0!
| R0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0! |  | R0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0! |  | R0 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) | [5 empty] | 0 | #DIV/0! |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0! |  |  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) | [5 empty] | 0 | #DIV/0! |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0! |  |  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0! |  |  | Left (LF) | [5 empty] | 0 | #DIV/0! |  |  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0! |  |  | Right (RT) | [5 empty] | 0 | #DIV/0! |  |  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0! |  |  | Side (SD) | [5 empty] | 0 | #DIV/0! |  |  | Side (SD) | [5 empty] | 0 | #DIV/0!
| L1 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! |  | L1 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! |  | L1 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
| L2 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! |  | L2 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! |  | L2 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! |  |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! |  |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! |  |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!

## BSFM
Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain
L1-BS | M1 - BOSEM | Face 1 (F1) | 532 |  |  |  | 28394 | -14197 | 1.0566
|  | Face 2 (F2) | 526 |  |  |  | 13656 | -6828 | 2.1968
|  | Face 3 (F3) | 677 | X |  |  | 22373 | -11186.5 | 1.3409
|  | Left (LF) | 529 |  |  |  | 26084 | -13042 | 1.1501
|  | Right (RT) | 531 |  |  |  | 26758 | -13379 | 1.1212
|  | Side (SD) | 452 |  |  |  | 21605 | -10802.5 | 1.3886
| M2 - BOSEM | Upper Left (UL) | 528 |  |  |  | 28739 | -14369.5 | 1.0439
|  | Lower Left (LL) | 530 |  |  |  | 27638 | -13819 | 1.0855
|  | Upper Right (UR) | 533 |  |  |  | 20775 | -10387.5 | 1.444
|  | Lower Right (LR) | 527 |  |  |  | 30590 | -15295 | 0.98071
H1-BS | M1 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!
| M2 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
I1-BS | M1 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!
| M2 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
H2-FMX | M1 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!
| M2 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
H2-FMY | M1 - BOSEM | Face 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Face 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Face 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!
| M2 - BOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
|  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
|  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
|  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!

## HLTS
Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain | Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain | Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain
L1-PR3 | M1 - BOSEM | Top 1 (T1) | 559 |  |  |  | 25061 | -12530.5 | 1.1971 | H1-PR3 | M1 - BOSEM | Top 1 (T1) | 73 |  |  |  | 26501 | -13250.5 | 1.132 | I1-PR3 | M1 - BOSEM | Top 1 (T1) | 42 |  |  |  | 25635.3 | -12817.65 | 1.1703
|  | Top 2 (T2) | 697 | X |  |  | 27906 | -13953 | 1.075 |  |  | Top 2 (T2) | 61 |  |  |  | 30538.6 | -15269.3 | 0.98236 |  |  | Top 2 (T2) | 68 |  |  |  | 25672.9 | -12836.45 | 1.1685
|  | Top 3 (T3) | 377 |  |  |  | 26262 | -13131 | 1.1423 |  |  | Top 3 (T3) | 64 |  |  |  | 29098.6 | -14549.3 | 1.031 |  |  | Top 3 (T3) | 171 |  |  |  | 25420 | -12710 | 1.1802
|  | Left (LF) | 219 |  |  |  | 31090 | -15545 | 0.96494 |  |  | Left (LF) | 69 |  |  |  | 26619.2 | -13309.6 | 1.127 |  |  | Left (LF) | 37 |  |  |  | 26762.8 | -13381.4 | 1.121
|  | Right (RT) | 558 |  |  |  | 29280 | -14640 | 1.0246 |  |  | Right (RT) | 63 |  |  |  | 27050.4 | -13525.2 | 1.109 |  |  | Right (RT) | 495 |  |  |  | 29999.9 | -14999.95 | 1
|  | Side (SD) | 335 |  |  |  | 30034 | -15017 | 0.99887 |  |  | Side (SD) | 700 | X |  |  | 29138.9 | -14569.45 | 1.0296 |  |  | Side (SD) | 686 |  |  |  | 25589.6 | -12794.8 | 1.1724
| M2 - AOSEM | Upper Left (UL) | 368 | PY-NY |  |  | 23916.5 | -11958.25 | 1.2544 |  | M2 - AOSEM | Upper Left (UL) | 488 | PY-NY |  |  | 18318 | -9159 | 1.6377 |  | M2 - AOSEM | Upper Left (UL) | 288 |  |  |  | 28762.3 | -14381.15 | 1.043
|  | Lower Left (LL) | 264 | PY-NY |  |  | 24980.4 | -12490.2 | 1.2009 |  |  | Lower Left (LL) | 601 | PY-NN |  |  | 20793 | -10396.5 | 1.4428 |  |  | Lower Left (LL) | 423 |  |  |  | 26012.5 | -13006.25 | 1.1533
|  | Upper Right (UR) | 370 | PY-NY |  |  | 26189.8 | -13094.9 | 1.1455 |  |  | Upper Right (UR) | 336 | PY-NN |  |  | 19426 | -9713 | 1.5443 |  |  | Upper Right (UR) | 339 |  |  |  | 26204.3 | -13102.15 | 1.1449
|  | Lower Right (LR) | 460 | PY-NY |  |  | 28640.2 | -14320.1 | 1.0475 |  |  | Lower Right (LR) | 413 | PY-NY |  |  | 18210 | -9105 | 1.6474 |  |  | Lower Right (LR) | 345 |  |  |  | 26307.3 | -13153.65 | 1.1404
| M3 - AOSEM | Upper Left (UL) | 277 | PY-NY |  |  | 26785.2 | -13392.6 | 1.12 |  | M3 - AOSEM | Upper Left (UL) | 489 | PY-NY |  |  | 17727 | -8863.5 | 1.6923 |  | M3 - AOSEM | Upper Left (UL) | 303 |  |  |  | 26562.8 | -13281.4 | 1.1294
|  | Lower Left (LL) | 324 | PY-NY |  |  | 26339.8 | -13169.9 | 1.139 |  |  | Lower Left (LL) | 269 | PY-NY |  |  | 23463 | -11731.5 | 1.2786 |  |  | Lower Left (LL) | 128 |  |  |  | 25727.1 | -12863.55 | 1.1661
|  | Upper Right (UR) | 600 | PY-NY |  |  | 27258.5 | -13629.25 | 1.1006 |  |  | Upper Right (UR) | 461 | PY-NY |  |  | 25154 | -12577 | 1.1927 |  |  | Upper Right (UR) | 282 |  |  |  | 27640.6 | -13820.3 | 1.0854
|  | Lower Right (LR) | 302 | PY-NY |  |  | 26080.8 | -13040.4 | 1.1503 |  |  | Lower Right (LR) | 287 | PY-NY |  |  | 17820 | -8910 | 1.6835 |  |  | Lower Right (LR) | 310 |  |  |  | 24742.2 | -12371.1 | 1.2125
L1-SR3 | M1 - BOSEM | Top 1 (T1) | 393 |  |  |  | 31002 | -15501 | 0.96768 | H1-SR3 | M1 - BOSEM | Top 1 (T1) | 55 |  |  |  | 20123.5 | -10061.75 | 1.4908 | I1-SR3 | M1 - BOSEM | Top 1 (T1) | 45 |  |  |  | 28457.2 | -14228.6 | 1.0542
|  | Top 2 (T2) | 590 | X |  |  | 27385 | -13692.5 | 1.0955 |  |  | Top 2 (T2) | 72 |  |  |  | 31413.2 | -15706.6 | 0.95501 |  |  | Top 2 (T2) | 144 |  |  |  | 26461.7 | -13230.85 | 1.1337
|  | Top 3 (T3) | 379 |  |  |  | 30918 | -15459 | 0.97031 |  |  | Top 3 (T3) | 39 |  |  |  | 31241.4 | -15620.7 | 0.96026 |  |  | Top 3 (T3) | 184 |  |  |  | 26536.6 | -13268.3 | 1.1305
|  | Left (LF) | 385 |  |  |  | 26784 | -13392 | 1.1201 |  |  | Left (LF) | 136 |  |  |  | 23018.6 | -11509.3 | 1.3033 |  |  | Left (LF) | 193 |  |  |  | 30762.8 | -15381.4 | 0.9752
|  | Right (RT) | 555 |  |  |  | 28622 | -14311 | 1.0481 |  |  | Right (RT) | 435 |  |  |  | 28071 | -14035.5 | 1.0687 |  |  | Right (RT) | 31 |  |  |  | 30158.3 | -15079.15 | 0.99475
|  | Side (SD) | 222 |  |  |  | 29589 | -14794.5 | 1.0139 |  |  | Side (SD) | 689 | X |  |  | 23469.8 | -11734.9 | 1.2782 |  |  | Side (SD) | 625 | X |  |  | 29402.8 | -14701.4 | 1.0203
| M2 - AOSEM | Upper Left (UL) | 459 | PY-NY |  |  | 24089 | -12044.5 | 1.2454 |  | M2 - AOSEM | Upper Left (UL) | 355 |  |  |  | 25624.4 | -12812.2 | 1.1708 |  | M2 - AOSEM | Upper Left (UL) | 275 |  |  |  | 30007.8 | -15003.9 | 0.99974
|  | Lower Left (LL) | 472 | PY-NY |  |  | 23381 | -11690.5 | 1.2831 |  |  | Lower Left (LL) | 387 |  |  |  | 24713.9 | -12356.95 | 1.2139 |  |  | Lower Left (LL) | 482 |  |  |  | 27113.4 | -13556.7 | 1.1065
|  | Upper Right (UR) | 422 | PY-NY |  |  | 23098 | -11549 | 1.2988 |  |  | Upper Right (UR) | 406 |  |  |  | 27586.9 | -13793.45 | 1.0875 |  |  | Upper Right (UR) | 484 |  |  |  | 26512 | -13256 | 1.1316
|  | Lower Right (LR) | 452 | PY-NY |  |  | 24775 | -12387.5 | 1.2109 |  |  | Lower Right (LR) | 444 |  |  |  | 26252.6 | -13126.3 | 1.1427 |  |  | Lower Right (LR) | 487 |  |  |  | 23963.5 | -11981.75 | 1.2519
| M3 - AOSEM | Upper Left (UL) | 191 | PY-NY |  |  | 23059 | -11529.5 | 1.301 |  | M3 - AOSEM | Upper Left (UL) | 350 |  |  |  | 25911.8 | -12955.9 | 1.1578 |  | M3 - AOSEM | Upper Left (UL) | 322 |  |  |  | 24998.2 | -12499.1 | 1.2001
|  | Lower Left (LL) | 271 | PY-NY |  |  | 23498 | -11749 | 1.2767 |  |  | Lower Left (LL) | 331 |  |  |  | 28339.7 | -14169.85 | 1.0586 |  |  | Lower Left (LL) | 286 |  |  |  | 26802.2 | -13401.1 | 1.1193
|  | Upper Right (UR) | 494 | PY-NY |  |  | 23070 | -11535 | 1.3004 |  |  | Upper Right (UR) | 206 |  |  |  | 25894.1 | -12947.05 | 1.1586 |  |  | Upper Right (UR) | 475 |  |  |  | 25949.6 | -12974.8 | 1.1561
|  | Lower Right (LR) | 416 | PY-NY |  |  | 23393 | -11696.5 | 1.2824 |  |  | Lower Right (LR) | 101 |  |  |  | 26057.9 | -13028.95 | 1.1513 |  |  | Lower Right (LR) | 411 |  |  |  | 25557.4 | -12778.7 | 1.1738

## HSTS
Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. (mH) | Open Light Value [counts] | Offset [counts] | Gain | Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain | Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain
L1-MC1 | M1 - BOSEM | Top 1 (T1) | 208 |  |  |  | 31950 | -15975 | 0.93897 | H1-MC1 | M1 - BOSEM | Top 1 (T1) | 578 |  |  |  | 25960 | -12980 | 1.1556 | I1-MC1 | M1 - BOSEM | Top 1 (T1) | 416 |  |  |  | 29357.8 | -14678.9 | 1.0219
|  | Top 2 (T2) | 592 | X |  |  | 25960 | -12980 | 1.1556 |  |  | Top 2 (T2) | 436 |  |  |  | 29277.9 | -14638.95 | 1.0247 |  |  | Top 2 (T2) | 441 |  |  |  | 25503.7 | -12751.85 | 1.1763
|  | Top 3 (T3) | 376 |  |  |  | 30563 | -15281.5 | 0.98158 |  |  | Top 3 (T3) | 440 |  |  |  | 27858.1 | -13929.05 | 1.0769 |  |  | Top 3 (T3) | 473 |  |  |  | 31260.6 | -15630.3 | 0.95967
|  | Left (LF) | 384 |  |  |  | 27553.3 | -13776.65 | 1.0888 |  |  | Left (LF) | 433 |  |  |  | 27758.7 | -13879.35 | 1.0807 |  |  | Left (LF) | 444 |  |  |  | 31563.2 | -15781.6 | 0.95047
|  | Right (RT) | 402 |  |  |  | 30757.9 | -15378.95 | 0.97536 |  |  | Right (RT) | 424 |  |  |  | 24529.4 | -12264.7 | 1.223 |  |  | Right (RT) | 447 |  |  |  | 31900.5 | -15950.25 | 0.94042
|  | Side (SD) | 535 |  |  |  | 28276 | -14138 | 1.061 |  |  | Side (SD) | 658 | X |  |  | 25139.5 | -12569.75 | 1.1933 |  |  | Side (SD) | 652 | X |  |  | 26982.8 | -13491.4 | 1.1118
| M2 - AOSEM | Upper Left (UL) | 318 | PY-NY |  |  | 27112.1 | -13556.05 | 1.1065 |  | M2 - AOSEM | Upper Left (UL) | 453 | PY-NY |  |  | 27359.6 | -13679.8 | 1.0965 |  | M2 - AOSEM | Upper Left (UL) | 251 |  |  |  | 26163.5 | -13081.75 | 1.1466
|  | Lower Left (LL) | 357 | PY-NY |  |  | 23647.3 | -11823.65 | 1.2686 |  |  | Lower Left (LL) | 320 | PY-NY |  |  | 24780.7 | -12390.35 | 1.2106 |  |  | Lower Left (LL) | 201 |  |  |  | 25059 | -12529.5 | 1.1972
|  | Upper Right (UR) | 325 | PN-NY | 20.2 | 3.09 | 28352.7 | -14176.35 | 1.0581 |  |  | Upper Right (UR) | 228 | PY-NY |  |  | 25117.9 | -12558.95 | 1.1944 |  |  | Upper Right (UR) | 157 |  |  |  | 27691.2 | -13845.6 | 1.0834
|  | Lower Right (LR) | 457 | PY-NY |  |  | 27492.6 | -13746.3 | 1.0912 |  |  | Lower Right (LR) | 438 | PY-NY |  |  | 25521.6 | -12760.8 | 1.1755 |  |  | Lower Right (LR) | 114 |  |  |  | 25538.9 | -12769.45 | 1.1747
| M3 - AOSEM | Upper Left (UL) | 113 | PY-NY |  |  | 26658.3 | -13329.15 | 1.1254 |  | M3 - AOSEM | Upper Left (UL) | 227 | PY-NY |  |  | 28354.2 | -14177.1 | 1.058 |  | M3 - AOSEM | Upper Left (UL) | 138 |  |  |  | 24790.9 | -12395.45 | 1.2101
|  | Lower Left (LL) | 317 | PY-NY |  |  | 25853.3 | -12926.65 | 1.1604 |  |  | Lower Left (LL) | 353 | PY-NY |  |  | 26509.5 | -13254.75 | 1.1317 |  |  | Lower Left (LL) | 165 |  |  |  | 30108.1 | -15054.05 | 0.99641
|  | Upper Right (UR) | 378 | PN-NY | 19.9 | 3.21 | 26637.5 | -13318.75 | 1.1262 |  |  | Upper Right (UR) | 276 | PY-NY |  |  | 27920.9 | -13960.45 | 1.0745 |  |  | Upper Right (UR) | 126 |  |  |  | 26371.6 | -13185.8 | 1.1376
|  | Lower Right (LR) | 196 | PY-NY |  |  | 26148.3 | -13074.15 | 1.1473 |  |  | Lower Right (LR) | 299 | PY-NY |  |  | 28435.2 | -14217.6 | 1.055 |  |  | Lower Right (LR) | 129 |  |  |  | 25052 | -12526 | 1.1975
L1-MC2 | M1 - BOSEM | Top 1 (T1) | 406 |  | 42 | 11.83 | 30307 | -15153.5 | 0.98987 | H1-MC2 | M1 - BOSEM | Top 1 (T1) | 92 |  |  |  | 27702.4 | -13851.2 | 1.0829 | I1-MC2 | M1 - BOSEM | Top 1 (T1) | 70 |  |  |  | 27828.4 | -13914.2 | 1.078
|  | Top 2 (T2) | 596 | X |  |  | 26652 | -13326 | 1.1256 |  |  | Top 2 (T2) | 43 |  |  |  | 22923.7 | -11461.85 | 1.3087 |  |  | Top 2 (T2) | 57 |  |  |  | 30706.5 | -15353.25 | 0.97699
|  | Top 3 (T3) | 273 |  |  |  | 31266 | -15633 | 0.95951 |  |  | Top 3 (T3) | 30 |  |  |  | 28098.7 | -14049.35 | 1.0677 |  |  | Top 3 (T3) | 11 |  |  |  | 30323.2 | -15161.6 | 0.98934
|  | Left (LF) | 221 |  | 43.3 | 11.9 | 30241 | -15120.5 | 0.99203 |  |  | Left (LF) | 23 |  |  |  | 26766.2 | -13383.1 | 1.1208 |  |  | Left (LF) | 65 |  |  |  | 28982.9 | -14491.45 | 1.0351
|  | Right (RT) | 551 |  |  |  | 29842 | -14921 | 1.0053 |  |  | Right (RT) | 86 |  |  |  | 27693.4 | -13846.7 | 1.0833 |  |  | Right (RT) | 29 |  |  |  | 28070.4 | -14035.2 | 1.0687
|  | Side (SD) | 209 |  | 43.8 | 11.9 | 29114 | -14557 | 1.0304 |  |  | Side (SD) | 613 | X |  |  | 30169.8 | -15084.9 | 0.99437 |  |  | Side (SD) | 680 | X |  |  | 27231.3 | -13615.65 | 1.1017
| M2 - AOSEM | Upper Left (UL) | 134 | PY-NY |  |  | 26424 | -13212 | 1.1353 |  | M2 - AOSEM | Upper Left (UL) | 307 | PY-NY |  |  | 24825 | -12412.5 | 1.2085 |  | M2 - AOSEM | Upper Left (UL) | 143 |  |  |  | 26770.3 | -13385.15 | 1.1206
|  | Lower Left (LL) | 283 | PY-NY | 20 | 3.04 | 28967 | -14483.5 | 1.0357 |  |  | Lower Left (LL) | 474 | PY-NY |  |  | 24060 | -12030 | 1.2469 |  |  | Lower Left (LL) | 156 |  |  |  | 26376.8 | -13188.4 | 1.1374
|  | Upper Right (UR) | 367 | PY-NY | 20.1 | 3.08 | 29043 | -14521.5 | 1.033 |  |  | Upper Right (UR) | 439 | PY-NY |  |  | 24304 | -12152 | 1.2344 |  |  | Upper Right (UR) | 493 |  |  |  | 27128.1 | -13564.05 | 1.1059
|  | Lower Right (LR) | 117 | PY-NY |  |  | 25961 | -12980.5 | 1.1556 |  |  | Lower Right (LR) | 305 | PY-NY |  |  | 25954 | -12977 | 1.1559 |  |  | Lower Right (LR) | 273 |  |  |  | 24980.7 | -12490.35 | 1.2009
| M3 - AOSEM | Upper Left (UL) | 496 | PY-NY | 19.5 | 3.19 | 24706 | -12353 | 1.2143 |  | M3 - AOSEM | Upper Left (UL) | 207 | PY-NY |  |  | 24822 | -12411 | 1.2086 |  | M3 - AOSEM | Upper Left (UL) | 481 |  |  |  | 25232.9 | -12616.45 | 1.1889
|  | Lower Left (LL) | 408 | PY-NY |  |  | 29293 | -14646.5 | 1.0241 |  |  | Lower Left (LL) | 311 | PY-NY |  |  | 23630 | -11815 | 1.2696 |  |  | Lower Left (LL) | 260 |  |  |  | 27287.8 | -13643.9 | 1.0994
|  | Upper Right (UR) | 124 | PY-NY |  |  | 28422 | -14211 | 1.0555 |  |  | Upper Right (UR) | 446 | PY-NY |  |  | 26177 | -13088.5 | 1.146 |  |  | Upper Right (UR) | 139 |  |  |  | 25201.6 | -12600.8 | 1.1904
|  | Lower Right (LR) | 135 | PY-NY | 19.6 | 3.22 | 26225 | -13112.5 | 1.1439 |  |  | Lower Right (LR) | 188 | PY-NY |  |  | 16687 | -8343.5 | 1.7978 |  |  | Lower Right (LR) | 328 |  |  |  | 25243.1 | -12621.55 | 1.1884
L1-MC3 | M1 - BOSEM | Top 1 (T1) | 395 |  |  |  | 26205 | -13102.5 | 1.1448 | H1-MC3 | M1 - BOSEM | Top 1 (T1) | 571 |  |  |  | 30360.8 | -15180.4 | 0.98812 | I1-MC3 | M1 - BOSEM | Top 1 (T1) | 488 |  |  |  | 27719.8 | -13859.9 | 1.0823
|  | Top 2 (T2) | 618 | X |  |  | 28545 | -14272.5 | 1.051 |  |  | Top 2 (T2) | 565 |  |  |  | 28718.9 | -14359.45 | 1.0446 |  |  | Top 2 (T2) | 492 |  |  |  | 30461.8 | -15230.9 | 0.98484
|  | Top 3 (T3) | 548 |  |  |  | 25117 | -12558.5 | 1.1944 |  |  | Top 3 (T3) | 312 |  |  |  | 30896.5 | -15448.25 | 0.97098 |  |  | Top 3 (T3) | 498 |  |  |  | 26278.9 | -13139.45 | 1.1416
|  | Left (LF) | 218 |  |  |  | 29979 | -14989.5 | 1.0007 |  |  | Left (LF) | 74 |  |  |  | 27143.9 | -13571.95 | 1.1052 |  |  | Left (LF) | 505 |  |  |  | 26470.7 | -13235.35 | 1.1333
|  | Right (RT) | 404 |  |  |  | 30305 | -15152.5 | 0.98994 |  |  | Right (RT) | 313 |  |  |  | 30112.7 | -15056.35 | 0.99626 |  |  | Right (RT) | 506 |  |  |  | 29147.7 | -14573.85 | 1.0292
|  | Side (SD) | 536 |  |  |  | 32105 | -16052.5 | 0.93443 |  |  | Side (SD) | 694 | X |  |  | 25573.4 | -12786.7 | 1.1731 |  |  | Side (SD) | 623 | X |  |  | 24190.6 | -12095.3 | 1.2402
| M2 - AOSEM | Upper Left (UL) | 327 | PY-NY |  |  | 26867.3 | -13433.65 | 1.1166 |  | M2 - AOSEM | Upper Left (UL) | 252 |  |  |  | 27239.3 | -13619.65 | 1.1013 |  | M2 - AOSEM | Upper Left (UL) | 147 |  |  |  | 27907.1 | -13953.55 | 1.075
|  | Lower Left (LL) | 602 | PY-NY |  |  | 25442.8 | -12721.4 | 1.1791 |  |  | Lower Left (LL) | 412 |  |  |  | 27451.2 | -13725.6 | 1.0928 |  |  | Lower Left (LL) | 163 |  |  |  | 26988.8 | -13494.4 | 1.1116
|  | Upper Right (UR) | 243 | PY-NY |  |  | 26222.2 | -13111.1 | 1.1441 |  |  | Upper Right (UR) | 205 |  |  |  | 25912.8 | -12956.4 | 1.1577 |  |  | Upper Right (UR) | 122 |  |  |  | 25236.6 | -12618.3 | 1.1887
|  | Lower Right (LR) | 217 | PY-NY |  |  | 26557 | -13278.5 | 1.1296 |  |  | Lower Right (LR) | 284 |  |  |  | 25128.2 | -12564.1 | 1.1939 |  |  | Lower Right (LR) | 178 |  |  |  | 26845.4 | -13422.7 | 1.1175
| M3 - AOSEM | Upper Left (UL) | 431 | PY-NY |  |  | 23980.5 | -11990.25 | 1.251 |  | M3 - AOSEM | Upper Left (UL) | 308 |  |  |  | 26301.3 | -13150.65 | 1.1406 |  | M3 - AOSEM | Upper Left (UL) | 169 |  |  |  | 25825.2 | -12912.6 | 1.1617
|  | Lower Left (LL) | 342 | PY-NY |  |  | 25861.9 | -12930.95 | 1.16 |  |  | Lower Left (LL) | 290 |  |  |  | 27520 | -13760 | 1.0901 |  |  | Lower Left (LL) | 167 |  |  |  | 28041.9 | -14020.95 | 1.0698
|  | Upper Right (UR) | 346 | PY-NY |  |  | 23755.6 | -11877.8 | 1.2629 |  |  | Upper Right (UR) | 162 |  |  |  | 25461.7 | -12730.85 | 1.1782 |  |  | Upper Right (UR) | 166 |  |  |  | 26546.7 | -13273.35 | 1.1301
|  | Lower Right (LR) | 373 | PY-NY |  |  | 25964.7 | -12982.35 | 1.1554 |  |  | Lower Right (LR) | 185 |  |  |  | 26232.7 | -13116.35 | 1.1436 |  |  | Lower Right (LR) | 125 |  |  |  | 26233.8 | -13116.9 | 1.1436
L1-PRM | M1 - BOSEM | Top 1 (T1) | 547 |  |  |  | 27156 | -13578 | 1.1047 | H1-PRM | M1 - BOSEM | Top 1 (T1) | 569 |  |  |  | 25291.2 | -12645.6 | 1.1862 | I1-PRM | M1 - BOSEM | Top 1 (T1) | 81 |  |  |  | 31280.1 | -15640.05 | 0.95908
|  | Top 2 (T2) | 598 | X |  |  | 31854 | -15927 | 0.9418 |  |  | Top 2 (T2) | 12 |  |  |  | 28022.7 | -14011.35 | 1.0706 |  |  | Top 2 (T2) | 309 |  |  |  | 28738.7 | -14369.35 | 1.0439
|  | Top 3 (T3) | 400 |  |  |  | 31403 | -15701.5 | 0.95532 |  |  | Top 3 (T3) | 40 |  |  |  | 27251.7 | -13625.85 | 1.1008 |  |  | Top 3 (T3) | 162 |  |  |  | 28471.7 | -14235.85 | 1.0537
|  | Left (LF) | 557 |  |  |  | 28352 | -14176 | 1.0581 |  |  | Left (LF) | 66 |  |  |  | 28020.2 | -14010.1 | 1.0707 |  |  | Left (LF) | 54 |  |  |  | 30398.5 | -15199.25 | 0.98689
|  | Right (RT) | 226 |  |  |  | 26037 | -13018.5 | 1.1522 |  |  | Right (RT) | 14 |  |  |  | 30290.4 | -15145.2 | 0.99041 |  |  | Right (RT) | 105 |  |  |  | 30157.9 | -15078.95 | 0.99476
|  | Side (SD) | 411 |  |  |  | 30530 | -15265 | 0.98264 |  |  | Side (SD) | 614 | X |  |  | 27229.7 | -13614.85 | 1.1017 |  |  | Side (SD) | 632 | X |  |  | 30377.4 | -15188.7 | 0.98758
| M2 - AOSEM | Upper Left (UL) | 490 | PY-NY |  |  | 25160 | -12580 | 1.1924 |  | M2 - AOSEM | Upper Left (UL) | 233 |  |  |  | 25221 | -12610.5 | 1.1895 |  | M2 - AOSEM | Upper Left (UL) | 175 |  |  |  | 26160.5 | -13080.25 | 1.1468
|  | Lower Left (LL) | 296 | PY-NY |  |  | 28730 | -14365 | 1.0442 |  |  | Lower Left (LL) | 485 |  |  |  | 26504.6 | -13252.3 | 1.1319 |  |  | Lower Left (LL) | 142 |  |  |  | 25702.3 | -12851.15 | 1.1672
|  | Upper Right (UR) | 257 | PY-NY |  |  | 24997 | -12498.5 | 1.2001 |  |  | Upper Right (UR) | 495 |  |  |  | 26713 | -13356.5 | 1.123 |  |  | Upper Right (UR) | 168 |  |  |  | 27877.1 | -13938.55 | 1.0762
|  | Lower Right (LR) | 278 | PY-NY |  |  | 23562 | -11781 | 1.2732 |  |  | Lower Right (LR) | 268 |  |  |  | 26766.5 | -13383.25 | 1.1208 |  |  | Lower Right (LR) | 127 |  |  |  | 26789.9 | -13394.95 | 1.1198
| M3 - AOSEM | Upper Left (UL) | 464 | PY-NY |  |  | 24360 | -12180 | 1.2315 |  | M3 - AOSEM | Upper Left (UL) | 246 |  |  |  | 25851 | -12925.5 | 1.1605 |  | M3 - AOSEM | Upper Left (UL) | 423 |  |  |  | 25865.7 | -12932.85 | 1.1598
|  | Lower Left (LL) | 226 | PY-NY |  |  | 28348 | -14174 | 1.0583 |  |  | Lower Left (LL) | 208 |  |  |  | 25623.2 | -12811.6 | 1.1708 |  |  | Lower Left (LL) | 119 |  |  |  | 27218 | -13609 | 1.1022
|  | Upper Right (UR) | 280 | PY-NY |  |  | 27465 | -13732.5 | 1.0923 |  |  | Upper Right (UR) | 300 |  |  |  | 27786.9 | -13893.45 | 1.0796 |  |  | Upper Right (UR) | 121 |  |  |  | 26020.9 | -13010.45 | 1.1529
|  | Lower Right (LR) | 360 | PY-NY |  |  | 27109 | -13554.5 | 1.1066 |  |  | Lower Right (LR) | 314 |  |  |  | 26341.7 | -13170.85 | 1.1389 |  |  | Lower Right (LR) | 182 |  |  |  | 25015.5 | -12507.75 | 1.1993
L1-PR2 | M1 - BOSEM | Top 1 (T1) | 408 |  | 42.2 | 11.89 | 27793 | -13896.5 | 1.0794 | H1-PR2 | M1 - BOSEM | Top 1 (T1) | 48 |  |  |  | 30234.4 | -15117.2 | 0.99225 | I1-PR2 | M1 - BOSEM | Top 1 (T1) | 412 |  |  |  | 25940 | -12970 | 1.1565
|  | Top 2 (T2) | 594 | X |  |  | 27637 | -13818.5 | 1.0855 |  |  | Top 2 (T2) | 97 |  |  |  | 28169 | -14084.5 | 1.065 |  |  | Top 2 (T2) | 413 |  |  |  | 28324 | -14162 | 1.0592
|  | Top 3 (T3) | 554 |  | 42.6 | 11.9 | 26482 | -13241 | 1.1328 |  |  | Top 3 (T3) | 60 |  |  |  | 28338 | -14169 | 1.0586 |  |  | Top 3 (T3) | 478 |  |  |  | 31125.7 | -15562.85 | 0.96383
|  | Left (LF) | 383 |  |  |  | 28906 | -14453 | 1.0378 |  |  | Left (LF) | 51 |  |  |  | 29158.1 | -14579.05 | 1.0289 |  |  | Left (LF) | 500 |  |  |  | 26572.2 | -13286.1 | 1.129
|  | Right (RT) | 549 |  |  |  | 25798 | -12899 | 1.1629 |  |  | Right (RT) | 87 |  |  |  | 31048.1 | -15524.05 | 0.96624 |  |  | Right (RT) | 508 |  |  |  | 27334.9 | -13667.45 | 1.0975
|  | Side (SD) | 556 |  | 42.1 | 11.92 | 26370 | -13185 | 1.1377 |  |  | Side (SD) | 653 | X |  |  | 27322.5 | -13661.25 | 1.098 |  |  | Side (SD) | 629 | X |  |  | 29323.7 | -14661.85 | 1.0231
| M2 - AOSEM | Upper Left (UL) | 241 | PY-NY |  |  | 27209 | -13604.5 | 1.1026 |  | M2 - AOSEM | Upper Left (UL) | 245 | PY-NY |  |  | 26789 | -13394.5 | 1.1199 |  | M2 - AOSEM | Upper Left (UL) | 137 |  |  |  | 26945.2 | -13472.6 | 1.1134
|  | Lower Left (LL) | 405 | PY-NY | 20.2 | 3.11 | 28824 | -14412 | 1.0408 |  |  | Lower Left (LL) | 238 | PY-NY |  |  | 29938.6 | -14969.3 | 1.0021 |  |  | Lower Left (LL) | 118 |  |  |  | 25280.1 | -12640.05 | 1.1867
|  | Upper Right (UR) | 456 | PY-NY | 20.1 | 3.13 | 27080 | -13540 | 1.1078 |  |  | Upper Right (UR) | 221 | PY-NY |  |  | 26124.4 | -13062.2 | 1.1484 |  |  | Upper Right (UR) | 414 |  |  |  | 26952.8 | -13476.4 | 1.1131
|  | Lower Right (LR) | 159 | PY-NY |  |  | 26590 | -13295 | 1.1282 |  |  | Lower Right (LR) | 234 | PY-NY |  |  | 25651.2 | -12825.6 | 1.1695 |  |  | Lower Right (LR) | 173 |  |  |  | 25271.8 | -12635.9 | 1.1871
| M3 - AOSEM | Upper Left (UL) | 409 | PN-NY | 19.7 | 3.24 | 27227 | -13613.5 | 1.1018 |  | M3 - AOSEM | Upper Left (UL) | 242 | PY-NY |  |  | 24033.7 | -12016.85 | 1.2482 |  | M3 - AOSEM | Upper Left (UL) | 202 |  |  |  | 26195 | -13097.5 | 1.1453
|  | Lower Left (LL) | 467 | PY-NY |  |  | 26424 | -13212 | 1.1353 |  |  | Lower Left (LL) | 215 | PY-NY |  |  | 25365 | -12682.5 | 1.1827 |  |  | Lower Left (LL) | 144 |  |  |  | 26546.5 | -13273.25 | 1.1301
|  | Upper Right (UR) | 604 | PY-NY |  |  | 29510 | -14755 | 1.0166 |  |  | Upper Right (UR) | 401 | PY-NY |  |  | 27214.8 | -13607.4 | 1.1023 |  |  | Upper Right (UR) | 181 |  |  |  | 28736.5 | -14368.25 | 1.044
|  | Lower Right (LR) | 265 | PY-NY | 19.5 | 3.25 | 29090 | -14545 | 1.0313 |  |  | Lower Right (LR) | 195 | PY-NY |  |  | 25800.2 | -12900.1 | 1.1628 |  |  | Lower Right (LR) | 172 |  |  |  | 28397.8 | -14198.9 | 1.0564
L1-SRM | M1 - BOSEM | Top 1 (T1) | 357 |  |  |  | 25872 | -12936 | 1.1596 | H1-SRM | M1 - BOSEM | Top 1 (T1) | 644 |  |  |  | 27948.8 | -13974.4 | 1.0734 | I1-SRM | M1 - BOSEM | Top 1 (T1) | 38 |  |  |  | 26396.2 | -13198.1 | 1.1365
|  | Top 2 (T2) | 621 | X |  |  | 28886 | -14443 | 1.0386 |  |  | Top 2 (T2) | 642 |  |  |  | 26275.4 | -13137.7 | 1.1418 |  |  | Top 2 (T2) | 59 |  |  |  | 29834.9 | -14917.45 | 1.0055
|  | Top 3 (T3) | 333 |  |  |  | 24751 | -12375.5 | 1.2121 |  |  | Top 3 (T3) | 431 |  |  |  | 29673.6 | -14836.8 | 1.011 |  |  | Top 3 (T3) | 153 |  |  |  | 28856.9 | -14428.45 | 1.0396
|  | Left (LF) | 343 |  |  |  | 25683 | -12841.5 | 1.1681 |  |  | Left (LF) | 419 |  |  |  | 28787.8 | -14393.9 | 1.0421 |  |  | Left (LF) | 277 |  |  |  | 29027.6 | -14513.8 | 1.0335
|  | Right (RT) | 371 |  |  |  | 26140 | -13070 | 1.1477 |  |  | Right (RT) | 432 |  |  |  | 27091.4 | -13545.7 | 1.1074 |  |  | Right (RT) | 287 |  |  |  | 28343.8 | -14171.9 | 1.0584
|  | Side (SD) | 326 |  |  |  | 26539 | -13269.5 | 1.1304 |  |  | Side (SD) | 685 | X |  |  | 26751 | -13375.5 | 1.1215 |  |  | Side (SD) | 566 | X |  |  | 27429.2 | -13714.6 | 1.0937
| M2 - AOSEM | Upper Left (UL) | 340 |  |  |  | 25196 | -12598 | 1.1907 |  | M2 - AOSEM | Upper Left (UL) | 334 |  |  |  | 28555 | -14277.5 | 1.0506 |  | M2 - AOSEM | Upper Left (UL) | 133 |  |  |  | 24533.5 | -12266.75 | 1.2228
|  | Lower Left (LL) | 223 |  |  |  | 24166 | -12083 | 1.2414 |  |  | Lower Left (LL) | 341 |  |  |  | 26662 | -13331 | 1.1252 |  |  | Lower Left (LL) | 266 |  |  |  | 25324.5 | -12662.25 | 1.1846
|  | Upper Right (UR) | 272 |  |  |  | 26129 | -13064.5 | 1.1481 |  |  | Upper Right (UR) | 343 |  |  |  | 28021.4 | -14010.7 | 1.0706 |  |  | Upper Right (UR) | 123 |  |  |  | 26265.7 | -13132.85 | 1.1422
|  | Lower Right (LR) | 193 |  |  |  | 26807 | -13403.5 | 1.1191 |  |  | Lower Right (LR) | 383 |  |  |  | 25299.5 | -12649.75 | 1.1858 |  |  | Lower Right (LR) | 161 |  |  |  | 26036.6 | -13018.3 | 1.1522
| M3 - AOSEM | Upper Left (UL) | 120 |  |  |  | 24866 | -12433 | 1.2065 |  | M3 - AOSEM | Upper Left (UL) | 425 |  |  |  | 27335.1 | -13667.55 | 1.0975 |  | M3 - AOSEM | Upper Left (UL) | 455 |  |  |  | 27091.6 | -13545.8 | 1.1074
|  | Lower Left (LL) | 352 |  |  |  | 24570 | -12285 | 1.221 |  |  | Lower Left (LL) | 349 |  |  |  | 29221.5 | -14610.75 | 1.0266 |  |  | Lower Left (LL) | 131 |  |  |  | 27768.8 | -13884.4 | 1.0803
|  | Upper Right (UR) | 476 |  |  |  | 22098 | -11049 | 1.3576 |  |  | Upper Right (UR) | 323 |  |  |  | 27061.6 | -13530.8 | 1.1086 |  |  | Upper Right (UR) | 164 |  |  |  | 25567 | -12783.5 | 1.1734
|  | Lower Right (LR) | 333 |  |  |  | 27868 | -13934 | 1.0765 |  |  | Lower Right (LR) | 466 |  |  |  | 25756 | -12878 | 1.1648 |  |  | Lower Right (LR) | 171 |  |  |  | 25923.5 | -12961.75 | 1.1573
L1-SR2 | M1 - BOSEM | Top 1 (T1) | 394 |  |  |  | 26772 | -13386 | 1.1206 | H1-SR2 | M1 - BOSEM | Top 1 (T1) | 76 |  |  |  | 26125.3 | -13062.65 | 1.1483 | I1-SR2 | M1 - BOSEM | Top 1 (T1) | 62 |  |  |  | 30991 | -15495.5 | 0.96802
|  | Top 2 (T2) | 610 | X |  |  | 27545 | -13772.5 | 1.0891 |  |  | Top 2 (T2) | 423 |  |  |  | 27116.1 | -13558.05 | 1.1064 |  |  | Top 2 (T2) | 272 |  |  |  | 24826.8 | -12413.4 | 1.2084
|  | Top 3 (T3) | 399 |  |  |  | 26136 | -13068 | 1.1478 |  |  | Top 3 (T3) | 646 |  |  |  | 25764.8 | -12882.4 | 1.1644 |  |  | Top 3 (T3) | 286 |  |  |  | 29801.2 | -14900.6 | 1.0067
|  | Left (LF) | 382 |  |  |  | 26110 | -13055 | 1.149 |  |  | Left (LF) | 150 |  |  |  | 31458.2 | -15729.1 | 0.95365 |  |  | Left (LF) | 304 |  |  |  | 29092.1 | -14546.05 | 1.0312
|  | Right (RT) | 397 |  |  |  | 28531 | -14265.5 | 1.0515 |  |  | Right (RT) | 98 |  |  |  | 25390.6 | -12695.3 | 1.1815 |  |  | Right (RT) | 307 |  |  |  | 28284.7 | -14142.35 | 1.0606
|  | Side (SD) | 405 |  |  |  | 29859 | -14929.5 | 1.0047 |  |  | Side (SD) | 668 | X |  |  | 25655.6 | -12827.8 | 1.1693 |  |  | Side (SD) | 651 | X |  |  | 25362.5 | -12681.25 | 1.1828
| M2 - AOSEM | Upper Left (UL) | 140 | PY-NY |  |  | 23639 | -11819.5 | 1.2691 |  | M2 - AOSEM | Upper Left (UL) | 232 |  |  |  | 26261.4 | -13130.7 | 1.1424 |  | M2 - AOSEM | Upper Left (UL) | 149 |  |  |  | 27430.2 | -13715.1 | 1.0937
|  | Lower Left (LL) | 294 | PY-NY |  |  | 23174 | -11587 | 1.2946 |  |  | Lower Left (LL) | 213 |  |  |  | 27884.6 | -13942.3 | 1.0759 |  |  | Lower Left (LL) | 141 |  |  |  | 31448.1 | -15724.05 | 0.95395
|  | Upper Right (UR) | 372 | PY-NY |  |  | 22916 | -11458 | 1.3091 |  |  | Upper Right (UR) | 279 |  |  |  | 25761.6 | -12880.8 | 1.1645 |  |  | Upper Right (UR) | 158 |  |  |  | 25698.8 | -12849.4 | 1.1674
|  | Lower Right (LR) | 190 | PY-NY |  |  | 22846 | -11423 | 1.3131 |  |  | Lower Right (LR) | 366 |  |  |  | 26451.8 | -13225.9 | 1.1341 |  |  | Lower Right (LR) | 259 |  |  |  | 24949.5 | -12474.75 | 1.2024
| M3 - AOSEM | Upper Left (UL) | 354 | PY-NY |  |  | 23431 | -11715.5 | 1.2804 |  | M3 - AOSEM | Upper Left (UL) | 312 |  |  |  | 26656.6 | -13328.3 | 1.1254 |  | M3 - AOSEM | Upper Left (UL) | 198 |  |  |  | 27541.6 | -13770.8 | 1.0893
|  | Lower Left (LL) | 451 | PY-NN |  |  | 24650 | -12325 | 1.217 |  |  | Lower Left (LL) | 424 |  |  |  | 25976.2 | -12988.1 | 1.1549 |  |  | Lower Left (LL) | 136 |  |  |  | 25554.1 | -12777.05 | 1.174
|  | Upper Right (UR) | 437 | PY-NY |  |  | 29188 | -14594 | 1.0278 |  |  | Upper Right (UR) | 285 |  |  |  | 25678.3 | -12839.15 | 1.1683 |  |  | Upper Right (UR) | 180 |  |  |  | 28374.7 | -14187.35 | 1.0573
|  | Lower Right (LR) | 606 | PY-NN |  |  | 25618 | -12809 | 1.1711 |  |  | Lower Right (LR) | 304 |  |  |  | 25379.6 | -12689.8 | 1.1821 |  |  | Lower Right (LR) | 170 |  |  |  | 24631.6 | -12315.8 | 1.2179

## OMC
Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain
L1-OMC | M1 - BOSEM | Top 1 (T1) | 329 |  |  |  | 23102 | -11551 | 1.2986
|  | Top 2 (T2) | 609 | X |  |  | 22698 | -11349 | 1.3217
|  | Top 3 (T3) | 330 |  |  |  | 22760 | -11380 | 1.3181
|  | Left (LF) | 349 |  |  |  | 23291 | -11645.5 | 1.2881
|  | Right (RT) | 366 |  |  |  | 22740 | -11370 | 1.3193
|  | Side (SD) | 369 |  |  |  | 23330 | -11665 | 1.2859
H1-OMC | M1 - BOSEM | Top 1 (T1) | [5 empty] | 0 | #DIV/0!
|  | Top 2 (T2) | [5 empty] | 0 | #DIV/0!
|  | Top 3 (T3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!
I1-OMC | M1 - BOSEM | Top 1 (T1) | [5 empty] | 0 | #DIV/0!
|  | Top 2 (T2) | [5 empty] | 0 | #DIV/0!
|  | Top 3 (T3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!

## TMTS
Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain
L1-TMSX | M1 - BOSEM | Top 1 (F1) |  |  |  |  | 26243.4 | -13121.7 | 1.1431
|  | Top 2 (F2) |  |  |  |  | 22632.8 | -11316.4 | 1.3255
|  | Top 3 (F3) |  |  |  |  | 28289.9 | -14144.95 | 1.0604
|  | Left (LF) |  |  |  |  | 22671.6 | -11335.8 | 1.3232
|  | Right (RT) |  |  |  |  | 25339.8 | -12669.9 | 1.1839
|  | Side (SD) |  |  |  |  | 29658.6 | -14829.3 | 1.0115
L1-TMSY | M1 - BOSEM | Top 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Top 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Top 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!
H1-TMSX | M1 - BOSEM | Top 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Top 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Top 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!
H1-TMSY | M1 - BOSEM | Top 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Top 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Top 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!
I1-TMSX | M1 - BOSEM | Top 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Top 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Top 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!
I1-TMSY | M1 - BOSEM | Top 1 (F1) | [5 empty] | 0 | #DIV/0!
|  | Top 2 (F2) | [5 empty] | 0 | #DIV/0!
|  | Top 3 (F3) | [5 empty] | 0 | #DIV/0!
|  | Left (LF) | [5 empty] | 0 | #DIV/0!
|  | Right (RT) | [5 empty] | 0 | #DIV/0!
|  | Side (SD) | [5 empty] | 0 | #DIV/0!

## HAUX
Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain | Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. [Ω] | Coil Ind. [mH] | Open Light Value [counts] | Offset [counts] | Gain | Suspension | Stage/Type | Location | Serial Number | Fully Charact. | Coil Res. (Ω) | Coil Ind. (mH) | Open Light Value [counts] | Offset [counts] | Gain
L1 | M1 - AOSEM | Upper Left (UL) | 477 | PY-NY |  |  |  | 0 | #DIV/0! | H1 | M1 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! | I1 | M1 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
S/N 001 |  | Lower Left (LL) | 429 | PN-NY |  |  |  | 0 | #DIV/0! | S/N 001 |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! | S/N 001 |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
SM1 |  | Upper Right (UR) | 418 | PY-NY |  |  |  | 0 | #DIV/0! | SM1 |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! | SM1 |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
IM1 |  | Lower Right (LR) | 247 | PY-NY |  |  |  | 0 | #DIV/0! | IM1 |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! | IM1 |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
L1 | M1 - AOSEM | Upper Left (UL) | 274 | PY-NY |  |  |  | 0 | #DIV/0! | H1 | M1 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! | I1 | M1 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
S/N 002 |  | Lower Left (LL) | 479 | PY-NY |  |  |  | 0 | #DIV/0! | S/N 002 |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! | S/N 002 |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
SM2 |  | Upper Right (UR) | 255 | PY-NY |  |  |  | 0 | #DIV/0! | SM2 |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! | SM2 |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
IM4 |  | Lower Right (LR) | 448 | PY-NY |  |  |  | 0 | #DIV/0! | IM4 |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! | IM4 |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
L1 | M1 - AOSEM | Upper Left (UL) | 329 | PY-NY |  |  |  | 0 | #DIV/0! | H1 | M1 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! | I1 | M1 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
S/N 003 |  | Lower Left (LL) | 603 | PY-NY |  |  |  | 0 | #DIV/0! | S/N 003 |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! | S/N 003 |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
PMMT1 |  | Upper Right (UR) | 469 | PY-NY |  |  |  | 0 | #DIV/0! | PMMT1 |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! | PMMT1 |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
IM2 |  | Lower Right (LR) | 315 | PY-NY |  |  |  | 0 | #DIV/0! | IM2 |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! | IM2 |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!
L1 | M1 - AOSEM | Upper Left (UL) | 435 | PY-NY |  |  |  | 0 | #DIV/0! | H1 | M1 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0! | I1 | M1 - AOSEM | Upper Left (UL) | [5 empty] | 0 | #DIV/0!
S/N 004 |  | Lower Left (LL) | 297 | PY-NY |  |  |  | 0 | #DIV/0! | S/N 004 |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0! | S/N 004 |  | Lower Left (LL) | [5 empty] | 0 | #DIV/0!
PMMT2 |  | Upper Right (UR) | 462 | PY-NY |  |  |  | 0 | #DIV/0! | PMMT2 |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0! | PMMT2 |  | Upper Right (UR) | [5 empty] | 0 | #DIV/0!
IM3 |  | Lower Right (LR) | 415 | PY-NY |  |  |  | 0 | #DIV/0! | IM3 |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0! | IM3 |  | Lower Right (LR) | [5 empty] | 0 | #DIV/0!