LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1200632-v5	Advanced LIGO	02/6/2014

TwinCAT Library for the
ALS Laser

Alexa Staley, Daniel Sigg

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note
of the LIGO Laboratory.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 159 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

| Library |
| --- |
| Title | ALSLaser |
| Version | 4 |
| TwinCAT version | 2.11 |
| Name space | – |
| Author | Alexa Staley, Daniel Sigg |
| Description | Monitors the ALS Laser Controls the crystal and doubler temperature A slow feedback servo is implemented to offload the fast PZT feedback to the slow temperature controls. It implements the following equation: ui = (1 – h) ui-1 + ei (h – g) Or if the polarity is false: ui = (1 – h) ui-1 + ei (h + g) with g = πfugf ∆t and h = πfPf∆t ∆𝑡: sampling interval, ,fugf : unity gain frequency of integrator, ,fPf : Knee frequency of proportional gain. (only used for control of the laser crystal, not the doubler) |
| Error codes | 0x0001 – Safety interlock engaged 0x0002 – Laser diode 1 guard alarm 0x0004 – Laser diode 2 guard alarm 0x0008 – Laser diode 1 current out-of-range 0x0010 – Laser diode 2 current out-of-range 0x0020 – Laser crystal TEC error signal out-of-range 0x0040 – Doubler crystal TEC error signal out-of-range 0x0080 – Laser diode 1 TEC error signal out-of-range 0x0100 – Laser diode 2 TEC error signal out-of-range 0x0200 – Noise eater readback signal out-of-range 0x0400 – Unity gain frequency too high 0x0800 – Proportional gain knee frequency too high 0x1000 – Temperature feedback limit exceeded |
| Library dependencies | ReadADC, WriteADC, SaveRestore, Error |

| Hardware Input Type TYPE ALSLaserInStruct : STRUCT LaserDiode1PowerMonitor: INT; LaserDiode2PowerMonitor: INT; LaserCrystalTECErrorSignal: INT; DoublingCrystalTECErrorSignal: INT; LaserDiode1TECErrorSignal: INT; LaserDiode2TECErrorSignal: INT; NoiseEaterMonitor: INT; LaserDiode1TempGuard: BOOL; LaserDiode2TempGuard: BOOL; InterLock: BOOL; END_STRUCT END_TYPE |
| --- |
| Type name | ALSLaserInStruct |
| Description | Structure of the hardware input that are wired up for the ALS laser |
| Definition | STRUCT |
| Element | Name: LaserDiode1PowerMonitor Type: INT Description: Laser diode 1 power monitor |
| Element | Name: LaserDiode2PowerMonitor Type: INT Description: Laser diode 2 power monitor |
| Element | Name: LaserCrystalTECErrorSignal Type: INT Description: Laser crystal, TEC error signal |
| Element | Name: DoublingCrystalTECErrorSignal Type: INT Description: Doubling crystal, TEC error signal |
| Element | Name: LaserDiode1TECErrorSignal Type: INT Description: Laser diode 1, TEC error signal |
| Element | Name: LaserDiode2TECErrorSignal Type: INT Description: Laser diode 2, TEC error signal |
| Element | Name: NoiseEaterMonitor Type: INT Description: Noise eater monitor |
| Element | Name: LaserDiode1TempGuard Type: BOOL Description: Laser diode 1, temp guard |
| Element | Name: LaserDiode2TempGuard Type: BOOL Description: Laser diode 2, temp guard |
| Element | Name: InterLock Type: BOOL Description: InterLock |

| Hardware Output Type TYPE ALSLaserOutStruct : STRUCT CrystalTemperature: INT; DoublerTemperature: INT; NoiseEaterRelayOn: BOOL; NoiseEaterRelayOff: BOOL; END_STRUCT END_TYPE |
| --- |
| Type name | ALSLaserOutStruct |
| Description | Structure of the hardware output that are wired up for the ALS laser |
| Definition | STRUCT |
| Element | Name: CrystalTemperature Type: INT Description: Crystal Temperature |
| Element | Name: DoublerTemperature Type: INT Description: Doubler Temperature |
| Element | Name: NoiseEaterRelayOn Type: BOOL Description: External relay to flip noise eater ON |
| Element | Name: NoiseEaterRelayOff Type: BOOL Description: External relay to flip noise eater OFF |

| User Interface Type TYPE ALSLaserFrequencyControlsStruct: STRUCT On: BOOL; Enabled: BOOL; Run: BOOL; Reset: BOOL; Low: LREAL; High: LREAL; Range: BOOL; Ugf: LREAL; Pf: LREAL; Polarity: BOOL; END_STRUCT END_TYPE |
| --- |
| Type name | ALSLaserFrequencyControlsStruct |
| Description | Structure used in the user interface that are used for the slow temperature controls |
| Definition | STRUCT |
| In/out Tag | Name: On Type: BOOL Description: On/Off button for temperature feedback |
| Output Tag | Name: Enabled Type: BOOL Description: Enabled by the autolocker |
| Output Tag | Name: Run Type: BOOL Description: Servo is running |
| In/out Tag | Name: Reset Type: BOOL Description: Reset the integrator and zero the output |
| In/out Tag | Name: Low Type: LREAL Description: Low limit for feedback controls in Hz |
| In/out Tag | Name: High Type: LREAL Description: High limit for feedback controls in Hz |
| Output Tag | Name: Range Type: BOOL Description: Feedback controls exceeds range |
| In/out Tag | Name: Ugf Type: LREAL Description: Unity gain frequency of temperature servo in Hz |
| In/out Tag | Name: Pf Type: LREAL Description: Knee frequency of proportional feedback in Hz; zero is none |

| User Interface Type TYPE ALSLaserStruct : STRUCT Error: ErrorStruct; LaserDiode1PowerMonitor: LREAL; LaserDiode2PowerMonitor: LREAL; LaserDiodePowerNominal: LREAL; LaserDiodePowerTolerance: LREAL; LaserCrystalTECErrorSignal: LREAL; DoublingCrystalTECErrorSignal: LREAL; LaserDiode1TECErrorSignal: LREAL; LaserDiode2TECErrorSignal: LREAL; TECTolerance: LREAL; NoiseEaterMonitor: LREAL; NoiseEaterNominal: LREAL; NoiseEaterTolerance: LREAL; NoiseEaterRelay: BOOL; LaserDiode1TempGuard: BOOL; LaserDiode2TempGuard: BOOL; InterLock: BOOL; CrystalTemperature: LREAL; CrystalCalibration: LREAL; CrystalFrequency: LREAL; FrequencyControl: ALSLaserFrequencyControlsStruct; PZTTuningCoefficient: LREAL; PZTFrequency: LREAL; DoublerTemperature: LREAL; END_STRUCT END_TYPE |
| --- |
| Type name | ALSLaserStruct |
| Description | Structure of the user interface tags that are used to control the ALS Laser |
| Definition | STRUCT |
| Output Tag | Name: Error Type: ErrorStruct Description: For error handler |
| Output Tag | Name: LaserDiode1PowerMonitor Type: LREAL Description: Laser diode 1 power monitor in A |
| Output Tag | Name: LaserDiode2PowerMonitor Type: LREAL Description: Laser diode 2 power monitor in A |
| Input Tag | Name: LaserDiodePowerNominal Type: LREAL Description: Laser diode power nominal in A |
| Input Tag | Name: LaserDiodePowerTolerance Type: LREAL Description: Laser diode power tolerance in A |
| Output Tag | Name: LaserCrystalTECErrorSignal Type: LREAL Description: Laser crystal, TEC error signal in C |
| Output Tag | Name: DoublingCrystalTECErrorSignal Type: LREAL Description: Doubling crystal, TEC error signal in C |
| Output Tag | Name: LaserDiode1TECErrorSignal Type: LREAL Description: Laser diode 1, TEC error signal in C |
| Output Tag | Name: LaserDiode2TECErrorSignal Type: LREAL Description: Laser diode 2, TEC error signal in C |
| Input Tag | Name: TECTolerance Type: LREAL Description: TEC error signal tolerance in C |
| Output Tag | Name: NoiseEaterMonitor Type: LREAL Description: Noise eater monitor in V |
| Input Tag | Name: NoiseEaterNominal Type: LREAL Description: Noise eater nominal value in V |
| Input Tag | Name: NoiseEaterTolerance Type: LREAL Description: Noise eater tolerance in V |
| Output Tag | Name: NoiseEaterRelay Type: BOOL Description: External relay to switch noise eater on/off |
| Output Tag | Name: LaserDiode1TempGuard Type: BOOL Description: Laser diode 1, temp guard, high is alarm |
| Output Tag | Name: LaserDiode2TempGuard Type: BOOL Description: Laser diode 2, temp guard, high is alarm |
| Output Tag | Name: InterLock Type: BOOL Description: InterLock, high represent an interlock |
| In/out Tag | Name: CrystalTemperature Type: LREAL Description: Crystal Temperature in C or V |
| Output Tag | Name: CrystalCalibration Type: LREAL Description: Crystal temperature coefficient at 1064nm in MHz/C; nominal -3000 MHz/C |
| In/out Tag | Name: CrystalFrequency Type: LREAL Description: Laser frequency as set by crystal temperature in MHz; updating the CrystalFrequency will update the CrystalTemperature and vis versa |
| In/out Tag | Name: FrequencyControl Type: ALSLaserFrequencyControlsStruct Description: Controls parameters for slow temperature feedback network |
| Output Tag | Name: PZTTuningCoefficient Type: LREAL Description: PZT tuning coefficient at 1064nm in MHz/V; nominal 1.5 MHz/V |
| Output Tag | Name: PZTFrequency Type: LREAL Description: Laser frequency as set by the PZT actuator in MHz |
| Input Tag | Name: DoublerTemperature Type: LREAL Description: Doubler Temperature |

| Function Block FUNCTION_BLOCK ALSLaserFB VAR_INPUT Request: SaveRestoreEnum; ALSLaserIn: ALSLaserInStruct; ConstrolsEnable: BOOL:= FALSE; PZTVoltage: LREAL := 0.0; END_VAR VAR_OUTPUT ALSLaserOut: ALSLaserOutStruct; END_VAR VAR_IN_OUT ALSLaserInit: ALSLaserStruct; ALSLaser: ALSLaserStruct; END_VAR |
| --- |
| Name | ALSLaserFB |
| Description | Monitors the ALS laser and computes the slow controls feedback |
| Input argument | Name: Request Type: SaveRestoreEnum Description: Request for save/restore/safemode or noop |
| Input argument | Name: ALSLaserIn Type: ALSLaserInStruct Description: Input hardware structure |
| Input argument | Name: ControlsEnable Type: BOOL Description: Enables the slow controls feedback Default: FALSE |
| Input argument | Name: PZTVoltage Type: LREAL Description: PZT controls voltage in V (error signal for slow controls feedback) Default: 0 |
| Output argument | Name: ALSLaserOut Type: ALSLaserOutStruct Description: Output hardware structure |
| In/out argument | Name: ALSLaser Type: ALSLaserStruct Description: User Interface structure |
| In/out argument | Name: ALSLaserInit Type: ALSLaserStruct Description: Save/restore variable in persistent memory |

| Visual |
| --- |
| Name | ALSLaserVis |
| Description | Displays power monitors, TEC error signals, noise eater status, crystal temperatures, slow temperature controls parameters, and alarms for temperature guards, interlock and error. |
| Placeholder | Name: ALSLaser Type: ALSLaserStruct Description: ALS laser structure |