aLIGO One Arm Test Commissioning Update: So far, so good.
Kawabe + many people

---
Initially Proposed Straw Man (T1100080)

---
Reality
1wk
Done
Done
Done
NOW

---
Initial alignment
- 2 days to shoot the beam down through the tube
- 5 additional days until the first lock
- Both baffle diodes and ccds are essential for alignment.
- Working PDH diode and demod are essential for alignment.
  - Use it instead of looking at the video for flashing, as you can't see it “flashing” until the laser is half-locked to the cavity.
- Working PLL is essential for alignment.
  - Free running freq. noise is large enough to make the PDH error blurry.

---
ISC: Almost everything works

---
ISC Status
- PLL UGF: 22kHz (about the same as design)
- PDH UGF: about 12kHz (as opposed to 2kHz design)
- Cavity lock lasts for hours
- VCO offloading to ETM path for “slowing” cavity (M0-only and BW very small for now)
- Calibrated spectrum (more on this later, G1200752)
  - Very good length calibration using VCO
  - Noise dominated by TM vibration f<1.2 Hz or so (as opposed to f<0.5Hz initial estimate)

---
PLL and PDH locking
Initial design
- PDH performance (BW~12kHz ) better than design
- Gives more headroom for ALS comm and diff feedback

---
Quad Feedback
Quad Feedback
- ETMY Longitudinal-to-Pitch coupling mitigation.
- Longitudinal feedback (at the TOP) couples into small Pitch signal (inherent in the mechanical suspension).
- Filter design to reduce this coupling
  - Model shows 76 pole-zero pairs, this is to hard to implement.
  - Reducing pairs which cancel, less then 10% different and above ~3.5 Hz.
- Work in progress.

---
ISC Status 2
- Active pointing control was implemented
  - Open loop pointing error meets the spec (Stochino)
- Fringe visibility of REFL 50-60%
- WFS as a sensor is being implemented
  - Gouy telescope redone (initial design produced tiny beam)
  - Possibility of control later
- Slow control (Beckhoff, OPC)
  - Most of the things work
  - No automation yet (auto locker for PLL and PDH)
  - OPC-EPICS translation issues being investigated

---
Note: data is the equivalent OPEN LOOP ERROR, and the requirement is for CLOSED LOOP. UGF=10Hz, so it looks like both angle and displacement already meets the requirements.

---
SUS
- Damping characterization/optimization effort.
  - Plot by Kissel
  - Spec-wise, long way to go: IY-EY < 10nm rms for f<0.5Hz with slow signal from PDH offloaded to EY
- Length-PIT actuation decoupling effort.
- Sensing diagonalization effort.

---
SEI
- Some trouble in ISI isolation.
  - Need a better high freq. (e.g. 1Hz) performance
- See Vincent/Fabrice's presentation.
- Despite these, we're already close to low frequency spec (IY-EY<250 nm rms w/o global feedback).

---
TCS
See Aidan's presentation.