## Slide 1: Dual PD Amp Circuit Board Test Results
By Alexa Staley
LIGO G1200781-v1

## Slide 2: What is the Dual PD Amplifier?
Will be placed in the five ALS Fiber Distribution Chassis
Has two channels that amplify a given source via four different gain settings
D1200543-v1

## Slide 3: ALS Fiber Distribution (brief)
Shifts the frequency of the PSL laser light after the reference cavity
Splits the input into four output fibers

## Slide 4: Some Specs.
When connected to a power supply with no source
    DC offset is measured to be less than 10mV for each gain settings
    See less than 20uV AC variations
    Minimal signal drift (but needs to be tested for longer period of time)

## Slide 5: Optical Transfer Function

## Slide 6: Transfer Functions

## Slide 7: Optical Transfer Function with Long Cable

## Slide 8: Shot Noise

## Slide 9: Total, Shot, & Electronics Noise

## Slide 10: Noise for Different Cable Lengths

## Slide 11: Noise Data

## Slide 12: Next Steps…
Possibly take measurements with more data points
Take shot noise measurements with a better light source
Replace 1206 10k resistors that remain with 0805 size – could reduce noise even further
Test long term drift
Produce 5 PCBs and place into ALS Fiber Distribution Chassis  test chassis