# Tested By: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ Date: \_\_\_\_\_\_\_\_\_

# ASC (WFS) Style Single Frequency Detector Measured Parameters

All transimpedance measurements are referred to plane of the physical output connector and include the effect of the voltage divider created by the 50 Ω termination. The notch rejection ratios are relative to the magnitude of the transimpedance at the respective RF detection center frequency of the given RF output port. The notation, Q1 to Q4 refers to the specific quadrant of a four section (Quad) diode.

Table 1

|  |  |
| --- | --- |
| **Unit identification** | **Value** |
| Photodetector serial number |  |
| Detector schematic D# and revision |  |
| Diode element manufacturer’s Part  and serial number |  |

Table

|  |  |
| --- | --- |
| **DC Parameters** | **Value** |
| Quiescent DC current (amps at +18 VDC) |  |
| Quiescent DC current (amps at -18 VDC) |  |
| PD bias regulator output voltage (VDC) |  |
| RF opamp positive voltage regulator (VDC) |  |
| RF opamp negative voltage regulator (VDC) |  |
| Audio opamp positive voltage regulator (VDC) |  |
| Audio opamp negative voltage regulator (VDC) |  |

Table

|  |  |
| --- | --- |
| **DC readout transimpedance**  **(Ω at differential output)** | **Value** |
| Q1 |  |
| Q2 |  |
| Q3 |  |
| Q4 |  |

Table

|  |  |
| --- | --- |
| **Global RF parameters** | **Value** |
| RF detection center frequency (MHz) |  |
| 2ω Notch frequency (MHz) |  |

Table

|  |  |
| --- | --- |
| **Q1 RF notch** | **Value** |
| Measured DC photocurrent (mA) |  |
| Rejection at 2ω notch (dB) |  |

Table

|  |  |
| --- | --- |
| **Q2 RF notch** | **Value** |
| Measured DC photocurrent (mA) |  |
| Rejection at 2ω notch (dB) |  |

Table

|  |  |
| --- | --- |
| **Q3 RF notch** | **Value** |
| Measured DC photocurrent (mA) |  |
| Rejection at 2ω notch (dB) |  |

Table

|  |  |
| --- | --- |
| **Q4 RF notch** | **Value** |
| Measured DC photocurrent (mA) |  |
| Rejection at 2ω notch (dB) |  |

Table

|  |  |
| --- | --- |
| **RF transimpedance** | **Value** |
| Q1 Transimpedance (Ω) |  |
| Q2 Transimpedance (Ω) |  |
| Q3 Transimpedance (Ω) |  |
| Q4 Transimpedance (Ω) |  |

Table

|  |  |
| --- | --- |
| **Shot-noise limited input sensitivity** | **Value** |
| Q1 (mA) |  |
| Q2 (mA) |  |
| Q3 (mA) |  |
| Q4 (mA) |  |