# Tested By: Jay Copti					               Date: 04/04/2014

# LSC Style Detector Measured Parameters

All transimpedance measurements are referred to plane of the physical output connector and include the effect of the voltage divider created by the 50 Ω termination.  All notch rejection ratios are relative to the magnitude of the transimpedance at the respective RF detection center frequency of the given RF output port.

| Parameter | Value |
| --- | --- |
| Detector serial number | S1203398 |
| Detector schematic D# and revision | D1101124-v7 |
| Diode element manufacturer’s serial number | A6896 |
| Quiescent DC current (amps at +18 VDC) | 60.8 mA |
| Quiescent DC current (amps at -18 VDC) | 93.2 mA |
| PD bias regulator output voltage (VDC) | 5.05 VDC |
| RF opamp positive voltage regulator (VDC) | 5.85 VDC |
| RF opamp negative voltage regulator (VDC) | -6.01 VDC |
| Audio opamp positive voltage regulator (VDC) | 14.80 VDC |
| Audio opamp negative voltage regulator (VDC) | -15.32 VDC |
| DC path transimpedance and zero light offset (Ω/mVDC at BNC out) | 099 Ω | 0.2 mVDC |
| DC path transimpedance and zero light offset (Ω/mVDC at differential out) | 199 Ω | 0.0 mVDC |
| DC path zero frequency (Hz) | 0.2 Hz |
| DC path pole frequency (Hz) | 2.4 Hz |
| Inferred DC path shot noise limited input photo sensitivity (mA) at 100Hz measured at differential output | 2.4 mA |
| RF detection center frequency (MHz), f low | 24.078 MHz |
| RF detection center frequency (MHz), f hi | N/A |
| Notch frequencies (MHz) used in design | 48.156 MHz (3 Notches, Same Freq) |
| F low feedback notch frequency | N/A |
| F hi feedback notch frequency | N/A |
| Rejection (dB) at notch1 (f low) | -36.9 |
| Rejection (dB) at notch2 (f low) | N/A |
| Rejection (dB) at notch3 (f low) | N/A |
| Rejection (dB) at notch4 (f low) | N/A |
| Rejection (dB) f low to f hi | N/A |
| Rejection (dB) at notch1 (f hi) | N/A |
| Rejection (dB) at notch2 (f hi) | N/A |
| Rejection (dB) at notch3 (f hi) | N/A |
| Rejection (dB) at notch4 (f hi) | N/A |
| Rejection (dB) f hi to f low | N/A |
| Transimpedance (Ω) at f low (note PD Current) | 489 Ω | 10.0 mA |
| RF dark/light noise used for f low Trans-Z | -125.7 dBm/Hz | -117.1 dBm/Hz |
| Transimpedance (Ω) at f hi (note PD Current) | N/A Ω | N/A mA |
| RF dark/light noise used for f hi Trans-Z | N/A dBm/Hz | N/A dBm/Hz |
| RF preamp used during testing (noise/gain) | 143.2 dBm/Hz | 20.4dB |
| f low, shot-noise limited input sensitivity (mA) | 1.56 mA |
| f hi, shot-noise limited input sensitivity (mA) | N/A |
| Test input transconductance at f1(mA/V) | 17.4 mA/V |
| Test switch isolation at f1 (dB) | 27.6 dB |
| Test input transconductance at f2(mA/V) | N/A |
| Test switch isolation at f2 (dB) | N/A |