# Tested By:  Jay Copti						        Date: 08/30/2013

# LSC Style Detector Measured Parameters

All transimpedance measurements are referred to plane of the physical output connector and include the effect of the voltage divider created by the 50 Ω termination.  All notch rejection ratios are relative to the magnitude of the transimpedance at the respective RF detection center frequency of the given RF output port.

| Parameter | Value |
| --- | --- |
| Detector serial number | S1300534 |
| Detector schematic D# and revision | D1101124-v7 |
| Diode element manufacturer’s serial number | A6906 |
| Quiescent DC current (amps at +18 VDC) | 70.1 mA |
| Quiescent DC current (amps at -18 VDC) | 92.5 mA |
| PD bias regulator output voltage (VDC) | 5.06 VDC |
| RF opamp positive voltage regulator (VDC) | 5.85 VDC |
| RF opamp negative voltage regulator (VDC) | -6.06 VDC |
| Audio opamp positive voltage regulator (VDC) | 14.77 VDC |
| Audio opamp negative voltage regulator (VDC) | -15.35 VDC |
| DC path transimpedance and zero light offset (Ω/mVDC at BNC out) | 98.6 Ω | 0.9 mVDC |
| DC path transimpedance and zero light offset (Ω/mVDC at differential out) | 198 Ω | 1.9 mVDC |
| DC path zero frequency (Hz) | 0.2 Hz |
| DC path pole frequency (Hz) | 2.4 Hz |
| Inferred DC path shot noise limited input photo sensitivity (mA) at 100Hz measured at differential output | 3.4 mA |
| RF detection center frequency (MHz), f low | 9 MHz |
| RF detection center frequency (MHz), f hi | 45 MHz |
| Notch frequencies (MHz) used in design | 18, 36, 54, 90 MHz |
| F low feedback notch frequency | N/A |
| F hi feedback notch frequency | N/A |
| Rejection (dB) at notch1 (f low) | -31.5 dB |
| Rejection (dB) at notch2 (f low) | -35.8 dB |
| Rejection (dB) at notch3 (f low) | -37.4 dB |
| Rejection (dB) at notch4 (f low) | -33.5 dB |
| Rejection (dB) f low to f hi | -34.3 dB |
| Rejection (dB) at notch1 (f hi) | -51.7 dB |
| Rejection (dB) at notch2 (f hi) | -34.8 dB |
| Rejection (dB) at notch3 (f hi) | -30.5 dB |
| Rejection (dB) at notch4 (f hi) | -34.5 dB |
| Rejection (dB) f hi to f low | -57.2 dB |
| Transimpedance (Ω) at f low (note PD Current) | 595 Ω | 10.5 mA |
| RF dark/light noise used for f low Trans-Z | -124.7 dBm/Hz | -115.3 dBm/Hz |
| Transimpedance (Ω) at f hi (note PD Current) | 731 Ω | 10.5 mA |
| RF dark/light noise used for f hi Trans-Z | -121.3 dBm/Hz | -113.3 dBm/Hz |
| RF preamp used during testing (noise/gain) | 143 dBm/Hz | 20.4 dB |
| f low, shot-noise limited input sensitivity (mA) | 1.34 mA |
| f hi, shot-noise limited input sensitivity (mA) | 1.96 mA |
| Test input transconductance at f1(mA/V) | 14.6 mA/V |
| Test switch isolation at f1 (dB) | 31.9 dB |
| Test input transconductance at f2(mA/V) | 14.4 mA/V |
| Test switch isolation at f2 (dB) | 32.5 dB |