## Sheet1
Data taken on S1300520 LSC RFPD, page 64 of notebook #4. Nominal component values per D1101124-v7 component variations sheet T1200428‐v6
9MHz Measured Transimpedance | 523
9MHz Measured Shot Noise Equivalent Photocurrent | 1.35mA
Dark Noise measured at 9MHz at Output Connector | -146.2 dBm/Hz | 11nV/rtHz
| 18MHz | 36MHz | 54MHz | 45MHz | 90MHz
Transfer function measurement relative to 9MHz Output (dB) | -31 | -36 | -37 | -36 | -33
Calculated 9MHz Output Transimpedance (Ω) | 14.7401 | 8.289 | 7.3876 | 8.289 | 11.7085
45MHz Measured Transimpedance | 688
45MHz Measured Shot Noise Equivalent Photocurrent | 1.98mA
Dark Noise measured at 45MHz at Output Connector | -142.2 dBm/Hz | 17.4nV/rtHz
| 9MHz | 18MHz | 36MHz | 54MHz | 90MHz
Transfer function measurement relative to 45MHz Output (dB) | -58 | -52 | -35 | -30 | -35
Calculated 45MHz Output Transimpedance (Ω) | 0.86614 | 1.7282 | 12.2346 | 21.7565 | 12.2346
Photocurrent (mA) associated with each optical power level assuming QE = 95% at REFL Port
| Carrier | 9MHz | 18MHz | 27MHz | 36MHz | 45MHz | 54MHz | 63MHz | 72MHz | 81MHz | 90MHz
| 40.2978 | 0.032524 | 7.9034 | 0 | 4.5209 | 0.0081311 | 1.7401 | 0 | 0.048787 | 0 | 2.3255
Anticipated Signal Level (mVrms) at output pin of RF opamp | 18MHz | 9MHz: 36MHz; 18MHz: 54MHz; 27MHz: 45MHz; 36MHz: 90MHz
9 MHz Input | 232.996 | 9MHz: 74.9475; 18MHz: 25.7096; 27MHz: 0.1348; 36MHz: 54.4563
Anticipated Signal Level (mVrms) at output pin of RF opamp | 9MHz | 9MHz: 18MHz; 18MHz: 36MHz; 27MHz: 54MHz; 36MHz: 90MHz
45MHz Output | 0.056342 | 9MHz: 27.3171; 18MHz: 110.6225; 27MHz: 75.7151; 36MHz: 56.903