## Sheet1
E1400152-v1 aLIGO LLO INSTALLATION "PUNCH LIST"
Action Item |  |  |  | Installation Subset |  | Report DCC#
# | Priority | Description | Assigned | DCC # | Description
1 | M | Revise the LBSC2 installation procedure E1200329: a) revise BSC cable tray to as-built b) pg. 10, sect 4.8, cable as per <missing DCC #> c) pg.11, sect 4.14, install BS elliptical baffles per <missing DCC #> d) pg. 14, sect 5.2, install field cabling per <missing DCC #> e) no comments or signoffs in the procedure; either remove or add notes (presumably from the elog) f) pg 10, complete weight of BS suspension |  | E1300847-v7 | LBSC2 | E1300967-v1
2 | L | Safety officer to consider revision to the LBSC2 cartridge installation hazard analysis (E1200327-v4): a) pg. 9, section 3.8, identify anchor points for fall arrest system b) pg. 11, section 5.1, why limit visibility with safety glasses during a critical, sight-dependent, operation? |  | E1300847-v7 | LBSC2 | E1300967-v1
3 | H | E1300847-v7 is missing reference to a BS Optical Lever testing report |  | E1300847-v7 | LBSC2 | E1300967-v1
4 | M | E1300847-v7 points to the L1 BS SUS Fabrication Acceptance document (E1300699) for the phase 2 & 3 test results for the BS Suspension. However this document only gives links to the elog, not to a report(s) (This report(s) should be linked to E1000496, aLIGO SUS BSFM Testing and Commissioning Documentation) |  | E1300847-v7 | LBSC2 | E1300967-v1
5 | L | There are some issues with the ICS entry for D0900428 for LBSC2. Some TCS and SLC records have been added but do not appear. Still need to add viewports, Oplev periscope and perhaps misc. other items to the ICS records. |  | E1300847-v7 | LBSC2 | E1300967-v1
6
7
8
9