## Individual Acceptance
|  | Status of SUS Individual Acceptance Reports |  |  |  | T1400181-v16
|  | Norna A Robertson | Calum I Torrie
6th June 2014
item number | SITE | Suspension Type | Suspension name | report DCC # | Filecard prepared | Review started | Approved | Count | Punchlist Items*,** | Update on Punchlist Status | Items to Complete before Acceptance? | Assignment of punchlist to aLIGO Project or Ops (project unless stated) and status | Assignee | Due Dates | Punchlist Closed | Open Bugzilla references
1 | LHO-H1 | QUAD | H1 ITMX | E1400146 | [6 empty] | not reviewed yet
2 |  | QUAD | H1 ITMY | E1400147 | [6 empty] | not reviewed yet
3 |  | QUAD | H1 ETMX | E1300908 | yes | yes | yes | 1 |  | (For completeness - info on optics, ears and fibers has been added to ICS entry -not noted on punchlist) | NO | N/A |  |  | YES
4 |  | QUAD | H1 ETMY | E1300866 | yes | yes | yes | 1 |  | (For completeness - info on optics, ears and fibers has been added to ICS entry -not noted on punchlist) | NO | N/A |  |  | YES
5 | LHO - 3rd IFO | QUAD | 3rd IFO ITMX | E1400148 | [6 empty] | not reviewed yet
6 |  | QUAD | 3rd IFO ITMY | E1400149 | [6 empty] | not reviewed yet
7 |  | QUAD | 3rd IFO ETMX | E1400150 | yes | [5 empty] | not reviewed yet
8 |  | QUAD | 3rd IFO ETMY | E1400151 | [6 empty] | not reviewed yet
9 | LLO-L1 | QUAD | L1 ITMX | E1300916 | yes | yes | yes | 1 |  |  | NO | N/A |  |  | YES | 758: L1 ITMX L2 (PUM) LL OSEM magnet & flag alignment stability
10 |  | QUAD | L1 ITMY | E1300917 | yes | yes | yes | 1 |  |  | NO | N/A |  |  | YES
11 |  | QUAD | L1 ETMX | E1400080 | yes | yes | yes | 1 |  |  | NO | N/A |  |  | YES | 680: LBSC4 Quad has a bad BOSEM L1 UR
12 |  | QUAD | L1 ETMY | E1400131 | yes | yes | yes | 1 | 1) Need "excel spreadsheet" with input data from build and or other information from fiber, ears, optics etc ... 2) No electronic rack data presented in a traditional sense. However, if one goes to S1202953 (as below) the content of the rack are the related document. Included for completeness and to spell out technique used. Included in punch-list to confirm this is good. 3) Attach link to alog for violin mode measurements. | Item 1) Info on fibers etc added to ICS and info on total mass added to ICS. Item 3) DONE S.Aston - 04/07/2014 | NO | All closed |  |  | YES
13 | LHO-H1 | BS | H1 BS | E1300712 | yes | yes | yes | 1 | Calibrated OSEM sensor ASDs Page 12 of 1st document in alog entry 7723, we see excess motion with control loops on. It's difficult for the casual observer to understand why. Arnaud can retake data after vent (maybe weeks) |  | NO | Follow up needed on excess motion seen. | Arnaud Pele | 2014-07-31
14 | LHO - 3rd IFO | BS | L1 BS | E1300699 | yes | yes | yes | 1 |  |  | NO |  |  |  | YES
15 | LLO-L1 | BS | 3rd IFO BS | E1400157 | yes | yes | yes | 1 | 1) Transfer function data noisy. This is due to result performed on free standing stand. Accepted but noted. No Action. |  | NO |  |  |  | YES
16 | LHO-H1 | HLTS | H1 PR3 | E1300844 | yes | yes | yes | 1 | 1) SUS TO UPDATE OPTIC ASSEMBLY DWG. 2) RACK DWG (See below) 3) M3 to M3 transfer function. 4. Check and add excel spreadsheet to ICS entry. Note added 16th April 2014 (NAR) - item 2) has now been addressed, rack drawings are updated. | 1) Optics assembly drawing has been updated as per DCN E1300890. 2) has now been addressed as noted in previous column. 3) - this has been passed to install acceptance since it is in-chamber measurement. 4) has been done - Excel spreadsheet attached to ICS record. | NO | 3) will be added to chamber acceptance, otherwise closed. |  |  | YES
17 | LHO-H1 | HLTS | H1 SR3 | E1400161 | yes | yes | yes | 1 | No Punch-list items. Note Phase 2b testing only. Beyond phase 2b refer to chamber acceptance. |  | NO | Item will be added to chamber acceptance |  |  | YES
18 | LHO - 3rd IFO | HLTS | 3rd IFO PR3 | E1400159 | yes | yes | yes | 1 | 1) Note transfer function taken on non-ideal table. So are noisy. Noted only. No action. 2) Y YAW has cross coupling from long into yaw. It goes away under damping. Noted. No Action. |  | NO |  |  |  | YES
19 | LHO - 3rd IFO | HLTS | 3rd IFO SR3 | E1400160 | yes | yes | yes | 1 |  |  | NO |  |  |  | YES
20 | LLO-L1 | HLTS | L1 PR3 | E1300836 | yes | yes | yes | 1 | 1) SUS TO UPDATE PR3 OPTIC ASSEMBLY DWG. 2) SUS to add mass information to ICS i.e. EXCEL spreadsheet with build numbers . | 1) Optics assembly drawing has been updated as per DCN E1300890, 2) Mass information extracted from alog entry has been added as an attachment to ICS entry | NO | 2) Add excel sheet to ICS | 1) Janeen Romie (Matt Heintze) | 2014-07-31
21 | LLO-L1 | HLTS | L1 SR3 | E1400115 | yes | yes | yes | 1 | 1) Need excel spreadsheet with input data from build. 2) Refer Bug 733 for cross coupling issue highlighted below. No fix. 3) Magnet missing from SR3. Refer to bug 175. No fix planned for now. | 1) Info on total mass added to ICS. | NO | 1) Add excel sheet to ICS. 2) and 3) will be tracked through bug list, otherwise closed. | 1) Janeen Romie (Matt Heintze) | 2014-07-31 |  | 733: L1 SR3 (HLTS) pitch to vertical cross coupling, 175: LHAM5 - SR3 magnet came off optic lower left position 831: L1 SR3 (HLTS) M2 UL channel high frequency turn-up
22 | LHO-H1 | HSTS | H1 MC1 | E1400118 | yes | yes | yes | 1 |  |  | NO |  |  |  | YES
23 | LHO-H1 | HSTS | H1 MC2 | E1400119 | yes | yes | yes | 1 | 1) Pitch / roll mode coupled to vertical. Recommend add to bug-list and track. Refer to alog 6552. | Added to bug list | NO | 1) will be tracked through bug list. |  |  | YES | 847: Excess cross-coupling in MC2 suspension under vacuum
24 | LHO-H1 | HSTS | H1 MC3 | E1400120 | yes | yes | yes | 1 |  |  | NO |  |  |  | YES
25 | LHO-H1 | HSTS | H1 PRM | E1400121 | yes | yes | yes | 1 | 1) Refer to alog entry 11121 and note 60Hz peak is dominant for top stage osems of PRM. Recommend check for grounding at next opportunity. Please add to bug list. | Added to bug list | NO | 1) will be tracked through bug list. |  |  | YES | 843: 60Hz spikes in OSEM spectra for H1 PRM suspension
26 | LHO-H1 | HSTS | H1 PR2 | E1400122 | yes | yes | yes | 1 | 1) Refer to alog entry 7568 and observe Vertical to pitch cross coupling @ 2.8Hz. No action required. |  | NO |  |  |  | YES
27 | LHO-H1 | HSTS | H1 SRM | E1400123 | yes | yes | yes | 1 | No punch list items. Note that data to phase 2b only. Final review of data post 2b in chamber acceptance review. |  | NO | Item will be added to chamber acceptance |  |  | YES
28 | LHO-H1 | HSTS | H1 SR2 | E1400124 | yes | yes | yes | 1 | No punch list items. Completed to 2b. Phase 3a and 3b review will be in chamber acceptance review. |  | NO | Item will be added to chamber acceptance |  |  | YES
29 | LHO - 3rd IFO | HSTS | 3rd IFO MC1 | E1300514 | yes | yes | yes | 1 | 1) Minor cross coupling from vertical to roll Alog entry 5684. Minor. Suppressed with damping. SUS note this was the last built and these had least matched blades. Noted. No Action. |  | NO |  |  |  | YES
30 | LHO - 3rd IFO | HSTS | 3rd IFO MC2 | E1400125 | yes | yes | yes | 1 |  |  | NO |  |  |  | YES
31 | LHO - 3rd IFO | HSTS | 3rd IFO MC3 | E1400126 | yes | yes | yes | 1 |  |  | NO |  |  |  | YES
32 | LHO - 3rd IFO | HSTS | 3rd IFO PRM | E1300497 | yes | yes | yes | 1 | 1) Cross coupling observed of vertical into roll / pitch. Suppressed when damping on. Noted. Accepted. No Action. |  | NO |  |  |  | YES
33 | LHO - 3rd IFO | HSTS | 3rd IFO PR2 | E1400127 | yes | yes | yes | 1 |  |  | NO |  |  |  | YES
34 | LHO - 3rd IFO | HSTS | 3rd IFO SRM | E1400128 | yes | yes | yes | 1 |  |  | NO |  |  |  | YES
35 | LHO - 3rd IFO | HSTS | 3rd IFO SR2 | E1400129 | yes | yes | yes | 1 |  |  | NO |  |  |  | YES
36 | LLO-L1 | HSTS | L1 MC1 | E1400109 | yes | yes | yes | 1 | Punch list item - need excel spreadsheet with input data from build. | Info on total mass added to ICS | NO | Add excel sheet to ICS | Janeen Romie (Matt Heintze) | 2014-07-31
37 | LLO-L1 | HSTS | L1 MC2 | E1201042 | yes | yes | yes | 1 | Punch list item - need excel spreadsheet with input data from build. (Not noted on filecard) | Info on total mass added to ICS | NO | Add excel sheet to ICS | Janeen Romie (Matt Heintze) | 2014-07-31
38 | LLO-L1 | HSTS | L1 MC3 | E1400113 | yes | yes | yes | 1 | Punch list item - need excel spreadsheet with input data from build. | Info on total mass added to ICS | NO | Add excel sheet to ICS | Janeen Romie (Matt Heintze) | 2014-07-31
39 | LLO-L1 | HSTS | L1 PRM | E1400114 | yes | yes | yes | 1 | Punch list item - need excel spreadsheet with input data from build. | Info on total mass added to ICS | NO | Add excel sheet to ICS | Janeen Romie (Matt Heintze) | 2014-07-31
40 | LLO-L1 | HSTS | L1 PR2 | E1300513 | yes | yes | yes | 1 | 1) Need excel spreadsheet with input data from build (not noted on filecard). 2) All three stages of transfer functions show a change of slope in magnitude, a "turn up," at the highest frequencies shown, typically above 10 [Hz] - to be investigated | 1) Info on total mass added to ICS | NO | 1) Add excel sheet to ICS 2) will be tracked through bug list. | 1) Janeen Romie (Matt Heintze) | 2014-07-31 |  | 327: Odd High-frequency behavior from all SUS Top2Top Transfer Functions (note this applies to other suspensions too)
41 | LLO-L1 | HSTS | L1 SRM | E1400116 | yes | yes | yes | 1 | 1) need excel spreadsheet with input data from build, 2) A cross-coupling of the 2nd roll mode at ~2.1 Hz into the transverse DOF is present, see https://services.ligo-wa.caltech.edu/integrationissues/show_bug.cgi?id=90 - this does not warrant a rebuild, thus we should proceed with install, 3) Verifying that the sensor read-back issue is only at the M2 stage and that M2-M3 actuation is OK. See below. This needs a bug. 4) M2 to M3 and M3 to M3 yaw shows some cross coupling. See links below. Included for completeness. No action recommended or required. | 1) Info on total mass added to ICS. | NO | 1) Add excel sheet to ICS. 2) and 3) will be tracked through bug list. | 1) Janeen Romie (Matt Heintze) | 2014-07-31 |  | 90: Excess cross-coupling in SRM at LLO 830: L1 SRM (HSTS) M2 stage sensor read-back issue
42 | LLO-L1 | HSTS | L1 SR2 | E1400117 | yes | yes | yes | 1 | Punch list item - need excel spreadsheet with input data from build. | 1) Info on total mass added to ICS. | NO | 1) Add excel sheet to ICS | 1) Janeen Romie (Matt Heintze) | 2014-07-31
43 | LHO-H1 | OMCS | H1 OMCS | E1400055 | yes | yes | yes | 1 | 1. SYS - Item #II below Top level assembly. Action on Systems (Eddie Sanchez). 2. SUS - Item #II below ICS. Action on Suspensions to find out why cables, bosems not present. (Janeen Romie).3. SUS - Item #II below ICS. Need to check that spreadsheet with as built masses, lengths etc has been attached to ICS. (Janeen Romie) | 1) This was revision of D1300240 . This has been done (confirmed by NAR with E Sanchez). 3) EXCEL spreadsheet is attached to the ICS. | NO | 2) to be completed/checked: check that info on BOSEMs and cables are in ICS | 2) Janeen Romie | 2014-07-31
44 | LHO - 3rd IFO | OMCS | 3rd IFO OMCS | E1400158 | yes | yes | yes | 1 | 1) It is noted that it is not always obviously clear (in test results) if X1 result is LHO or 3rd IFO. It is clear in text. No action. |  | NO |  |  |  | YES
45 | LLO-L1 | OMCS | L1 OMCS | E1400034 | yes | yes | yes | 1 | 1. SYS - Item #II below Top level assembly. Action on Systems (Eddie Sanchez). 2. SUS - Item #II below ICS. Action on Suspensions to find out why cables, bosems not present. (Janeen Romie).3. SUS - Item #II below ICS. Need to check that spreadsheet with as built masses, lengths etc has been attached to ICS. (Janeen Romie) | 1) This refers to revision of D1300077. This has been done (confirmed by NAR with E Sanchez). 2) BOSEMs are present in ICS. 3) EXCEL spreadsheet is attached to ICS | NO | 2) to be checked that cables are in ICS | 2) Janeen Romie | 2014-07-31
|  |  |  | total prepared | 40 |  | total approved | 39 | [6 empty] | colour scheme
[14 empty] | red | open punchlist item(s)
[9 empty] | * In general punchlist items brought up in individual acceptance reviews but which are common to all of a certain type of suspension, or to all suspensions, are listed on next page. |  |  |  |  | green | punchlist closed (or no punchlist)
[9 empty] | ** For 3rd IFO suspensions, the ICS assembly link on the filecards is to the generic assembly number for that type of suspension (HSTS etc) with uniquely identifying serial number, since a chamber assembly number does not exist. See E1400187 for list of the ICS identifier numbers. |  |  |  |  | yellow | punchlist closed, with item or items being tracked elsewhere

## General Acceptance
Status of General Acceptance Documentation
Item number (continued from previous page) | Suspension Type | DCC # | Filecard prepared | Review Started | Approved | Punchlist Items | Punchlist Items | Update on Punchlist Status | Items to Complete before Acceptance? | Assignment of punchlist to aLIGO Project or Ops and status | Assignee | Due Dates | Punchlist Closed
[6 empty] | common* | specific
46 | QUAD | E1201038 | yes | yes | yes | 1) Ops Manual (Ops instuctions), 2) Spares storage organisation, 3) Shipping/transportation Plans, 4) AOSEM FDD document T0900286 update, 5) LHO electronics doc tree | 6) Finish/Check Quad top level drawings including optics assemblies ALL updates will be covered via DCN E1300970 and signed off by SUS and SYS. | 1) extensive work has been done on Ops manual - currently being reviewed. Now reviewed (4 June 2014) by Rich A. He is happy with it, Calum agrees action closed. 4) AOSEM doc T0900286 has now been updated to v3 (6th June 2014). | NO | 1) and 4) closed. 2) and 3) are under Operations. 5) and 6) are Project and to be done/finished | 2) - Detector group (Janeen Romie, Vern Sandberg). 3) - To be assigned. 5) - Rich Abbott. 6) - Janeen Romie (Eddie Sanchez). | 2) End August 2014. 3) Open. 5) end August 2014. 6) end July 2014.
47 | BSFM | E1201039 | yes | yes | yes | ditto | 1) Finish/Check BSFM top level drawings |  | NO | 1) underway | 1) Janeen Romie (Eddie Sanchez) | end July 2014
48 | HLTS | E1201041 | yes | yes | yes | ditto | 1) Finish/Check HLTS top level drawings. 2) Install doc E1000045 update as needed | 1) Eddie Sanchez reports optics assemblies done. See DCN E1300890 | NO | 1) closed. 2) underway | 2) - Janeen Romie | end July 2014
49 | HSTS | E1201040 | yes | yes | yes | ditto | 1) Finish/Check HSTS top level drawings. 2) Install doc E0900334 update as needed | 1) Eddie Sanchez reports optics assemblies done. See DCN E1300890 2) updated on DCC 16th June 2014, and signed off. | NO | 1) and 2) closed |  |  | YES
50 | OMCS | E1201043 | yes | yes | yes | ditto | 1) Finish/Check OMCS top level drawings. 2) Control range doc and noise model doc to be done (Jeff K). 3) assembly doc - pictures to be added (Jeff B) | 3) OMCS Assembly doc is now updated including figures and posted to DCC. | NO | 1) underway, 2) to be done, 3) closed | 1) Janeen Romie (Eddie Sanchez), 2) Jeff Kissel. | 1) End July 2014, 2 ) end July 2014
* where an item has come up in one review but affects most or all types of suspensions, it is included under "common".