## Slide 1: A Preview of Future Cryogenic Suspensions for aLIGO Upgrades
Brett Shapiro
Stanford cryo people:
Brian Lantz, Tim MacDonald, Dakota Madden-Fong (summer ‘13)
LIGO
Notes: Don’t hold onto the podium so much
Look around the audience more, don’t just stare at one person
Take a few more moments to explain plots. E.g. this is a plot of the temperature vs time …. The goal is for the test mass to reach the dotted line at 124 Kas fast as possible
Explain the steady state vs rapid cooling operation more clearly

## Slide 2: Summary
Moving from Advanced LIGO to LIGO III
LIGO III quad suspension design
Steady state cooling (science mode)
  System layout
  Length of heat shield extending into beam tube
Initial Cool down
  Stanford initial cool down experiments
Future work

## Slide 3: Advanced LIGO Timeline
1st science run
Livingston, LA
Hanford, WA
Sept-Dec 2015
Last core optics
March 2014
Full lock testing
Aug-Sept 2014
Build LIGO III
GW
astronomy
Soon after
Far after

## Slide 4: Predicted Advanced LIGO Sensitivity

## Slide 5: Predicted Advanced LIGO Sensitivity
LIGO III aims to do better by a factor of 3

## Slide 6: Predicted Advanced LIGO Sensitivity
LIGO III aims to do better by a factor of 3
Use cryogenics to push thermal noise down

## Slide 7: LIGO III cryo work distribution
Courtesy of Nicolas Smith
Caltech - cryogenic reference cavities; direct thermal noise measurements
Jena/Glasgow/Moscow - mechanical loss
MIT – high emissivity coatings
KAGRA – 20 K sapphire suspensions
INPE Brazil – Cryogenic multi-nested pendulum
Stanford – optical coatings; cryogenic technology

## Slide 8: Cyro Test Mass Problem Statement
* For LIGO III, reduce suspension and coating thermal noise by cooling the lower quad to 124 K (-149.15 C)
  Si test masses (blue team in LIGO-T1200031)
2 cooling regimes:
  Initial cool down - get to 124 K quickly <- Stanford experiment
  Steady state - maintain 124 K once you get there
Include a warm-up scheme (don’t forget!)
Use the same seismic isolation platforms (ISIs, HEPIs)
  Limit the amount of extra weight on these plaforms
  Leave the rest of the vacuum chamber warm
Do not increase the test mass lossiness
  Emissive coatings, heat links, thick sus fibers, optical coatings, substrate, suspension fiber bonding, etc
Do not compromise passive seismic isolation
  Cables, hoses, links, etc

## Slide 9: Quad Suspension Design
Must incorporate much larger test mass

## Slide 10: LIGO III Quad Suspension Design
10 kg
20 kg
Main
chain
Reaction
chain
143 kg
Main
chain
Reaction
chain
Advanced LIGO quad pendulum
Preliminary LIGO III quad pendulum
2.14 m
1.626 m
40 kg
20 kg
40 kg
20 kg
Fused silica, 295 K
Silicon, 124 K (-149 C)
? kg
? kg
? kg

## Slide 11: Quad Design Requirements

## Slide 12: Quad Design Requirements
Payload P ≤ 270 kg for the main chain for aLIGO BSC-ISI

## Slide 13: Quad Design Requirements
Payload P ≤ 270 kg for the main chain for aLIGO BSC-ISI
Test mass m4 ≈ 143 kg to minimize radiation pressure and thermal noise

## Slide 14: Quad Design Requirements
Payload P ≤ 270 kg for the main chain for aLIGO BSC-ISI
Total length L ≤ 2.14 m to fit in vacuum system
Test mass m4 ≈ 143 kg to minimize radiation pressure and thermal noise

## Slide 15: Quad Design Requirements
Payload P ≤ 270 kg for the main chain for aLIGO BSC-ISI
Total length L ≤ 2.14 m to fit in vacuum system
Test mass m4 ≈ 143 kg to minimize radiation pressure and thermal noise
Seismic isolation as good as aLIGO

## Slide 16: Quad Design Requirements
Payload P ≤ 270 kg for the main chain for aLIGO BSC-ISI
Total length L ≤ 2.14 m to fit in vacuum system
Test mass m4 ≈ 143 kg to minimize radiation pressure and thermal noise
Seismic isolation as good as aLIGO
Vertical bounce mode fbounce ≤ aLIGO

## Slide 17: Quad Design Requirements
Payload P ≤ 270 kg for the main chain for aLIGO BSC-ISI
Total length L ≤ 2.14 m to fit in vacuum system
Test mass m4 ≈ 143 kg to minimize radiation pressure and thermal noise
Seismic isolation no worse than aLIGO
Vertical bounce mode fbounce ≤ aLIGO
Si fiber stress σ4 ≤ 1400 Mpa
Very rough guess

## Slide 18: 3 Optimal Quad Designs
See T1300786
Higher payload
Lighter
Test mass
PUM Si springs
Note: these results are highly dependent on the allowed Si fiber stress, σ4. There is still much uncertainty in this value.
10Hz

## Slide 19: All designs permit top mass damping

## Slide 20: Steady State Cooling
Keeping the test mass temperature at the operating point

## Slide 21: LIGO III Steady State Cooling
See also G1200246-v1
124 K
295 K
Heat shield at
80 K maintains 124 K through radiative coupling
Active Seismic Isolation Table
ISI stage 2
Vacuum flange
Flexible liquid nitrogen hoses,
1 mL/s laminar flow:
Supply (65 K) and return (75 K);
or Cu cables
clamp
1560 nm MW laser
≈ 3 W absorption at 124 K
≤ 10 m
143 kg Si test mass
…
…
chiller

## Slide 22: Pros and Cons of LN2 pipe vs. Cu cable
Cu cable
Pros:
No fluid to make noise
No LN2 pumping mechanism
No risk of N2 leaks
Cons:
Low heat transfer
...
LN2 pipe
Pros:
High heat transfer
No length / heat transfer trade-off
Low weight
…
Cons:
Risk of leaking
Fluid flow could contribute noise
…
Heat shield
ISI
Copper cable
Heat shield
ISI
Liquid N2 pipe, supply and return

## Slide 23: Pros and Cons of LN2 pipe vs. Cu cable
Cu cable
Pros:
No fluid to make noise
No LN2 pumping mechanism
No risk of N2 leaks
Cons:
Low heat transfer
Cryo refrigerator must be placed near feedthrough
High bulk: stiffness, weight, etc
  Q = K(A/L)ΔT
    Big L means big A
    Can reduce A by making cold end less than LN2 (77 K). E.g. Cryomech’s PT407 can pull 25W at 55 K.
  - Minimize stiffness by using lots of this wires, but wire dia must be > electron m.f.p.
Thermal conductivity decreases when wire dia becomes smaller than electron m.f.p.
Hysteresis
  - Lots of small wires sliding past each other
High difficulty in minimizing seismic shorting
- Minimize using:
    Lots of thin wires
    Intermediate masses along length
LN2 pipe
Pros:
High heat transfer
Low bulk
Moderate difficulty in minimizing seismic shorting
Length of pipe in vacuum not an issue for net heat transfer (longer pipes do require more pressure to push fluid)
Vibrating cryorefrigerator can be placed further from vacuum feedthrough.
Cons:
Complex LN2 pumping mechanism
Risk of leaking
Fluid flow could contribute noise
  Minimize by:
    Cooling the LN2 so it doesn’t boil
    Ensuring laminar flow
Pressure of flow and thermal contraction can influence table displacement
  - Minimize by locating pipes in strategic locations

## Slide 24: Beam Tube Heat Shield Length
L
What is the optimal heat shield length in the beam tube?
Beam tube
Blackbody disc model
Heat Shield
rj = 0.28 m
ri = 0.6225m
Test Mass
Tj = 120 K
Ti = 295 K
T = 80K
Εi = 1
Εj = 0.1
See T1300556

## Slide 25: Beam Tube Heat Shield Length
Beam tube shield length L (m)
Black body power from beam tube arriving at test mass
≈1/3 drop for every factor of 2 L
See T1300556
Test mass emissivity = 0.1

## Slide 26: End Station Vacuum System
10 m
5 m
0 m
Gate valves
End Station

## Slide 27: Initial Cooldown
Acquiring the operating temperature

## Slide 28: A Lot of Heat to Remove
60 W
143 kg Test Mass:
14 MJ to get cold (warm)
2.7 days to 124 K (295 K) at a constant 60 W
Maybe there is another way to cool the test mass quickly
* Radiation to the heat shield alone would take ≈ 2 weeks.

## Slide 29: Initial Cool Down Cold Link – 2 Designs
return
Test mass
Conductive cooling, low pressure N2 gas
Pros and Cons
liquid N2 pipe flexures
supply
thermally conductive plate with variable gap
Pros
Operates in partial vacuum.
Low heat transfer between cold and warm parts of vacuum system.
Fine temperature control – just back away when at desired temperature
Versatility, design permits both conductive and convective cooling.
Cons
Requires moving parts:
  flexible pipes
  actuators
Physically contacts barrel of test mass

## Slide 30: Initial Cool Down Cold Link – 2 Designs
Pros and Cons
Convective cooling, up to 1 atm N2 gas
Test mass
liquid N2
pipes
thermally conductive plate with large fixed gap
return
supply
convective N2 gas
Pros
No moving parts or actuators
No contact with test mass
Faster cooling than conduction
Cons
Convection between cold and warm parts of vacuum system
No fine temperature control – must return to vacuum to ‘turn off’ cold link.
Does not operate under vacuum

## Slide 31: Experimental Setup
Test dewar
Heat shield

## Slide 33: Close up of cold link
Test
mass
Cold link
Flexure ‘point’ contacts
Variable gap

## Slide 34: Flexible cold link evolution
Time

## Slide 35: Measurement – cold link engaged
140 min time constant

## Slide 36: Test mass temperature modeling

## Slide 37: Test mass temperature modeling

## Slide 38: Test mass temperature modeling
Heat transfer goes through residual N2 gas between test mass and cold link!

## Slide 39: Test mass temperature modeling
Heat transfer goes through residual N2 gas between test mass and cold link!
This is a perfect exponential decay…

## Slide 40: Exponential Temperature Decay
These are both functions of temperature.
In general, the solution is not truly exponential since the time ‘constant’ changes.

## Slide 41: Exponential Temperature Decay

## Slide 42: Exponential Temperature Decay

## Slide 43: Exponential Temperature Decay

## Slide 44: Measurement – cold link disengaged
160 min time constant

## Slide 45: Measurement – cold link disengaged
160 min time constant
Simulation assuming only N2 gas conduction

## Slide 46: Conductive vs Convective Regimes
Convection

## Slide 47: Finite Element Modeling
Images courtesy of Tim MacDonald
≈ 43 min into cool down
Cold link
Test mass
N2 gas
Due to complexity, LIGO III designs must be verified with modeling or FEM
Below: FEM of conduction through N2 gas to cold link for Stanford experiment
Convective FEM is proving to require large amounts of computing power

## Slide 48: Cold link on a LIGO III test mass
Test mass radius = 280 mm
Test mass thickness = 250 mm
Cold link covers 2/3 of top barrel
Conduction limit, no convection

## Slide 49: Stanford Next Gen Experiment

## Slide 50: Stanford Next Gen Experiment Layout
ISI Stage 2
ISI Stage 1
LN2 return
LN2 supply
Cold link
motor
Vacuum wall
LN2 feedthrough
GN2 feedthrough
Stainless steel legs
Heat shield
Test mass
LN2 pipe around shield
Cold link
Maraging steel blade spring
chiller

## Slide 51: Future work
Next generation experiment using the Stanford ETF (experimental test facility)
More realistic LIGO setup
Measure temperature drifts on LIGO hardware, e.g. blade springs
Measure seismic noise of nitrogen delivery and/or copper cables
Test heat shield design
  Black coatings
  baffles
Test a variety of cooling techniques
System integration: how to make all this stuff work together
Implement in stages
  Cables/hoses first – test seismic noise
  Heat shield and suspended optic
  Install cryogenic refrigerator
  Cavity?
  Anything we haven’t thought of yet
LIGO

## Slide 52: LVC STANFORD August 25-29 2014

## Slide 53
LVC STANFORD August 25-29 2014
Carothers Hall
(30) Single Rooms: $77.20 per person/night
(85) Double Rooms: $56.70 per person/night
Benefits:
5 minute walk to conference.
Outdoor facilities available:
Telephone in each
Comfortable and clean facility.
Onsite management.
Free laundry facilities.
Linens provided
Quiet areas including business center/lounge.
Crothers Hall is exclusive to LVC.
Free shuttle service to shopping/restaurants
Alternate accommodations (hotel)
$180-350/night
Registration – webpage coming soon
$250 Student rate
$500 Non-student rate
Includes:
Breakfast and lunch
Banquet dinner
Conference rooms
AV and AV Technicians
Poster session
Does not include:
Parking passes ($10/day)
Transportation to/from airport
Poster printing
Accommodations
Contact Claudette Earl
cearl@stanford.edu

## Slide 54: Backups
LIGO

## Slide 55: Test Mass in Next Gen Prototype
LN2 pipe flexure
(2nd one not shown)
Cold plate
LN2 pipe
Si test mass
6” dia, 1” thick, 1 kg

## Slide 56: Experimentation lessons learned
Air dominates most heat flow across contacts
Cold links should have distributed contacts
Solder is not leak tight against high pressure cryogenic fluids – welding is probably best
Cryogenic fluid should have 1 flow path
Send fluid from bottom up
Use fatter pipes to minimize fluid pressure
Minimize the number of materials in the plumbing – joining and contraction issues
Leave room for differential contraction
Silicon diode temperature sensors are great

## Slide 57: Steady State LN2 flow rate
LN2 flow
Simplified LN2 shield pipe
L
What flow rate do we need to maintain the shield at about ≈ 80 K?
r
Differential fluid element
HL [W/m]
dx

## Slide 58
LN2 flow
Simplified LN2 shield pipe
L
r
Differential fluid element
HL [W/m]
dx

## Slide 59
LN2 flow
Simplified LN2 shield pipe
L
r
Differential fluid element
HL [W/m]
dx

## Slide 60: Scaling laws of LN2 pipes vs. Cu cables
Cu cable
LN2 pipe

## Slide 61: Scaling laws cont.
Cu cable
LN2 pipe

## Slide 62: Advanced LIGO Layout
FI
Signal Recycling Cavity
ITM
ETM
Input
Mode
Cleaner
Output
Mode
Cleaner
Power Recycling
BS
4 km
Laser
Φ
m
PD
GW readout
FI
ITM
ETM
125 W
5.7 kW
815 kW
CP
Installation complete in 2014
Observing GWs in 2015

## Slide 63: How to get a LN2 Hose to ST2
Stage 0
Stage 1
Stage 2
Stage A
LN2 hoses: Supply and return
Actuators
Sensors
Extra stage, A, in parallel with stage 1 carries hose. Stage A is actuated to follow stage 2 so the hose has does not short seismic isolation. Stage A sensor noise is set by the stage 2 isolation requirement (so it follows stage 2 and not the sensor noise).
Vacuum flange
clamps

## Slide 64: Test Mass Radiation Simulation
Ideal radiation shield:
Fully encloses test mass
77 K (liquid N2)
Shield emissivity = 1
Test mass barrel emissivity = 0.5
Laser power = 0
Power α (Temperature)^4

## Slide 65: Dewar pressure during measurements

## Slide 66: Effect of pressure on test mass temp

## Slide 68: N2 gas therm. cond. vs pressure

## Slide 69: Temperature Sensor Locations

## Slide 70: Test Mass Temperature Equations

## Slide 71
ref: Harry, Bodiya, Desalvo. Optical Coatings and Thermal Noise in Precision Measurement. 2012. pg 113.

## Slide 72: Si CTE vs Temperature
Thermoelastic component of thermal noise goes to zero with CTE.

## Slide 73: Si Thermal Conductivity vs Temp.

## Slide 74: Si Specific Heat vs Temperature

## Slide 75: Thermal Conductivity of Materials
References: Stainless: NIST, CRYOGENICS TECHNOLOGIES GROUP, http://cryogenics.nist.gov/
Al, Cu, Si: Thermal Conductivity: metallic elements and alloys, Touloukian 1970
N2 gas: Thermal conductivity: nonmetallic liquids and gases, Touloukian 1970

## Slide 76: Other Problems To Solve
Flexibility if liquid N2 hoses or Cu cables
Temperature/height control of blade springs
Test mass temperature control
Test mass temperature tolerance
How to measure temperature?
  Measure acoustic modes – Young’s modulus is temp. dependent
  Measure test mass diameter – combined with CTE data gives temperature
  Infrared camera
Emissivity of optical coatings
Lossiness of emissive coatings
Good emissivity estimates/measurements of Si?
Power absorption in Si and Si coatings (ppm, W, etc)?
How noisy is flowing laminar liquid nitrogen: seismic, Newtonian?
Optical coating thermal noise at 124 K
How to actuate the test mass – is the ESD out?
Can we put viewports in the heat shield?