H1 TMSX Telescope focal tuning results
Keita Kawabe
LIGO-G1400514

---
Relevant documents

---
TMS Telescope
Change from H1 TMSY (pilot) that is applicable to all other units including H1 TMSX
Changed a vendor of some of flat optics for coating quality.
- Different spec for the primary optic thickness (thinner than pilot unit) though the vendor is the same.
  - Spec for the curved surface didn't change.
- No change in the tuning target parameter nor procedure.

---
Source Parameters (MM results)
Source Parameters (MM results)
IR beam from the source was measured using Mode Master (MM). MM to the last lens of launcher telescope is 75.8cm.
Relevant part of the measurement file (0725_src_doublet.std) is shown below. The RED marked values are used for the model.
[EXTERNAL RESULTS]
Min	Max	Mean	Std Dev	Dim
---------------------------------------------------------------------------------
M≤x	1.09	1.11	1.10	0.007	-
M≤y	1.12	1.14	1.13	0.006	-
M≤r	1.11	1.12	1.11	0.005	-
2Wox	4.212	4.673	4.480	0.1526	mm
2Woy	4.061	4.505	4.325	0.1544	mm
2Wor	4.137	4.590	4.403	0.1531	mm
2Wex	5.726	5.786	5.744	0.0190	mm
2Wey	5.748	5.806	5.784	0.0217	mm
2Wer	5.742	5.790	5.764	0.0167	mm
Zox	10.998	10.427	10.774	-0.1935	m
Zoy	11.011	10.636	10.839	-0.1257	m
Zor	11.005	10.538	10.808	-0.1500	m
Zrx	11.860	14.720	13.488	0.9208	mm
Zry	10.755	13.362	12.260	0.8669	mm
Zrr	11.287	14.017	12.851	0.8917	mm
Divergencex	0.32	0.36	0.33	0.012	mr
Divergencey	0.34	0.38	0.35	0.013	mr
Divergencer	0.33	0.37	0.34	0.012	mr
Astigmatism(Zoy-Zox)/Zrr	-0.4	2.3	0.5	0.86	%
Waist Asymmetry(2Woy/2Wox)	0.958	0.976	0.965	0.0053
Divergence Asymmetry Thetay/Thetax	1.051	1.071	1.062	0.0061

---
Tele Measurement - 3 (Single Lens)
Tele Measurement
IR beam from the Pedestal breadboard into the Tele, retro-reflected by the 8” flat ETM, back to the breadboard into the MM.
The RED marked values are used in the model are return values.
[EXTERNAL RESULTS]
    Min	Max	Mean	Std Dev	Dim
---------------------------------------------------------------------------------
M2x	1.17	1.28	1.23	0.037	-
M2y	1.03	1.16	1.07	0.040	-
M2r	1.32	1.42	1.36	0.033	-
2Wox	0.637	0.671	0.652	0.0128	mm
2Woy	0.630	0.679	0.655	0.0140	mm
2Wor	0.742	0.797	0.768	0.0171	mm
2Wex	2.209	2.387	2.285	0.0580	mm
2Wey	1.263	1.469	1.320	0.0594	mm
2Wer	1.817	1.929	1.867	0.0378	mm
Zox	-0.827	-0.890	-0.856	-0.0209	m
Zoy	-0.494	-0.562	-0.515	-0.0208	m
Zor	-0.687	-0.741	-0.709	-0.0178	m
Zrx	0.245	0.261	0.255	0.0059	mm
Zry	0.276	0.312	0.295	0.0113	mm
Zrr	0.307	0.332	0.320	0.0081	mm
Divergencex	2.48	2.63	2.56	0.046	mr
Divergencey	2.12	2.35	2.22	0.069	mr
Divergencer	2.36	2.45	2.40	0.026	mr
Astigmatism(Zoy-Zox)/Zrr 99.1	112.9	106.3 4.46	%
Waist Asymmetry(2Woy/2Wox)	0.960	1.062	1.005	0.0322
Divergence Asymmetry Thetay/Thetax	0.813	0.945	0.869	0.0388

---
Telescope de-tuning Results
(using TeleScan3 measurements)
Telescope de-tuning Results
Input into the TeleModel are the measured Source Parameters from the ModeMaster.
The GREEN markers are from the measured ModeMaster results from ‘Tele Measurement 3’ (they should be crossing the red calculated line).
The vertical lines indicate the tolerance.
X-axis
Y-axis