## PUNCH LIST
# | Priority | Description | Holding this review? | Punch-list | Assigned or Action | Installation Subset | Asignee | Due | Notes
1 | H | Electronic Rack Designations need to be filled in for all chambers/systems. | No | Yes | Project | All Chambers | Richard McCarthy | 2015-02-27 | Keith Thorne has listed the racks, docs should be updated and Keith's info cross-checked.
3 | H | Obtain pointers to all available oplev testing documents | Yes | Yes | Project | all oplevs | Eric Gustafson | 2015-02-27
4 | H | Add LHO viewport summary document to all relevant chamber instance docs. Should include Leak and Pressure testing. | Yes | Yes | Project | VP | Gerardo Moreno | 2015-02-27
5 | L | Complete as-built IO table drawings and enter into DCC record D1300357 | No | Yes | Project | IO | Cheryl Vorvick | 2015-02-27
10 | L | Record any deviations from viewport install plan T1000746 | No | Yes | Project | AOS | Gerardo Moreno | 2015-02-27
11 | L | IAS documents should note that deviations in the as-built location of the chambers may be large enough to affect the IAS tolerances | No | Yes | Project | AOS | Doug Cook | 2015-02-27
15 | L | alog #10593 indicates an installation issue on WHAM4. Fix should be documented for future installs. | No | Yes | Project | SEI | Hugh Radkins | 2015-02-27
16 | L | HEPI testing reports are missing for most (if not all) chambers. | No | Yes | Project | SEI | Hugh Radkins | 2015-02-27
21 | L | Is there an equivalent to D1300347 in ICS for H1? (This is the PSL table IO layout). | No | Yes | Project | PSL | Mike Landry | 2015-02-27
22 | L | Item #10 from the list at LHO alog #3060 needs to be completed: "Set up final configuration power/rf supplies (cables need to be run too)" | No | Yes | Project | PSL | Peter King | 2015-03-10
32 | L | There is a temporary T240 installed at the LHO X-end instead of a STS-2 | No | Yes | Project | SEI/WBSC9 | Mike Landry | 2015-02-27
36 | L | HWSY has the wrong in-vacuum lens. There is a temporary fix in place, but testing is not complete. | No | Yes | Project | AOS/TCS | Aidan Brooks | 2015-02-27
37 | L | SUS Hardware Watchdog is not yet implemented. | No | Yes | Project | SUS/SYS | Rolf Bork | 2015-02-27
38 | L | E1000708 contains "TBD" heights. Are they recorded somewhere? (LLO doc has same issue) | No | Yes | Project | VE | John Worden | 2015-02-27
39 | L | The Timing Comparator and GPS receiver need to be installed in the H1 MSR | No | Yes | Project | DAQ | David Barker | 2015-02-27
40 | L | Atomic clock needs to be installed and used to verify timing system. | No | Yes | Project | DAQ | Dennis Coyne | 2015-02-27
41 | L | H1 TCS and ISC racks have different designations than L1 (and documentation). This should be fixed. | No | Yes | Project | DAQ | Filiberto Clara | 2015-02-27
42 | L | The S-number records in the DCC for H1 racks are incomplete | No | Yes | Project | DAQ | Filiberto Clara | 2015-02-27

## CLOSED PUNCHLIST ITEMS
# | Priority | Description | Holding this review? | Punch-list | Assigned or Action | Installation Subset | Asignee | Due | Notes
2 | L | E1200566-V6 is heavily redlined. This should be updated. | No | Yes | Project | WHAM2 | Calum Torrie | 2014-12-08 | Red-Lines are acceptable. They have captured information. If Operations chooses to do so they will pretty up this document.
6 | L | Update redlined oplev floor-occupancy doc - G1000740-V2 | No | Yes | Project | oplevs | Eric Gustafson | 2015-01-31 | Red-Lines are acceptable. They have captured information. If Operations chooses to do so they will pretty up this document.
7 | L | Produce pier drawing PDF from D1001291 and post to D1001291 | No | Yes | Project | oplevs | Eric Gustafson | 2015-01-31 | Refer to LIGO-D1001334-v2 for pdf. Note added to D1001291
8 | L | Every WHAM is missing ICS records for HEPI. | No | Yes | Project | WHAM1 and SEI | Mike Landry | 2015-01-31 | ICS for key components as per ICS user guide. We checked and HAM HEPI's are in ICS, look unde WHAM2 i.e. D0901083. Also use HEPI test reports as reference.
12 | L | Does WHAM2 ICS record have all major components (LHAM2 record is longer) | No | Yes | Project | SEI | Mike Landry | 2015-01-31 | We checked and key componets are in the WHAM2 record.
13 | L | Installation Procedure (E1200450), refers to D1101249, which it notes is "not currently available for H1. In the mean time use L1 documents". Implies LHO are awaiting this doc, but is being worked on? or can LHO just use the LLO doc in it's place? | No | Yes | Project | SYS | Calum Torrie | 2015-01-31 | Red-lines are aceptable. They have captured information. If operations chooses to do so they can make it pretty.
14 | L | As-built/Aligned Procedure, E1200470, provides clear data for PR2 (alog 5784), however, MC2 (alog 4811) is seriously lacking. Is this data available? | No | Yes | Project | AOS | Doug Cook | 2015-01-31 | Stuart Aston updated WHAM3 document with alignment numbers.
17 | L | WHAM4 modal testing results needed (if any) | No | Yes | Project | SEI/SUS | Mike Landry | 2015-01-31 | Modal testing was done on SR2. Results detailed in E1400142, the acceptance report for SR2.
18 | L | HWS alignment and tests on WHAM4 need to be documented | No | Yes | Project | AOS | Aidan Brooks | 2015-01-31 | Closed after receiving information from A. Brooks on which aLOGs were relevant.
19 | L | OMC ICS record for WHAM6 is not complete. Missing ISC components. | No | Yes | Project | SUS/ISC | Calum Torrie | 2015-01-31 | Calum via email thinks this looks ok. It may have some incompleteness, but will go with SYS verdict and close item. (BO'R)
23 | L | Where are the test results for the PSL table legs? | No | Yes | Project | PSL | Mike Landry | 2015-01-31 | Added links to iLIGO pre-aLIGO test of table legs, plus H1/L1/H2 testing links.
25 | L | Do we need a H1 PSL issue tracker in the integration issues list? L1 does not have one, perhaps it should too. Otherwise how do we track issues that are listed in PSL installation acceptance document? | No | Yes | Project | PSL | Mike Landry | 2015-01-31 | Consensus is that PSL issues are handled well w/o a separate issue tracker.
24 | L | Has the high power beam dump been tested? Where are the results? | No | Yes | Project | PSL | Mike Landry | 2015-01-31 | The high power beam dump was tested at LLO and this test is deemed sufficient for the project phase acceptance.
26 | L | Edit Install procedure for WBSC1 and remove redlines (E1200915-V1) | No | Yes | Project | SYS/WBSC1 | Calum Torrie | 2015-02-27 | Red-lines are aceptable. They have captured information. If operations chooses to do so they can make it pretty.
27 | L | No modal test results for the Cryopump Baffle on the vertex X-arm or Y-arm | No | Yes | Project | AOS/WBSC1/WBSC3 | Mike Landry | 2015-01-31 | This was a missed opportunity. We rely on L1 similarity.
28 | L | No TCS testing data/documentation for WBSC1-CO2P or WBSC3-CO2P | No | Yes | Project | WBSC1/WBSC3/TCS | Mike Landry | 2015-01-31 | Document updated with links provided by A. Brooks.
29 | L | No alignment documentation for WBSC3 CO2P | No | Yes | Project | WBSC3/TCS | Mike Landry | 2015-01-31 | Document updated with links provided by A. Brooks.
30 | L | E1200971 has broken links in TOC and many redlines (missing docs, refers to L1 etc.). Needs updating. | No | Yes | Project | SYS/WBSC2 | Calum Torrie | 2015-02-27 | Red-lines are aceptable. They have captured information. If operations chooses to do so they can make it pretty.
31 | L | Test results for the BSC-ISI in WBSC2 are missing. E1100847 is empty | No | Yes | Project | SEI/WBSC2 | Mike Landry | 2015-01-31 | Now posted. Closed.
33 | L | AOS/TCS/RHx on WBSC9 and WBSC10 need acceptance testing measurements | No | Yes | Project | AOS/WBSC9/WBSC10 | Aidan Brooks | 2015-02-16 | Acceptance testing of RH in LHO aLOG #16654
35 | L | Was B&K testing for the CMB on WBSC9 and WBSC10 waived? | No | Yes | Project | AOS/WBSC9/WBSC10 | Mike Landry | 2015-02-27 | This was a missed opportunity. We rely on L1 similarity.

## PUNCHLIST ITEMS FOR OPS
# | Priority | Description | Holding this review? | Punch-list | Assigned or Action | Installation Subset | Asignee | Due | Notes
9 | L | CPS readout boxes beside each chamber need an official designation | No | Yes | Ops | SEI | Calum Torrie | 2015-01-31 | Ben Abbott is working on this.
34 | L | TMSX, no B&K measurements have been taken | No | Yes | Ops | SUS/TMSx | Mike Landry | 2015-02-27 | Will be done during an incursion if time/circumstances permit/warrant.

## WHAM1
| DONE | PUNCH | ONES MARKED PUNCH NOW ADDED TO PUNCH LIST
| TO DO | NOTE | THE ONES MARKED NOTE ARE FOR AUTHORS TO CONSIDER. THE ONES MARKED PUNCH WERE MOVED TO PUNCH LIST
| Chamber | Section | Paragraph(s) | Comment | Calum's Remarks 5/7/14 | BML reply
WHAM1 NOTES FROM KEITH
Changes 2014-12-01 for v2
NOTE | Section 1 "covers the HAM chamber LHAM1" to "covers the HAM chamber WHAM1"
PUNCH | Section 1 "Electronics Racks Designation"
Just said TBD. I went ahead and entered all the racks
NOTE | Section 1 "Optical Table/Enclosure Designation"
Removed hyper-link to drawing (moved to Section 3)
NOTE | Section 2: "Initial Alignment Procedure"
Needs an (empty) card-file entered in DCC
NOTE | Section 2: "As-Built/Aligned Procedure"
Despite claims, I found the following aLOG entries regarding alignment,
noting coordination with IAS. These methods should be documented in a procedure.
ALOG 4131 - WHAM1 Support Table Coarse Alignment
ALOG 4223 - SEI Optical Table in Place - report coordination with IAS
ALOG 6689 - WHAM1 Support Table/Optical Leveling (Survey notes)
ALOG 8057 - WHAM1 Optical Table Level/Elevation checked/set
NOTE | Section 3: Electronics Rack Drawings:
I also went ahead and entered hyperlinks for the drawing documents for all the racks.
NOTE | Section 3: Optics Table/Enclosure Drawings:
Inserted document link previously in Section 1
NOTE | Section 4: Serial Number Records
Again, I created entries for each rack(s) drawing, then DCC link to S-number.
None of these are yet in ICS (but my search technique may not be sufficient)
NOTE | Also, added the Optical Table (not in ICS either)
PUNCH | Section 5: Testing
** Still no records on viewport testing on WHAM1
NOTE | Section 7:
** Issue 761 should be removed. It only covers inspections at LLO.
Given Issue 846, I suspect the inspection regime in 761 for LLO should be done at LHO.
WHAM1 Note from Hugh
NOTE | LHO alog 8057 for final level & details (Comment made in email to Calum and circled in PDF of install document section 2)

## WHAM2
| DONE | PUNCH | ONES MARKED PUNCH NOW ADDED TO PUNCH LIST
| TO DO | NOTE | THE ONES MARKED NOTE ARE FOR AUTHORS TO CONSIDER. THE ONES MARKED PUNCH WERE MOVED TO PUNCH LIST
WHAM2 NOTES FROM CALUM
PUNCH | Electronic Rack Designation (page 1) - LHO rack designations TBD - _ACTION - THIS SHOULD BECOME A PUNCH-LIST ITEM. _
PUNCH | 1) Electronic Rack Designation (page 1) - Note on CPS readout boxes does apply to H1. (SYS working this and design of shelf for these)
NOTE | 2) Optic Table Enclosures (page 1) - Believe fine as is so note on "update for H1?" can be removed.
NOTE | 3) As-built Install procedures (page 2) - Note associated with LIGO-E1200566-v6 <https://dcc.ligo.org/LIGO-E1200566> should be improved to explain that this is the WHAM2 top level install procedure. Seems to be complete list of alog entires.
NOTE | 4) Alignment procedures (page 3) - all seems complete although QUESTION - _Are there any logs for alignment work?_ (Perhaps question for Doug / Jason?)
NOTE | 5) Drawings (page 3) - all seems complete apart from Rack drawings which is an action above. Oh and also - the note on Pr2 Op Lev "no doc in filecard" can be removed. D1001166 is the parent of D1002207 which is a related doc of D1001166 and has drawings with detail.
NOTE | 6) Serial Number Records (page 4) - Confirmed top level present and populated. Re note on "Electronic rack info needed here". Refer to punch-list item above.
NOTE | 7) Testing (page 4 and 5)
PUNCH | i) Viewports - _Is there an LHO equivalent to E1200445?_
PUNCH | ii) Op Lev - _Are there completed reports filed on log or DCC?_ (Perhaps question for Eric before decide if punch-list or missed opportunity.)
NOTE | 8) Installation Completeness (page 5) - This is good to see!! I also agree with it.
NOTE | 9) Bugs (page 5, 6, 7 and 8) - _what is plan with bugs? In doc or on bugzilla?_ I haven't reviewed this list yet.
WHAM2 NOTES FROM BRIAN
NOTE | D1100170 in Electronics Rack Designation(s) in 1. is called obsolete in DCC notes
PUNCH | - E1200566-V6 has a lot of redlines still. Either this goes to the punchlist or should be finished soon.
NOTE | - elog #3182 for baffles mentions a galled screw and a broken fastener. Should we record this on issues or some other note? - No, standard waiver from MRL
NOTE | - alog #3980 is for the output MC baffle. #3182 seems to cover both IMC baffles.
PUNCH | - alog #4714 on Viewport installation is only for the last viewport. The corresponding entry at LLO referred to the document that was
followed. There was also a discussion/confusion about viewports when we did this at LLO. Did LHO follow LLO, try to follow the
document or something else?
PUNCH | - alog #4465 seems to indicate that HEPI install messed up the alignment. Is this true? Was corrective action taken, was alignment re-checked?
NOTE | - "closeout LHAM2" should be "closeout WHAM2"
NOTE | - alog #8088 appears to not be relevant.
NOTE | - Is there an alog entry for "#xxxx". This is for PR2 and LHAM2 (should be WHAM2) OpLevs.
NOTE | - What does "early version" mean relevant to E1200470. The DCC indicated V5 is the asbuilt.
NOTE | - D1300357 is an empty filecard (under 3. Drawings).
NOTE | - G1000740 appears to be unofficial. It is also heavily redlined.
NOTE | - D1001291 only has an edrawing, no pdf. -Now has html too
NOTE | - In the ICS record under serial numbers the ISI unit for WHAM2 is #4. Is this correct? I only ask because at LLO unit #1 was
also the first one installed (in LHAM3, LHAM2 got #2).
PUNCH | - The WHAM2 ICS entry is much shorter than LHAM2. This is probably a M. Heintze effect, maybe it isn't critical to have all
clamps, dumps etc., although it is nice. Someone should do a careful comparative check (maybe I will!).
PUNCH | - The leak and pressure testing of viewports def. applies to LHO as well as LLO. The new viewports should have been tested
prior to installation. If it wasn't performed this is a big deal. But my guess is that it was.

## WHAM3
| DONE | PUNCH | ONES MARKED PUNCH NOW ADDED TO PUNCH LIST
| TO DO | NOTE | THE ONES MARKED NOTE ARE FOR AUTHORS TO CONSIDER. THE ONES MARKED PUNCH WERE MOVED TO PUNCH LIST
WHAM3_acceptance_notes_2014-12-01 (S.Aston)
Changes requested from E1400440-v1 to v2
Section 1
NOTE | Change "covers the HAM chamber LHAM3" to WHAM 3 in the introduction.
NOTE | Electronics Rack Designations have now been populated.
Section 2
PUNCH | Installation Procedure (E1200450), refers to D1101249, which it notes is "not currently available for H1. In the mean time use L1 documents". Implies LHO are awaiting this doc, but is being worked on? or can LHO just use the LLO doc in it's place?
PUNCH | As-built/Aligned Procedure, provides clear data for PR2 (5784), however, MC2 (4811) is seriously lacking. Is this data available elsewhere, or can it be posted to the aLOG?
Section 3
NOTE | Electronics Rack Drawings have now been populated.
Section 4
NOTE | If exists, ICS records for HEPI need to be linked here.
Section 5
PUNCH | Viewport leak and pressure tests need to be included.
NOTE | I understand that B&K hammering of scrapper baffles has not been carried out.
Section 7
NOTE | I noticed there was an open integration issue regarding a cross-coupling observed on MC2 (Bug#847), which has now been mitigated and I requested be closed.

## WHAM4
TO DO | THE ONES MARKED NOTE ARE FOR AUTHORS.
DONE | THE ONES MARKED PUNCH WERE MOVED TO PUNCH LIST
Review comments from Marty Levine were added to E1400441-V1 filecard in the DCC under "Other Files". Here are some of Marty's explicitly noted.
PUNCH | alog #10593 indicates an installation issue. Fix should be documented for future installs.
PUNCH | HEPI test report E1300830 is an empty filecard. Report needs to be added.
PUNCH | Need modal testing results for baffles. Note originally this referred to an LHAM4 LLO alog (#7144) entry. That entry is for LHAM5.
PUNCH | Need HWS alignment and measurement notes
Comments from Aidan. These were used to close punchlist #18 and open #36
Item 18: HWS Alignment and tests on WHAM4 need to be documented:
Response: The alignment of HWSX and HWSY to the ITMs, necessary for aLIGO Installation, was completed around Sept/Oct 2014 as documented below.
Subsequent alignment changes in the core optics in the SRC disrupted the HWS alignments. The correction of the HWS alignments is an Operations activity.
Alignment process:
1. [24-Sept 2014]: Leakage of ALS green beam from WHAM4 makes it to the HWS table (https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=14123)
2. [24-Sept-2014]: HWSX (~800nm) probe beam matched to ALS axis and return beam from ITMX is recorded. HWSY in-vacuum lens discovered to be wrong (https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=14135)
3. [25-Sept-2014]: HWSX alignment is fine-tuned using without the Hartmann plate on the camera. A slight mis-alignment is introduced to remove the stray probe beam returning from ETMX. (https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=14158)
4. [26-Sept-2014]: HWSX is moved to within ~9mm of conjugate plane of ITMX. (https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=14171)
(https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=14180)
5. [26-Sept-2014]: Confirmation of measurement of RH thermal lens with HWSX (https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=14172)
6. [29-Oct-2014]: Calibration of HWSX RH thermal lens & localization of ITMX center in HWSX coordinate system (https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=14698)
7. [27-Oct-2014]: HWSY fix installed. Confirmation of HWSY probe beam return from ITMY. (https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=14648)
8. [28-Oct-2014]: Fixed probe beam leakage between HWSY and HWSX (https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=14683)
9. No calibration of HWSY yet.
PUNCH So, the LHO HWS are installed, but HWSY is not installed as designed due to the wrong in-vacuum lens. The existing HWSY system has a
temporary fix to give us some sensitivity and is aligned to ITMY but not 100% tested.

## WHAM5
TO DO | THE ONES MARKED NOTE ARE FOR AUTHORS.
DONE | THE ONES MARKED PUNCH WERE MOVED TO PUNCH LIST
WHAM5 NOTES FROM BRIAN
PUNCH | LHO Rack Designations is in punchlist
NOTE | SR3 Install is alog #11754, harder to search for SRM and OFI. aLOG entries for SRM and OFI found by Mike L.
NOTE | E1200954-v4 appears to be a comprehensive account of IAS. Red text can be removed.
PUNCH | Note on E1200954 says there was an error in WHAM5 position. Should be mentioned or discussed (refer to an alog?)

## WHAM6
TO DO | THE ONES MARKED NOTE ARE FOR AUTHORS.
DONE | THE ONES MARKED PUNCH WERE MOVED TO PUNCH LIST
WHAM6 NOTES FROM CALUM
My notes on WHAM6 i.e. on LIGO-E1400443-v1 <https://dcc.ligo.org/LIGO-E1400443>
PUNCH | 0) Page 1 - Rack stuff. Discussed before. Action.
NOTE | 1) Page 1 - /"install procedure"/ - what about LIGO-E1200099-v4 <https://dcc.ligo.org/LIGO-E1200099>?
NOTE | 2) Page 2 - /"baseline or initial alignment"/ - if note on LIGO-E1200955-x0 <https://dcc.ligo.org/LIGO-E1200955> is true then LIGO-E1200955-x0 <https://dcc.ligo.org/LIGO-E1200955> should be updated to match.
NOTE | 3) Page 3 - /"D0900421"/ - need to update as D0900421 is for HAM4.
NOTE | 4) Page 3 - /"D1000342. Note that this ICS entry seems to have no entries for ISC"/ - not true I checked and there are entries. : However I (ML) changed this to the H1 version D0901822, which does not have ISC ientries.
BO'R just checked this ICS entry and I find ISC entries in the record.
NOTE | 5) Page 3 - I note there is also a aLIGO WHAM6-H1 Top Lev Cham Assy ICS record (populated) and not listed https://ics-redux.ligo-la.caltech.edu/JIRA/browse/ASSY-D0901822-NA.
NOTE | 6) Page 4/5 - Testing - E1200445 is LLO doc. Where is LHO one?
NOTE | 7) Page 4 - See note above re ICS. Not true it is populated.
NOTE | 8) Page 5 - Bugs not reviewed.

## WBSC1
DONE | PUNCH ITEMS NOW INTEGRATED INTO PUNCH LIST
TO DO | THE ONES MARKED NOTE ARE FOR THE AUTHORS TO CONSIDER.
Comments/Edits by BO'R
PUNCH | The installation procedure E1200915 has many redlines that should be fixed.
NOTE | Need to add designations for Electronics racks, including remote rack(s)
NOTE | Need links to drawings for OpLev and Baffles.
NOTE | Need links for CMB install
NOTE | Need links for TCS CO2P alignment
NOTE | Missing ICS record
NOTE | Many links are for LBSC1, need to be change to WBSC1 appropriate ones.
PUNCH | No modal test results for the Cryopump Baffle on the vertex Y-arm
PUNCH | No TCS testing data for WBSC1-CO2P

## WBSC2
DONE | PUNCH ITEMS NOW INTEGRATED INTO PUNCH LIST
TO DO | THE ONES MARKED NOTE ARE FOR THE AUTHORS TO CONSIDER.
Review of WBSC2 (E1400461) K. Thorne 12/16/14
Section 1
NOTE: Electronics Racks not done (as expected). Started E1400461-v2 by adding racks. [in V2]
Section 2
NOTE - Baseline Installation Procedure - Hyperlink should be to specific version (E1200971-v1) [in V2]
NOTE - As-Built Installation Procedure - E1200971-v3 has wrong rev # in header (should be v3) [in V2] - Good catch, but I recommend ignoring it. Made note in metadata for V3 (BO'R)
PUNCH ** Table of Contents in E1200971-v3 PDF has missing links - needs to be updated
PUNCH ** E1200971-v3 notes deficiencies (in red) in documents it cites. Where are we tracking that those get corrected?
NOTE - Baseline Alignment Procedure - Hyperlink should be to specific version (E1200975-v1) [in V2] - The link is to the as-built alignment, so I think this is good (V6) (BO'R)
Section 3
NOTE Electronic Rack Drawings - missing, here is what should be there [in V2]
H1-SEI-C5 D0901301 BSC SEI System Wiring Schematic
H1-SUS-C5,C6,R2 D1100022 aLIGO SUS ITM and BS System Wiring Diagrams
NOTE Section 4 Serial Number Records
- removed HEPI as include in ICS for D0901142 [in V2]
- added rack Records [in V2]
D0901301 H1-SEI-C5 S1301866
D1100022 H1-SUS-C5 S1301872
D1100022 H1-SUS-C6 S1301873
D1100022 H1-SUS-R2 S1301887
NOTE Viewport serial numbers are in T1200235
Section 5 Testing
NOTE SEI ISI: ICS record identified BSC-ISI Unit#3 as installed in WBSC2, updated ref to E100990[in V2]
PUNCH ** SEI HEPI: E1300839 is still empty
PUNCH ** Viewports - Phase 3 (in-situ testing) is Integration Bug 761 (no LHO testing yet)
PUNCH ** AOS/OptLev - Phase 2 testing - May be linked to E1200094 somewhere
NOTE ** Electronics Testing? - Not in scope for this document, maybe in DAQ? (BO'R)
Section 6
NOTE ** Is everything installed? (I think it is - need definitive from LHO)
Section 7
NOTE Issues: Mostly closed - #989 is overall issue tracker [in V2]

## WBSC3
DONE | PUNCH ITEMS NOW INTEGRATED INTO PUNCH LIST
TO DO | THE ONES MARKED NOTE ARE FOR THE AUTHORS TO CONSIDER.
The WBSC3 issues were very similar to the WBSC1 ones. Instead of separate comments the WBSC3 document was edited with fixes and the punchlist items were changed to include WBSC3

## WBSC9
DONE | PUNCH ITEMS NOW INTEGRATED INTO PUNCH LIST
TO DO | THE ONES MARKED NOTE ARE FOR THE AUTHORS TO CONSIDER.
WBSC9_acceptance_notes_2014-12-17 (S.Aston)
Change list from E1400463-v1 to v2
Section 1 (Installation)
NOTE - Electronics Rack Designations have now been populated for SUS and ISC. But, SEI & TCS are still missing.
PUNCH - Noted that there is a temporary T240 Ground Seismometer at X end, in place of a STS-2 D1400077.
Section 2 (Procedures)
NOTE - Added OpLev aLOG links for installation, alignment and calibration.
Section 3 (Drawings)
NOTE - Added Electronics Rack Drawings for SUS, SEI and ISC.
Section 4 (Serial Numbers)
NOTE - Added ICS link for HEPI, D0901150, sub-assembly there of.
Section 5 (Testing)
PUNCH - Viewports, removed link to old ICS record E1300447, replaced with more suitable T1200363, but still needs updating, and test status not shown.
NOTE - AOS/ACB, added photodiode tested in-situ using baffle OpLev alignment & calibration tests, aLOG link 10331.
PUNCH - AOS/TCS/RHx, added initial testing link to aLOG 9272, but need acceptance testing measurements.
PUNCH - TMSX, confirmed that no B&K measurements have been taken.
NOTE - ESD, added aLOG links to ESD initial and TF tests.
Section 6 (Installation Completeness)
NOTE - ICS assembly records to be updated.
Section 7 (Integration Issues / ECRs)
NOTE - I noticed there was an open integration issue regarding a noisy channel on the ETMX L2 UL OSEM (Bug#929).

## WBSC10
DONE | PUNCH ITEMS NOW INTEGRATED INTO PUNCH LIST
TO DO | THE ONES MARKED NOTE ARE FOR THE AUTHORS TO CONSIDER.
Review comments from Marty Levine were added to E1400464-V1 filecard in the DCC under "Other Files", S. Aston transcribed these and other issues to the "Notes" field for E1400464-V4.
There were issues that made it into the punchlist. I list these here:
PUNCH - AOS/TCS/RHx, added initial testing link to aLOG 9272, but need acceptance testing measurements.
PUNCH - Was B&K testing for the CMB on WBSC10 waived?
v1 is first draft by Norna, using LLO equivalent document E1400082 as a template. Items in red need addressed. This includes Bug list.
v2 changes (18 Dec 2014). These are based on updates which Stuart made in v2 version of E1400463 (WBSC9) and comments provided by Marty Levine (see his redlined comments under Other Files above.
Section 1:
a) Electronics rack designations with DCC links added for SUS, ISC, and TCS. SEI rack designation missing.
b) Links added to drawings for the cryo-pump baffle, optical lever and photon calibrator, and ISCTEY table. Link to drawing showing position of ground seismometer at end Y added.
Section 2 :
a) Removed mention of B and K hammering (it appears under testing).
b) Removed one of alog links to ACB install and made clearer the other link notes the installation has taken place.
c) Added Oplev install alog. Could not track down alignment alog.
Section 3:
a) Added ( repeated) the electronic rack designations here (Rich A’s suggestion) and added rack layout links and SEI wiring schematic.
SEI rack designation missing.
b) Updated note with respect to PCal pier assembly.
Section 5
a) Changed viewport link to more suitable document. However the document needs updating.
b) Ring heater testing. Contacted Aidan – this is now on his to-do list.
c) Added links to alogs on ESD testing in vacuum and added link to general ESD documentation page.
Section 7.
Removed the LLO list that was here and replaced it with the appropriate link to the WBSC10 issue tracker (at bug 986).
v3 changes (22 Dec 2014) - added in SEI electronics rack info in sections 1 and 3.
v4 changes (30 Jan 2015) S. Aston
- replaced LBSC4 with WBSC10 in introduction to section 1.
- added links to further ETMY OpLev entries in section 2.
- removed statement "This is called TX but probably correct for Y as well " since doc certainly contains both in section 3.
- added link to HEPI ICS record in section 4.
- added TCS acceptance testing doc link to section 5.
- added AOS/ACB baffle photodiodes tested operable link to section 5.
- added AOS/CMB added link to LLO cryobaffle aLOG entry in section 5.

## WPSL
DONE | PUNCH ITEMS NOW INTEGRATED INTO PUNCH LIST
TO DO | THE ONES MARKED NOTE ARE FOR THE AUTHORS TO CONSIDER.
WPSL_acceptance_notes_2014-12-08 (S.Aston)
Changes requested from E1400466-v1 to v2
Section 1 - Installation
NOTE - Electronics Rack Designations. Look correct, but DCC search comes up empty. Do they exist? If so, meta-data links need updating.
Section 2 - Procedures
PUNCH - Punch list for PSL items following completion of installation (#3060). Are any still outstanding?
Section 3 - Drawings
NOTE - Electronics Rack Drawings. DCC links are to LLO racks, need to be updated to LHO racks.
Section 4 - Serial Numbers
NOTE - Assembly DCC file-cards for electronic racks are empty. - Note that electronics is not tracked in ICS like parts are (BO'R).
Section 5 - Testing
PUNCH - Testing of Optics Table legs : documented in PSL doc - ML
NOTE - No documentation found testing (installed) rotation stage.
PUNCH- No documentation found testing (installed) high power beam-dump.
Are any of the above of sufficient concern that it should be elevated to the punch list? (Added some to Punchlist, some were taken care of (BO'R))

## WDAQ
TO DO | THE ONES MARKED NOTE ARE FOR AUTHORS.
DONE | THE ONES MARKED PUNCH WERE MOVED TO PUNCH LIST
Review comments on E1400457 H1 DAQ Installation Acceptance K. Thorne 12/9/14
Scope Question
- Does DAQ include just DAQ or DAQ, PEM, Timing?
Section 1:
Building(s)/Room(s):
List is too extensive - should only include where DAQ components, front-end computers,
IO chassis and timing systems are. Likely should not include LVEA.
(i.e. - The per-chamber acceptance should include up through AA/AI chassis
I am assuming, otherwise we have overlap.
Electronics Racks:
- Needs to be limited to racks with DAQ components. Could be
CER:H1-ISC-C[1,3], H1-SUS-C[2,3,5,8], H1-SEI-C[1-6], <PSL rack>, <OAF rack>
MSR: H1-MSR[1-10]
XEND:H1-ISC-XC[3,4], H1-SUS-XC2, H1-SEI-XC1, H1-DAQ-XC1
YEND:H1-ISC-YC[3,4], H1-SUS-YC2, H1-SEI-YC1, H1-DAQ-YC1
- Missing racks with PSL, OAF IO Chassis, Timing in CER
Section 3:
Electronic Rack Drawings:
- G1001032 is a drawing of LLO - does not have H1 racks.
Should have links to design drawings of racks above (and be the drawings in Section 4)
ISC LVEA Equipment Room Racks (D1001427)
ISC End Station Remote Racks (D1001459)
SUS racks (See links in E1100337)
- aLIGO SUS HAM2 System Wiring (D0902810)
- aLIGO SUS HAM3-4 System Wiring (D1000599)
- aLIGO SUS HAM5-6 System Wiring (D1002740)
- aLIGO SUS ITM BS System Wiring (D1100022)
- aLIGO SUS ETM System Wiring (D1002741)
SEI racks (See links in T1300173)
- aLIGO BSC SEI System Wiring (D0901301)
- aLIGO HAMs 2&3 ISI System Wiring (D1101576)
- aLIGO HAMs 4&5 ISI System Wiring (D1000298)
- aLIGO HAMs 1&6 ISI System Wiring (D1101584)
PSL-C1 (could use D1201118)
OAF-C1 (could use D1200941)
PEM, TCS see D1300226
DAQ racks - Need design drawings for these
Section 4:
- Some drawing numbers are missing.
- End station ISC racks in D1001459
- End station SEI racks in D0901301
- PEM, TCS - D1300226
- Some S-numbers of some racks are missing
- No DAQ rack serial/ S-numbers
- The S-number records are correct, but none of these exist in ICS (based on my searching).
Section 5:
Is testing to include AA/AI electronics or just IO chassis ADC/DAC?
Above comments have been taken care of save for the following items:
PUNCH - | Install Timing Comparator and GPS receiver in MSR
PUNCH - | S-Records are needed for some DAQ racks and some S-number records are empty
PUNCH | H1 TCS and ISC racks have different designations than L1 (and documentation). This should be fixed.
PUNCH | Atomic clock needs to be installed and used to verify timing system.

## WVE
DONE | PUNCH ITEMS NOW INTEGRATED INTO PUNCH LIST
TO DO | THE ONES MARKED NOTE ARE FOR THE AUTHORS TO CONSIDER.
N.B. Items in green are all addressed. : documented in V2, V3 of Vac doc DCC record E1400467
The alog search suggested at various parts of the document returns 82 hits!
Only #3782 appears relevant to vacuum mods.
Several are however relevant to the testing section.
specific alogs are:
840: BSC6 move
877: End station spool delivery (second shipment from GNB)
869: BSC chamber move Y-arm
788: MX and MY spool pieces arrive
596: Shot of H2 diagonal
589: HAM11 move
Seems to be a lot of missing alogs. Some of the referred ones are
specific to H2, not H1. I would have expected alogs for the BSC moves
etc. at the X arm. Some closeouts, septum plate installs etc.
E1000708 contains "TBD" heights. Are they recorded somewhere? (LLO doc
has same issue). ACTTION : created Punchlist item no. 38
The Scott McCormick reference at the beginning of Section 7 should be
removed.
integration issues:
Maybe add: 81, 761, 846, 941 and 949. Either in document or in a
tracker.