## Slide 1: Noise Analysis of OSEM Actuator to Sensor Coupling - G1401205
Coupling observed in measured transfer functions.
AKA, the ‘turnups’.
https://services.ligo-wa.caltech.edu/integrationissues/show_bug.cgi?id=327
These slides will show that the coupling between the OSEM coil and sensor will not harm our noise performance, given certain assumptions.

## Slide 2: Damping model with OSEM coupling
P
-D
B
+
+
e
y
s
P = top mass to top mass plant
D = damping filter
B = OSEM actuator to sensor coupling
y = top mass displacement
e = top mass TF excitation for measurements
f = OSEM force
s = OSEM sensor output
n = OSEM sensor noise
OSEM sensor output
OSEM force output
Top mass displacement
n
f
Assumptions:
The actuation signal couples to the sensor signal
The sensor noise is summed after this coupling
The coupling is linear

## Slide 3: The measured open loop TFs
If
0.004
B = -----------------
(s+1000*2*pi)
This shape has the right phase
to generate the notch feature.
However, LHO alog 1753 shows
an f^1 shape. We’ll see the shape
does not matter for noise in general.
In this case the damping filter D = 0. So the sensor signal from an excitation is.
Simulated top mass L to top mass L TF = (P+B)
The measured top mass to top mass TF is then
The notch sometimes seen in the TFs

## Slide 4: Derivation of closed loop noise transmission with OSEM coupling
* Key Result: OSEM sensor noise to top mass displacement TF with OSEM coupling. This results suggests the shape of B is unimportant, provided (P+B)*D is less than 1 outside the damping band.
Measured closed loop displacement with OSEM sensor
- It helps to reference the block diagram on slide 2 for this slide.
The top mass displacement, y, expressed in terms of s
The damping force, f, expressed in terms of s
Plugging y and f back in
Factoring out the sensor signal s
The measured closed loop TF between sensor noise and sensor
Then plugging gives us the top mass performance

## Slide 5: Simulink check
A closed loop simulation with the coupling term B and damping filter D on slide 6.

## Slide 6: Filters in the Simulink simulation
D
damp =
1e05 s (s+8.162) (s+0.9069)
----------------------------------
(s+15*2*pi)^2 (s^2 + 1.814s + 7.402)
B
coupling =
0.004
--------
(s+1000*2*pi)
Simulated top mass L to top mass L TF = (P+B)
Velocity damping with a two pole roll-off at 15 Hz and a resgain bump at the 0.43 Hz mode.

## Slide 7: Top mass loop gain in simulink
Showing the loop gain P*D to provide some sense of the damping design and performance.

## Slide 8: Simulink result: Top mass displacement from OSEM sensor noise with various gains on the coupling
The OSEM coupling B in general has negligible influence on the top mass displacement. In the extreme case where B is very large, red, the top mass displacement is actually suppressed.