|  |  |
| --- | --- |
| **Test Date** | Mar 4-6, 2015 |
| **Author(S)** | Liyuan Zhang, GariLynn Billingsley |
| **Approval(s)** |  |
| **Specification Doc.** | LIGO-E0900041 | Specification | < 50 ppm |
| **Procedure Doc.** | LIGO-E1000863 | Mean ± Error\* | 41 ± 1 ppm |
| **Conclusion** | Qualified. |

\* Error is the calibration error, which is ~2%.

**Discussions and Comments:**

 **The scan was done over a 160 mm diameter aperture with the beam and step sizes of 1 mm while keeping the arrow on barrel at Y+ direction. The calibration is done by normalizing the AR reflection signal to the signal from a 1” HR mirror (T=70 ppm) and the variation of laser power during scan is monitored and corrected, the result is summarized in Fig.1.**

****

**Fig. 1 ITM06 S2 AR reflection over a 160 mm diameter aperture.**