*LIGO Laboratory / LIGO Scientific Collaboration*

LIGO-T1500404-v1 *LIGO* July 25, 2015

Test Procedure for EOM/AOM Driver

Koji Arai and Daniel Sigg

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

|  |  |
| --- | --- |
| **California Institute of Technology**  **LIGO Project – MS 100-36**  **1200 E. California Blvd.**  **Pasadena, CA 91125**  Phone (626) 395-2129  Fax (626) 304-9834  E-mail: info@ligo.caltech.edu | **Massachusetts Institute of Technology**  **LIGO Project – NW22-295**  **185 Albany St**  **Cambridge, MA 02139**  Phone (617) 253-4824  Fax (617) 253-7014  E-mail: info@ligo.mit.edu |
| **LIGO Hanford Observatory**  **P.O. Box 159**  **Richland WA 99352**  Phone 509-372-8106  Fax 509-372-8137 | **LIGO Livingston Observatory**  **P.O. Box 940**  **Livingston, LA 70754**  Phone 225-686-3100  Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

The following Test Procedure describes the test of proper operation of the EOM/AOM Drive. **The RF output port should always be properly terminated with 50 Ohms.** Further information can be found on the [wiki page](https://awiki.ligo-wa.caltech.edu/aLIGO/EomAomDriver).

S/N \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ Tester \_\_\_\_\_\_\_\_\_\_\_\_\_\_ Date \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

# Test Equipment

* Voltmeter
* Oscilloscope
* Stanford Research SR785 analyzer
* RF Network Analyzer (Agilent AG4395A or equivalent)
* RF Spectrum Analyzer (Agilent AG4395A is OK)
* RF function generator (Tektronix AFG3101, or equivalent)
* AF function generator (Stanford Research DS335/340/345 or similar)
* RF Power Meter (Agilent E4418A, Gigatronics 3410A, or similar)
* D1500091 EtherCAT EOM Driver Test Rig with DB25 Male/Male cable
* DB9 breakout board
* Two dual bench power supply (for +/24Volts, +/ 18 Volts)
* Single bench power supply (for +5V)
* Diagrams / Board Schematics, LIGO [D0900760](https://dcc.ligo.org/LIGO-D0900760), [D0900761](https://dcc.ligo.org/LIGO-D0900761), [D0900847](https://dcc.ligo.org/LIGO-D0900847), [D0900848](https://dcc.ligo.org/LIGO-D0900848)

# Tests

*The EOM/AOM Driver uses the RF AM Stabilization Power (D0900848, rev B) with the RF AM Stabilization Controller (D0900761, rev B) and RF AM Stabilization Servo (D0900847, rev B).*

***Power Supply Test***

1. **Verify the proper current draw with the output switch off:**

Turn off the excitation (left) and output (right) switches. Using a bench DC supply apply +/24Volts and +/ 18 Volts to P1 of the *RF AM Stabilization Power (D0900848) at the front panel*. Measure the current draw of the box. **Confirm the RF output port I sterminated.** Turn on the output switch and measure the current draw of the box.

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Supply**  **Voltage** | **With RF output SW “OFF”** | | **With RF output SW “ON”** | |
| **Current** | **Nominal** | **Current** | **Nominal** |
| +24V | \_\_\_\_\_\_\_\_\_\_ | 0.08A | \_\_\_\_\_\_\_\_\_\_ | 0.53A |
| -24V | \_\_\_\_\_\_\_\_\_\_ | 0.07A | \_\_\_\_\_\_\_\_\_\_ | 0.07A |
| +18V | \_\_\_\_\_\_\_\_\_\_ | 0.29A | \_\_\_\_\_\_\_\_\_\_ | 0.21A |
| -18V | \_\_\_\_\_\_\_\_\_\_ | 0.24A | \_\_\_\_\_\_\_\_\_\_ | 0.26A |

1. **Verify the internal supply voltages:**

On the 2nd PCB of the power board: check the voltage between TP 1 (GND) and the following test points.

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **Test Point** | **Nominal V** | **Measurement** | **Test Point** | **Nominal V** | **Measurement** |
| TP13 | -5V | \_\_\_\_\_\_\_\_\_\_ | TP2 | +24V | \_\_\_\_\_\_\_\_\_\_ |
| TP12 | -15V | \_\_\_\_\_\_\_\_\_\_ | TP3 | +18V | \_\_\_\_\_\_\_\_\_\_ |
| TP11 | -21V | \_\_\_\_\_\_\_\_\_\_ | TP9 | +10V | \_\_\_\_\_\_\_\_\_\_ |
| TP10 | -10V | \_\_\_\_\_\_\_\_\_\_ | TP17 | +24V | \_\_\_\_\_\_\_\_\_\_ |
| TP5 | -18V | \_\_\_\_\_\_\_\_\_\_ | TP14 | +21V | \_\_\_\_\_\_\_\_\_\_ |
| TP6 | -24V | \_\_\_\_\_\_\_\_\_\_ | TP15 | +15V | \_\_\_\_\_\_\_\_\_\_ |
|  |  |  | TP16 | +5V | \_\_\_\_\_\_\_\_\_\_ |

1. **Verify supply OK logic:**

Turn off the external power supply. Disconnect the power supply connector from the front panel. Remove the front panel. Remove the top two board from the inter PCB connector. Expose the D0900848 board. Make sure two boards have proper electrical isolation from the chassis or other conductive material. Connect and turn on the power supply again. If above TP 5, 6, 2, and 3 are correct then pin 5 on U1 (connected to R11) and U4 (connected to R23) should be

**Logic high ~3Volts. Confirm.\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

1. **Verify the relays for the power supply sequencing:**

Turn off +/-24 V. Confirm Pin 5 of K1 and K2 are not energized to +/-18V.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

Turn on +/-24 V again. Confirm Pin 5 of K1 and K2 are now energized to +/-18V.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

1. **Verify noise levels of the internal power supply voltages:**

Turn off +/-24 V. Disconnect the power supply from the front panel connector. Restore the two boards and the front panel.

The noise on the following test points should be measured with a SR785 using an

rms power spectrum.

TP13 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 15 nVrms/sqrt Hz at 140 Hz

TP12 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 30 nVrms/sqrt Hz at 140 Hz

TP11 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 40 nVrms/sqrt Hz at 140 Hz

TP10 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 20 nVrms/sqrt Hz at 140 Hz.

TP9 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 20 nVrms/sqrt Hz at 140 Hz.

TP14 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 30 nVrms/sqrt Hz at 140 Hz.

TP15 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 20 nVrms/sqrt Hz at 140 Hz.

TP16 noise \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ less than 15 nVrms/sqrt Hz at 140 Hz.

***RF Test***

1. **Test output switch functionality:**

Make sure the output switch is off. Connect the RF function generator to the input connector. Provide 9.1MHz signal at the power level of 10dBm. Connect the RF power monitor to the output connector. Set the output adjustment switches to have 13dBm output (i.e. 12dBm & 1.0dBm)

Confirm there is no output at this point.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

Turn on the output switch. Confirm there is the RF output about 13dBm.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

1. **Check stabilization servo stability:**

Turn off the RF output switch and connect the output to the oscilloscope with a 50 Ohm terminator. Turn the RF output again. Confirm the waveform is sinusoidal and there is no obvious amplitude oscillation.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

Observe the TP1 signal with the oscilloscope and SR785. Confirm there is no oscillation.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

Observe the CTRL signal with the oscilloscope and SR785. Confirm there is no oscillation.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

1. **Check exc port functionality:**

Observe the TP1 signal with the oscilloscope. Connect 1kHz 2Vpk signal from the audio frequency function generator to the exc port.

Turn off the exc switch and confirm there is no obvious 1kHz component on the TP1 signal.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

Turn on the exc switch and confirm there is visible 1kHz component on the TP1 signal.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

1. **Openloop transfer function measurement:**

Connect TP1 and TP2 to ChA and ChB of the network analyzer, respectively. Connect the signal generator output of the network analyzer to Exc. Turn on the exc switch.

Sweep the signal of 0dBm from 1kHz to 2MHz (801 line).

Measure the transfer function and check the frequency where the magnitude is 0dB (i.e. UGF: unity gain frequency). Check the phase at the UGF (i.e. P.M.: phase margin). Confirm the UGF is above 100kHz and the phase is more than 30deg.

**Unity Gain Frequency \_\_\_\_\_\_\_\_\_\_\_ (>100kHz nom)**

**Phase Margin \_\_\_\_\_\_\_\_\_\_\_ (>40deg nom)**

1. **Output calibration:**

Turn off the output switch. Connect the output to the power meter. Turn on the output switch again. The power adjustment is set to be 13dBm (12dBm+1dBm). Rotate the trimmer R6 on the front panel until the output becomes 13.00dBm+/-0.01dBm.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

1. **Check the output power with other output values:**

Set the power adjustment to 4dBm+0dBm. Measure and record the RF output power with the power meter, while sweeping the output from 4+0dBm to 26+0dBm.

|  |  |  |
| --- | --- | --- |
| **Setting**  **(dBm)** | **Measurement**  **(dBm)** | **Nominal**  **(dBm)** |
| 4.0 | **\_\_\_\_\_\_\_\_** | 4.0+/-0.1 |
| 6.0 | **\_\_\_\_\_\_\_\_** | 6.0+/-0.1 |
| 8.0 | **\_\_\_\_\_\_\_\_** | 8.0+/-0.1 |
| 10.0 | **\_\_\_\_\_\_\_\_** | 10.0+/-0.1 |
| 12.0 | **\_\_\_\_\_\_\_\_** | 12.0+/-0.1 |
| 14.0 | **\_\_\_\_\_\_\_\_** | 14.0+/-0.1 |
| 16.0 | **\_\_\_\_\_\_\_\_** | 16.0+/-0.1 |
| 18.0 | **\_\_\_\_\_\_\_\_** | 18.0+/-0.1 |
| 20.0 | **\_\_\_\_\_\_\_\_** | 20.0+/-0.1 |
| 22.0 | **\_\_\_\_\_\_\_\_** | 22.0+/-0.1 |
| 24.0 | **\_\_\_\_\_\_\_\_** | >23.5 |
| 26.0 | **\_\_\_\_\_\_\_\_** | >24.0 |

Set the coarse power adjustment to 4dBm, 16dBm, 22dBm and scan the fine adjustment from 0dBm to 2.2dBm. For each scan, measure and record the RF output power with the power meter.

|  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| **4dBm** | | | **16dBm** | | | **26dBm** | | |
| **Setting**  **(dBm)** | **Measurement**  **(dBm)** | **Nominal**  **(dBm)** | **Setting**  **(dBm)** | **Measurement**  **(dBm)** | **Nominal**  **(dBm)** | **Setting**  **(dBm)** | **Measurement**  **(dBm)** | **Nominal**  **(dBm)** |
| +0.0 | **\_\_\_\_\_\_\_\_** | 4.0+/-0.1 | +0.0 | **\_\_\_\_\_\_\_\_** | 16.0+/-0.1 | +0.0 | **\_\_\_\_\_\_\_\_** | >24.5 |
| +0.2 | **\_\_\_\_\_\_\_\_** | 4.2+/-0.1 | +0.2 | **\_\_\_\_\_\_\_\_** | 16.2+/-0.1 | +0.2 | **\_\_\_\_\_\_\_\_** | >24.5 |
| +0.4 | **\_\_\_\_\_\_\_\_** | 4.4+/-0.1 | +0.4 | **\_\_\_\_\_\_\_\_** | 16.4+/-0.1 | +0.4 | **\_\_\_\_\_\_\_\_** | >24.5 |
| +0.6 | **\_\_\_\_\_\_\_\_** | 4.6+/-0.1 | +0.6 | **\_\_\_\_\_\_\_\_** | 16.6+/-0.1 | +0.6 | **\_\_\_\_\_\_\_\_** | >24.5 |
| +0.8 | **\_\_\_\_\_\_\_\_** | 4.8+/-0.1 | +0.8 | **\_\_\_\_\_\_\_\_** | 16.8+/-0.1 | +0.8 | **\_\_\_\_\_\_\_\_** | >24.5 |
| +1.0 | **\_\_\_\_\_\_\_\_** | 5.0+/-0.1 | +1.0 | **\_\_\_\_\_\_\_\_** | 17.0+/-0.1 | +1.0 | **\_\_\_\_\_\_\_\_** | >24.5 |
| +1.2 | **\_\_\_\_\_\_\_\_** | 5.2+/-0.1 | +1.2 | **\_\_\_\_\_\_\_\_** | 17.2+/-0.1 | +1.2 | **\_\_\_\_\_\_\_\_** | >24.5 |
| +1.4 | **\_\_\_\_\_\_\_\_** | 5.4+/-0.1 | +1.4 | **\_\_\_\_\_\_\_\_** | 17.4+/-0.1 | +1.4 | **\_\_\_\_\_\_\_\_** | ~4.0 |
| +1.6 | **\_\_\_\_\_\_\_\_** | 5.6+/-0.1 | +1.6 | **\_\_\_\_\_\_\_\_** | 17.6+/-0.1 | +1.6 | **\_\_\_\_\_\_\_\_** | ~4.0 |
| +1.8 | **\_\_\_\_\_\_\_\_** | 5.8+/-0.1 | +1.8 | **\_\_\_\_\_\_\_\_** | 17.8+/-0.1 | +1.8 | **\_\_\_\_\_\_\_\_** | ~4.0 |
| +2.0 | **\_\_\_\_\_\_\_\_** | 6.0+/-0.1 | +2.0 | **\_\_\_\_\_\_\_\_** | 18.0+/-0.1 | +2.0 | **\_\_\_\_\_\_\_\_** | ~4.0 |
| +2.2 | **\_\_\_\_\_\_\_\_** | 6.2+/-0.1 | +2.2 | **\_\_\_\_\_\_\_\_** | 18.2+/-0.1 | +2.2 | **\_\_\_\_\_\_\_\_** | ~4.0 |

***Stabilization Test***

1. **Check monitor values:**

Keep the RF output connected to the power meter. Set the output level at 13dBm.

Check voltage values of the monitor channels with the voltage meter.

Note: BIAS1 and MON1 are located at the right half of the top monitor BNC connectors. BIAS2 and MON2 are at the left half.

|  |  |  |
| --- | --- | --- |
| **Channel** | **Measurement** | **Nominal** |
| CTRL | **\_\_\_\_\_\_\_\_** | ~3V |
| BIAS1 | **\_\_\_\_\_\_\_\_** | 1.7~1.8V |
| MON1 | **\_\_\_\_\_\_\_\_** | < +/-15mV |
| BIAS2 | **\_\_\_\_\_\_\_\_** | 1.7~1.8V |
| MON2 | **\_\_\_\_\_\_\_\_** | < +/-15mV |

1. **Closed loop transfer function of the internal loop for the 2nd RFAM detector**

Keep 13dBm output of the unit on the power meter. Connect the signal generator of SR785 to the exc port and inject 5Vpk random noise. Set the frequency range upto 25Hz and 800lines. Turn on the exc switch. Measure the transfer function from MON1 to MON2.

Confirm the transfer function is 0 dB +/- 0.5 dB above 10Hz.

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

1. **Power Spectral Density**

Keep 13dBm output of the unit on the power meter. Make sure that the exc switch is off and the signal generator of SR785 is disconnected from the exc port. Measure power spectral density of the following signals at the listed frequencies. Note: PSD Unit (/rtHz) should be turned **on** when reading the numbers.

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | **Measurement [V/rtHz]** | | | |
| **Channel** | **10Hz** | **100Hz** | **1kHz** | **10kHz** |
| CTRL | \_\_\_\_\_\_\_\_ | \_\_\_\_\_\_\_\_ | \_\_\_\_\_\_\_\_ | \_\_\_\_\_\_\_\_ |
| BIAS1 | \_\_\_\_\_\_\_\_ | \_\_\_\_\_\_\_\_ | \_\_\_\_\_\_\_\_ | \_\_\_\_\_\_\_\_ |
| MON1 | \_\_\_\_\_\_\_\_(14A) | \_\_\_\_\_\_\_\_(14B) | \_\_\_\_\_\_\_\_(14C) | \_\_\_\_\_\_\_\_(14D) |
| BIAS2 | \_\_\_\_\_\_\_\_ | \_\_\_\_\_\_\_\_ | \_\_\_\_\_\_\_\_ | \_\_\_\_\_\_\_\_ |
| MON2 | \_\_\_\_\_\_\_\_(14E) | \_\_\_\_\_\_\_\_(14F) | \_\_\_\_\_\_\_\_(14G) | \_\_\_\_\_\_\_\_(14H) |

1. **Calibration 1**

Keep 13dBm output of the unit on the power meter. Connect the signal generator of SR785 to the exc port and inject 5Vpk sinusoidal signal. Turn on the exc switch. Check the signal amplitude of the following ports at 1kHz. Note: PSD Unit (/rtHz) should be turned **off** when reading the numbers.

|  |  |  |
| --- | --- | --- |
| **Channel** | **Measurement [mVrms]** | **Nominal** |
| CTRL | \_\_\_\_\_\_\_\_ | ~3mVrms |
| MON1 | \_\_\_\_\_\_\_\_(15A) | ~350mVrms |
| MON2 | \_\_\_\_\_\_\_\_(15B) | ~350mVrms |

1. **Calibration 2**

Keep the above excitation setting. Turn off the output switch and connect the output to the spectrum analyzer. Turn on the output switch again. Measure the peak height of the carrier power at 9.1MHz and sideband power at 9.099MHz and 9.101MHz.

|  |  |
| --- | --- |
| **Signal** | **Measurement**  **[dBm]** |
| Carrier  (9.100MHz) | \_\_\_\_\_\_\_\_(16A) |
| Upper Sideband  (9.101MHz) | \_\_\_\_\_\_\_\_ (16B) |
| Lower Sideband  (9.099MHz) | \_\_\_\_\_\_\_\_(16C) |

1. **Calibration ~ Calculation**

Copy numbers from above calibration measurements and fill the blanks. Confirm the final numbers if they fulfill the requirements

Signal Amplitudes

ACR = Sqrt[ 10^(\_\_\_\_\_\_\_\_(16A) / 10) /1000\*50\*2] = \_\_\_\_\_\_\_\_ [Vpk]

AUSB = Sqrt[ 10^(\_\_\_\_\_\_\_\_(16B) / 10) /1000\*50\*2] = \_\_\_\_\_\_\_\_ [Vpk]

ALSB = Sqrt[ 10^(\_\_\_\_\_\_\_\_(16C) / 10) /1000\*50\*2] = \_\_\_\_\_\_\_\_ [Vpk]

Modulation Depth

mAM = ( AUSB + ALSB ) / ACR = \_\_\_\_\_\_\_\_ [RIN]

Calibrations of the monitor channels

calMON1 = \_\_\_\_\_\_\_\_(15A) \* Sqrt(2) / mAM = \_\_\_\_\_\_\_\_ [Vpk/RIN]

calMON2 = \_\_\_\_\_\_\_\_(15B) \* Sqrt(2) / mAM = \_\_\_\_\_\_\_\_ [Vpk/RIN]

Calibrated power spectral densities

MON1 (In-loop)

10Hz: \_\_\_\_\_\_(14A) / calMON1 = \_\_\_\_\_\_ [RIN / rtHz] less than 1.0x10-8 /rtHz (160dBc/Hz)

100Hz: \_\_\_\_\_\_(14B) / calMON1 = \_\_\_\_\_\_ [RIN / rtHz] less than 3.2x10-9 /rtHz (170dBc/Hz)

1kHz: \_\_\_\_\_\_(14C) / calMON1 = \_\_\_\_\_\_ [RIN / rtHz] less than 3.2x10-9 /rtHz (170dBc/Hz)

10kHz: \_\_\_\_\_\_(14D) / calMON1 = \_\_\_\_\_\_ [RIN / rtHz] less than 3.2x10-9 /rtHz (170dBc/Hz)

MON2 (Out-of-loop)

10Hz: \_\_\_\_\_\_(14E) / calMON2 = \_\_\_\_\_\_ [RIN / rtHz] less than 1.0x10-8 /rtHz (160dBc/Hz)

100Hz: \_\_\_\_\_\_(14F) / calMON2 = \_\_\_\_\_\_ [RIN / rtHz] less than 3.2x10-9 /rtHz (170dBc/Hz)

1kHz: \_\_\_\_\_\_(14G) / calMON2 = \_\_\_\_\_\_ [RIN / rtHz] less than 3.2x10-9 /rtHz (170dBc/Hz)

10kHz: \_\_\_\_\_\_(14H) / calMON2 = \_\_\_\_\_\_ [RIN / rtHz] less than 3.2x10-9 /rtHz (170dBc/Hz)

***DAQ channel Test***

1. **Check functionality of the DAQ port**

Measure the transfer functions between the following ports with SR785 with no excitation injected. The two pins of the DAQ port signals should be measured differentially: Use “A-B” input mode of the SR785. Connect two signal clips to the SR785 inputs and clip the signals with the positive clips.

|  |  |  |  |
| --- | --- | --- | --- |
| **Transfer Function** | | **Measurement** | **Nominal Gains** |
| **From (BNC Port)** | **To (DAQ DSUB Port)** |
| MON1 | Pin1(+) and Pin6 (-) | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 20? |
| CTRL | Pin2(+) and Pin7 (-) | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 2 (f<0.1Hz), 40 (f>100Hz) |
| MON2 | Pin3(+) and Pin8 (-) | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 20? |
| BIAS2 | Pin4(+) and Pin9 (-) | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 2 |

***Test with Remote Control RIG***

1. **Check indicators of EtherCAT EOM Driver Test Rig**

Continue to use the 13dBm output setting.

Turn off the exc and output switches. Make sure the Int/Ext switch on the main unit is at “Int”. Connect the test rig to the control connector with a DB25 M/M cable.

Make sure the Latch switch on the test rig is at “Dis”.

Toggle the Int/Ext switch on the main unit. Confirm the indicator “Int” on the test rig

correctly shows the state. **Confirm \_\_\_\_\_\_\_\_\_\_\_**

Turn the Int/Ext switch on the main unit to “Ext”.

Toggle the RF output switch on the main unit. Confirm the indicator “RF” on the test rig

correctly shows the state. **Confirm \_\_\_\_\_\_\_\_\_\_\_**

Make sure the exc switch is off. Toggle the Exc Enable on and off. Confirm the indicator

“Exc Readback” on the test rig correctly shows the state. **Confirm \_\_\_\_\_\_\_\_\_\_\_**

1. **Power setting test**

Turn off the RF output switch on the main unit. Connect the power meter to the RF output port. Then turn on the RF output switch. Turn off (up) all the value switches.

Measure the RF output power at this point and confirm it is still about 13dBm

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

Turn the Latch switch on the test rig to the “En (Enable)” position.

Make sure the RF output power is now about 4dBm

**Confirm \_\_\_\_\_\_\_\_\_\_\_**

Measure the RF output while turning on the switches one by one. (i.e. Only one switch is on at a time). Note: The panel marking (+1dB, +2dB, …) is obsolete and they should be scaled by a factor of 0.2. (i.e. +1dB = +0.2dB, +2dB = +0.4dB, and so on)

|  |  |  |
| --- | --- | --- |
| Switch | Measurement | Nominal [dBm] |
| +1dB on | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 4.2+/-0.1 |
| +2dB on | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 4.4+/-0.1 |
| +4dB on | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 4.8+/-0.1 |
| +8dB on | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 5.6+/-0.1 |
| +16dB on | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 7.2+/-0.1 |
| +32dB on | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 10.4+/-0.1 |
| +64dB on | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 16.8+/-0.1 |
| +128dB on | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | ~4.0 |

1. **Test on the signals via the test rig**

Turn on the +1, +4, +8, and +32dB switches on the test rig. This should yield the output of 13dBm. Measure the transfer functions between the following ports with SR785.

|  |  |  |  |
| --- | --- | --- | --- |
| **Transfer Function** | | **Measurement** | **Nominal Gains** |
| **From (BNC Port)** | **To (TEST RIG Port)** |
| BIAS1 | BIAS1 | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 1 |
| BIAS2 | BIAS2 | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | 1 |
| CTRL | CTRL | \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | Bandpass shape  20dB@15Hz 0deg |