## Slide 1: Tilt-correction and Sensor-correction at LHO Using offline data for designing and quantifying filters
Conor Mow-Lowry
Thanks to: Krishna, Jeff, Brian, Jim, Hugh, Robert

## Slide 2: Tilt versus Horizontal displacement
Conventional seismometers and tiltmeters cannot differentiate between horizontal displacement and ground tilt.
x  ax
x  g
Tilt response to horizontal displacement response for all seismometers = -
Tilt is confused with horizontal motion at low frequencies (below ~ 0.1 Hz).
Solution: Inertial rotation sensors, Tilt-free seismometers
Source: Krishna Venkateswara

## Slide 3
Schematic
Source: Krishna Venkateswara
Autocollimator
10-25 m-thick Cu-Be Flexures

## Slide 4: The ISI tilt problem, Stage 2
Calibration step | Effect on displacement | Relevant frequency
GS-13 measures velocity | 1/f | w = 1
Amplifier noise | 1/√f | fc = 1 Hz
Plant inversion | 1/f2 | f0 = 1 Hz
Tilt-baseline | 1 | --
Tilt-to-horizontal coupling | 1/f2 | g/w2 = 1 @0.5 Hz
Result | 1/f5.5 | Below 0.5 Hz
GS-13 tilt-noise and its effect on horizontal translation

## Slide 5: Some points
RMS velocity is a good figure of merit, and probably what we care about below ~0.3 Hz.
Sensor correction allows us to use a lower-tilt seismometer for low-frequency isolation.
Below 100mHz, ground motion is pretty common mode. The primary microseism ~70mHz sees at least a factor 4 direct subtraction.
We should design further ISI improvements in the IFO basis, and potentially using IFO readout.

## Slide 6: BRS Signal Path
Offline data was used to tune all the BRS filters for improved tilt-subtraction and reduced low-frequency noise injection.
Most benefit came from matching the high-pass filter to the STS2 AC-coupling.

## Slide 7: BRS Tilt-correction filter tuning (ETMX)
m/s
nm/s

## Slide 8: BRS Tilt-correction filter tuning (ETMY)
m/s
nm/s

## Slide 9: Figure of merit
We want to minimise the low-frequency RMS velocity transmitted through the Sensor-Correction filter.
This motion is typically dominated by (uncorrelated) tilt.
We want to see whether the BRS causes harm when there is no wind.
And measure the reduction of injected arm-length motion during high winds.

## Slide 10: Sensor Correction filter

## Slide 11: BRS impact, no wind (<4mph)

## Slide 12: BRS impact, X-arm, high wind (~25 mph)

## Slide 13: BRS impact, Y-arm, high wind (~25 mph)

## Slide 14: Conclusions
The BRS does no harm, even with no wind.
At high winds, it provides ~a factor of 5 reduction in arm-length RMS velocity.
The low-tilt corner station STS2 is nearly as good as tilt-corrected ETMY in pretty high winds.
Identical Sensor correction filters in all corner-station horizontal-translation Stage-1 DoFs should result in substantially less differential motion.
Sensor correction and blending should be tuned using real data with high microseismic motion.

## Slide 15: Does Sensor Correction help? (yes)
Tilt-corrected ground After sensor correction filter

## Slide 16: IFO-basis seismic motion is useful!

## Slide 17: IFO-basis seismic motion is useful!

## Slide 18: No-wind data
Tilt-corrected ground After sensor correction filter

## Slide 19: High-wind data
Tilt-corrected ground After sensor correction filter