LASER INTERFEROMETER GRAVITATIONAL WAVE OBSERVATORY

LIGO-India

Technical Note Oct, 2019

Vacuum System Requirement Document (LIGO-E1700313-v7)

Prepared by: S Sunil, R Kumar, A Prajapati, A Dasgupta, A Srivastava, H K Gulati, S Mukherjee (IPR) B C Pant, S Raja (RRCAT) T Souradeep (IUCAA) M Goverdhan (DCSEM) On behalf: LIGO-India Team

| Directorate of Construction, Services & Estate Management Anushakti Nagar, Mumbai, India 400094 E-mail: ligodcsem@gov.in | Inter University Centre for Astronomy and Astrophysics Ganeshkhind, SavitribaiPhule Pune University Campus Pune 411 07, India E-mail: tarun@iucaa.in |
| --- | --- |
| Institute for Plasma Research LIGO Division Bhat, Gandhinagar Gujarat, India 382428 E-mail: ligo.division@ipr.res.in | Raja Ramanna Centre for Advanced Technology LIGO Division P.O.:CAT, Indore – 452013, M.P., India E-mail: sendhil@rrcat.gov.in |

Content

1.	Laser Interferometer Gravitational wave Observatory	5

1.1	Introduction	5

1.2	Detector configuration	7

2.	LIGO-India Detector	8

2.1	Vacuum System of the interferometer	8

2.1.1.	Basic symmetric chamber (BSC) and Horizontal Access Module (HAM):	11

a.	Purpose of BSC and HAM	12

b.	Dimension of BSC	14

c.	Tolerance for BSC chamber	16

d.	Dimension of HAM	16

e.	Tolerances for HAM chamber	18

f.	Interfaces	18

2.1.2.	Mode-cleaner (MC) tube	18

a.	Purpose	18

b.	Dimension of MC tube	19

c.	Tolerances for MC tube	20

d.	Interfaces	20

2.1.3.	Spool & Adaptor	20

a.	Purpose:	20

b.	Dimension of Spool and Adaptors:	21

c.	Tolerances for Spools and Adaptors	22

d.	Interfaces	22

2.1.4.	Beam tube	22

a.	Purpose:	22

b.	Dimension:	25

c.	Tolerances for beam tube:	26

d.	Interfaces	27

2.1.5.	Metal Baffles (beam tube):	27

a.	Purpose	27

b.	Dimension	28

c.	Tolerances for Baffles	29

d.	Interfaces	30

2.1.6.	Gate valves	30

a.	Purpose	30

2.1.7.	80 K cryopump	31

a.	Purpose	31

b.	Dimension:	34

c.	Tolerances for cryopump	35

d.	Interfaces	35

2.2.	Vacuum pumps for LIGO-India interferometer	35

a.	LVEA Pumping and pump location	35

b.	LVEA Interfaces with vacuum pumps	37

c.	LVEA Power requirement for vacuum pumps	37

d.	VEA Pumping and pump location	37

e.	VEA Interfaces with vacuum pump	38

f.	VEA Power requirement for vacuum pumps	38

g.	Beam tube Pumping system	38

h.	Beam tube interfaces and location of vacuum pump	41

i.	Beam tube Power requirement for vacuum pumps	41

j.	Evacuation of assembled vacuum system	41

2.3.	Vacuum system basic design and fabrication requirement: Other than the beam tube.	43

a.	Codes & Standards	43

b.	Materials & Design loads	44

c.	Fabrication & Technical requirements	45

d.	Vent & Purge systems	49

e.	Reliability and Quality Assurance	50

f.	UHV Cleaning	50

g.	QA / QC plans	52

2.4.	Vacuum system basic design and fabrication requirement: Beam tube.	54

a.	Codes & Standards	54

b.	Materials & Design loads	55

c.	Fabrication & Technical requirements	56

d.	Vent & Purge systems	58

e.	Reliability and Quality Assurance	58

f.	UHV Cleaning	59

g.	QA / QC plans	60

2.5.	Vacuum Control and Monitoring System (VCMS)	62

2.6.	aLIGO-India detector upgrade	64

2.7.	ANNEXURE	65

LIGO-India Vacuum System Requirement Document (LI-VSRD)

The LI-VSRD intends to explain the key requirements of vacuum system for the proper functioning of a Laser Interferometer Gravitational wave Observatory (LIGO). The scope of this document is limited to individual vacuum system with its critical dimensions, materials used for fabrication, execution plan for setting up the LI-interferometer. The document is written based on our understanding and references collected of similar detector installed at Hanford and Livingston. This is our first basic level document based on which other technical documents like Conceptual Design, Preliminary Design and Final Design Reports of the detector will follow.

This document follows Adv. LIGO configuration. Interface of vacuum along with interferometer components, civil infrastructure and data acquisition and control would be covered in separate document. Also detailed QA/QC and timeline along with major milestone would cover in separate document.

In many places the dimensions are mentioned in inches as well as in mm/cm/m. The standard conversion of 1 inch = 25.4 mm is used. The mm/cm/m dimension is to be used for design/analysis/fabrication purposes.

- Laser Interferometer Gravitational wave Observatory

## 1.1	Introduction

Laser interferometer gravitational wave observatory (LIGO) is a detector to understand origin of the gravitational waves (LIGO-G1702087-v1). Gravitational waves are radiation which travel at the speed of light and are transparent to any matter and set into motion any matter they interact with. LIGO is a modified version of the Michelson interferometer which incorporates a pair of suspended Fabry-Perot cavity. The orthogonal arms of LIGO is 4-km each and is installed for studying astrophysical events which radiate energy in the form of gravitational waves. Some of the astrophysical events which emits strong gravitational wave radiations that can be detected by earth-based detectors are coalescing of binary black holes (B.P. Abbott et. al., PRL, 116, 241103), neutron stars (B. P. Abbott et. al. PRL, 119, 161101), exploding supernovas (arXiv:1709.00955v2) and spinning neutron stars (K. Riles, arXiv:1209.0667v3). Although the gravitational waves emitted at the source are strong, by the time these radiations travelling at the speed of light reaches earth, their amplitude decays by r-1 (where r is the distance of earth from the source that emits gravitational waves) and cannot be measured using normal measuring techniques.

The detector achieves its capability by using sophisticated engineering techniques which attenuates noise in high frequency range and tunes the detector to lower frequency region. The schematic of LIGO is shown in fig. 1.1. All the optical components shown in fig. 1.1 are suspended using multiple stage vibration isolation (F. Maticrad et. al., CQG 32, 185003) such that the ambient noises are cancelled. The measurement being extremely small, all the optical components are housed and the experiment is carried out in ultra high vacuum.

Fig.1.1: Experimental setup of LIGO (Courtesy: DCC-G1300494-V3)

Current version of the detector, termed as the Advanced LIGO (Adv-LIGO), is tailored such that the detector can cover the volume in space by a factor of 1000 compared to its earlier versions of detector termed as Initial LIGO and the Enhanced LIGO. The strain sensitivity (B.P. Abbott et. al., PRL, 116, 241103) of Adv-LIGO is measured to be h < 3.0  10–24 Hz–1/2 at 150 Hz as shown in fig. 1.2

Fig. 1.2: Strain sensitivity of Advanced LIGO compared at two observatories (Hanford-H1 and Livingston-L1) at USA.

## 1.2	Detector configuration

LIGO-US detector mainly comprises of the following three sections:

- Corner station

- Mid-station

- End-stations

All these three sections are interconnected with stainless steel beam tube.

Schematic of a LIGO-US detector is shown in fig. 1.3.

Fig. 1.3: Schematic of LIGO-US interferometer.

- LIGO-India Detector

At the time of writing this document, it is envisaged that, LIGO-India detector will have the same design as the Adv. LIGO-US detector. The challenge LIGO-India face is of making a detector at par with the LIGO-US detector in terms of vacuum requirement. The required vacuum inside the interferometer is better than 10-9 mbar with average partial pressure requirement for hydrogen gas to be less than 10-9 mbar (Dennis Coyne, LIGO-T040001-00-D) to have the detector strain sensitivity of 10–24 Hz−1/2.

  - Vacuum System of the interferometer

LIGO detector is an integrated system of many vacuum systems and components. Each of the stations described in fig. 1.4 comprises of following vacuum components as listed in table 2.1, 2.2 and 2.3. LIGO-India will follow the coordinate systems for the installation of vacuum chambers as per the LIGO-US detector at Livingston (D0901859-v3).

Fig. 1.4: Major vacuum components with interfaces of the interferometer along X-arm. The interfaces are same for Y-arm.

Vacuum components housed in corner station (D0901859-v2, V049-5-003) are described in table 2.1:

Table 2.1: Details of the vacuum components in Corner Station.

| Sr. No. | Vacuum Components | Quantity | Distribution |
| --- | --- | --- | --- |
| 1 | HAM Chamber | 06 | 03 (X-Arm) |
|  |  |  | 03 (Y-Arm) |
| 2 | BSC Chamber | 03 | 01 (Centre) |
|  |  |  | 01 (X-Arm) |
|  |  |  | 01 (Y-Arm) |
| 3 | LN2Cryo pump (Long) | 02 | 01 (X-Arm) |
|  |  |  | 01 (Y-Arm) |
| 4 | Large gate valves (GV) | 06 | 02 (Isolation of Cryo pump, X-Arm) 02 (Isolation of Cryo pump, Y-Arm) 02 (Vacuum components isolation) |
| 5 | Spool & Adaptors | - | For connecting of vacuum components in the Corner Station. |
| 6 | MC Tubes | 04 | For connecting HAM-2 & HAM-3, HAM-4 & HAM-5 (i.e. 2 2) |
| 7 | Bellows | - | In Spool and Adaptors |

Vacuum components housed in mid-station (D1600164) are described in table 2.2:

Table 2.2: Details of the vacuum components in Mid-Station.

| Sr. No. | Vacuum Components | Quantity | Distribution |
| --- | --- | --- | --- |
| 1 | Large gate valves (GV) | 02 | 01 (X-Arm) |
|  |  |  | 01 (Y-Arm) |

Vacuum components housed in end-station (V049-5-008, V049-5-009) are described in table 2.3:

Table 2.3: Details of the vacuum components in End-Station.

| Sr. No. | Vacuum Components | Quantity | Remark |
| --- | --- | --- | --- |
| 1 | BSC Chamber | 02 | 01 (X-Arm) |
|  |  |  | 01 (Y-Arm) |
| 2 | LN2Cryo pump (Short) | 02 | 01 (X-Arm) |
|  |  |  | 01 (Y-Arm) |
| 3 | Large Gate Valves (GV) | 04 | 02 (Isolation of each Cryo pump) |
| 4 | Spool & Adaptors | - | For connecting of vacuum components in the End station. |
| 5 | Bellows | - | In Spool and Adaptors |

In order to reduce different types of noises in the interferometer like (a) phase noise from forward scattering (b) thermal noise from mechanical damping (c) acoustic noise coupling (d) seismic noise coupling and (e) mirror contamination and movement of test mass due to air current, the experimental components has to be housed in ultra-high vacuum. Critical vacuum requirement for the interferometer is shown in table 2.4.

Table 2.4: Basic vacuum parameters for the detector (Reference T040001)

| Sr. No. | Parameters | Adv. LIGO Requirement | Goal |
| --- | --- | --- | --- |
| 1 | Vacuum | <1.0 10–9 mbar | < 1.0 10–10 mbar |
| 2 | Partial pressure for hydrogen gas | 1.0 10–9 mbar | 1.0 10–10 mbar |
| 3 | Partial pressure for water vapour | 1.0 10–10 mbar | 1.0 10–10 mbar |
| 4 | Partial pressure for nitrogen | 6.0 10–11 mbar | 6.0 10–11 mbar |
| 5 | Partial pressure for carbon monoxide | 5.0 10–11 mbar | 5.0 10–11 mbar |
| 6 | Partial pressure for carbon dioxide | 2.0 10–11 mbar | 2.0 10–11 mbar |
| 7 | Partial pressure for methane | 3.0 10–11 mbar | 3.0 10–11 mbar |
| 8 | Partial pressure for AMU 100 hydrocarbon | 2.0 10–12 mbar | 7.0 10–13 mbar |
| 9 | Partial pressure for AMU 300 hydrocarbon | 2.2 10–13 mbar | 5.0 10–14 mbar |
| 10 | Partial pressure for AMU 500 hydrocarbon | 9.0 10–14 mbar | 1.0 10–14 mbar |

The vacuum chambers for LIGO-India have to be compatible for all in-vacuum optical and mechanical structures. This include vibration isolation stages, end mirrors, beam splitter, power and signal recycling mirrors, thermal compensation plates, mode-cleaner optics, mode matching optics baffles, laser beam dump etc.. In this section we deal with the following vacuum systems and their vacuum requirements:

- Basic Symmetric Chamber

- Horizontal Access Module

- Mode-cleaner tubes

- Spools and adaptors

- Beam tube

- Baffles

- Gate valves

- 80K Cryopump

    - Basic symmetric chamber (BSC) and Horizontal Access Module (HAM):

  - Purpose of BSC and HAM

Basic Symmetric Chambers (BSC) are used to support core optical components while Horizontal Access Module (HAM) are used to house auxiliary optics(Ref. F. Matichard et. al. CQG, 32, 185003, 2015). All major optical components are housed in these two chambers. Both BSC and HAM chambers are provided with demountable cover flanges so that the servicing of the optical components could be carried out whenever needed. These chambers contain seismic isolators and alignment mechanism which support the optical elements and have internal attachment brackets. These brackets will be used to support lightweight optical components. The seals on these cover flanges shall be designed as double O-rings with a pumped annulus to reduce the gas load due to the permeability of Viton. A clean air vent and purge system shall be provided to break vacuum and maintain cleanliness of the optical components whenever a chamber is open.

3D representation of BSC and HAM in Adv. LIGO is shown in fig. 2.1 and 2.2 respectively.

Fig. 2.1: (a) 3-D view of a BSC chamber

Fig 2.1. (b) Actual picture of assembled BSC without HEPI

Fig. 2.2: (a) 3-D view of a HAM chamber,

Fig 2.2. (b) Actual picture of assembled HAM without HEPI

  - Dimension of BSC

BSC chamber (annexure A) is made in two parts consisting of the following components:

- Upper Part

  - Dished and Flanged head

  - Cylindrical shell

  - Lifting lug

  - Ports

  - Stiffener

  - Connecting flanges

- Lower Part

  - Connecting flange

  - Dished and Flanged head

  - Cylindrical shell

  - End cover

    - End cover type 1(Contain five port)

    - End cover type 2 (Contain one port)

  - Ports

  - Stiffener

  - Internal attachment brackets

- Support structure

The bottom section of the shell has many port openings for laser beam access, electrical feed-through and diagnostics etc. The few features of this single chamber are as follows

- Volume of each chamber 	               : ~ 27 m3

- Surface area exposed to vacuum    	   : ~ 52 m2

- Wall thickness (upper part)		   : 6.35 mm

- Wall thickness (lower part)		   : 12.7 mm

The overall dimensions of BSC chamber are shown in Fig.2.3 and also listed in table 2.5.

Isometric view of BSC chamber showing the ports and interfaces

Fig.2.3: Overall dimensions of BSC chamber.

Table 2.5: Dimensions of BSC Chamber.

| Sr. No. | Parameters | Values [mm (inch)] |
| --- | --- | --- |
| 1 | Inner Diameter | 2654.3 (104.5) |
| 2 | Upper Part |  |
|  | Height | 1994 (78.5) |
|  | No. of Nozzles / OD Size | 02 / 356 (14) |
| 3 | Lower Part |  |
|  | Height | 2823 (111.13) |
|  | Connecting flange ID | 2654.3 (104.5) |
|  | End cover ID | 1536.7 (60.5) |
| 4 | No. of Nozzles / OD Size | 12 / 356 (14) 06 / 254 (10) 12 / 203 (8) 01 / 64 (2.5) |

  - Tolerance for BSC chamber

The tolerance of BSC will be as per the reference document mentioned in the annexure A.

  - Dimension of HAM

HAM chamber (annexure A) consists of the following components:

- Cylindrical Part

  - Cylindrical shell

  - Access cover

  - Ports

  - Lifting lug

  - Internal attachment brackets

  - Connecting flanges

- Two big Cylindrical Ports

  - Cylindrical shell

  - End cover

  - Ports

  - Bellow with tie rod

  - Connecting flanges

- Support Structure

This chamber has > 70% opening for accessing optical components and many other port openings for laser beam access, electrical feed-throughs and for diagnostics setup etc. The few features of this single chamber are as follows

- Volume of each chamber 	               : ~ 10 m3

- Surface area exposed to vacuum    	   : ~ 27 m2

- Wall thickness				   : 12.7 mm

The overall dimensions of HAM chamber are shown in Fig.2.4 and also listed in table 2.6.

Top View

Isometric view of HAM chamber

Fig.2.4: Overall dimensions of HAM chamber.

Table 2.6: Dimensions of HAM chamber.

| Sr. No. | Parameters | Values [mm (inch)] |
| --- | --- | --- |
| 1 | Cylindrical Part |  |
|  | Inner Diameter | 2140 (84.25) |
|  | Length | 1854 (73) |
|  | No. of Nozzles / OD Size | 01 / 356 (14) 04 / 305 (12) 06 / 254 (10) 10 / 203 (8) 01 / 64 (2.5) |
| 2 | Two big Cylindrical Ports |  |
|  | Inner Diameter | 1537 (60.5) |
|  | No. of Nozzles / OD Size | 02/254 (10) |

  - Tolerances for HAM chamber

The tolerance of HAM will be as per the reference document mentioned in the Annexure A.

  - Interfaces

HAM and BSC have interfaces with long cryo-pump, MC tubes, spools & adaptors and large gate valves as shown in fig. 1.4. HAM1 and HAM2 are connected by a septum plate (ref Annexure A), HAM2 and HAM3 by mode cleaner tubes (section 2.1.2), while HAM3 and BSC2, BSC2 and BSC3 are connected with spools and BSC3 to gate valve GV2 with adaptor (section 2.1.3).

    - Mode-cleaner (MC) tube

- Purpose

The mode cleaner tubes are provided to reduce the high frequency amplitude noise of the laser light. These MC tubes act as an intermediate frequency reference for the laser and reduces the angular fluctuations of the laser beam.

Mode cleaner tubes consist of rolled sections of varying length. Supports for the cleaner tubes will be similar to that of Adv. LIGO and as shown in fig. 2.5. Bellows are also provided which allow thermal expansion during baking. The connecting flanges have dual Viton O-rings design with an inter-space pumping facility. These tubes are used for connecting of HAM-2 & HAM-3 and HAM-4 & HAM-5 chambers.

Fig. 2.5: 3-D view of Mode Cleaner tube and its support structure

- Dimension of MC tube

MC tube (annexure A) consists of following components:

- Cylindrical Shell

- Big Cylindrical shell

- Small Cylindrical shell

- Bellows with tie rods

- Flanges

- Stiffeners

- Ports

- Lifting Lugs

The few features of these MC Tube are as follows

- Volume of all MC Tube  	               : ~ 100 m3

- Surface area exposed to vacuum    	   : ~ 180 m2

- Wall thickness				   : 6.35 mm

The dimensional details for MC tube are mentioned in table 2.7.

Table 2.7:Dimensions of MC tube.

| Sr. No. | Parameters | Values [mm (inch)] |
| --- | --- | --- |
| 1 | Big Cylindrical Shell ID | 2146 (84.5) |
| 2 | Small Cylindrical Shell ID | 1553 (61.13) |
| 3 | Length | 6858 (270) |
| 3 | No. of Nozzles / OD Size | 04 / 356 (14) 18 / 254 (10) 48 / 203 (8) 02/ 152 (6) 02/ 64 (2.5) |

- Tolerances for MC tube

The tolerance of MC Tube will be as per the reference document mentioned in annexure A.

- Interfaces

Mode cleaner tube connect HAM2 to HAM3 and HAM4 to HAM5. The ends of the mode cleaner tubes have adaptive flanges to connect to the HAM chambers.

    - Spool & Adaptor

  - Purpose:

The Spools and adaptors are mainly used for two purposes

- For inter-connecting of vacuum equipments

- Optical requirement

The size and quantities depend upon the availability of space and the functional requirement of the interferometer components. A typical spool and adaptor is shown in fig. 2.6.

        - (b)

Fig. 2.6: 3-D view of (a) Spool and (b) Adaptor

  - Dimension of Spool and Adaptors:

As per current LIGO US (LLO) configuration details of spool and adaptor are:

Spool and Adaptor (annexure A) consist of the following components:

- Cylindrical Shell

- Bellows with tie rods

- Flanges

- Stiffeners

- Ports

- Lifting Lugs

Cylindrical shell of the Spool can have different internal diameter as per the adaptors, connecting equipment and its port. The few features of these connectors are as follows

- Total volume of spools & adaptors at LVEA 			: ~ 175 m3

- Total volume of spools & adaptors at each VEA			: ~ 10 m3

- Surface area of spools & adaptors exposed to vacuum at LVEA	: ~ 400 m2

- Surface area exposed to vacuum (VEA)  				: ~ 25 m2

- Wall thickness								: 6.35 & 9.5 mm

The dimensional details for Spool and Adaptor are mentioned in table 2.8.

Table 2.8: Dimensions of Spool and Adaptor.

| Sr. No. | Parameters | Values [mm (inch)] |
| --- | --- | --- |
| 1 | Cylindrical Shell Inner Diameters | 1835 (72.25) 1537 (60.5) 1226 (48.25) 1133 (44.62) |
| 2 | No. of Nozzles / Sizes | 02 / 356 (14) 04 / 254(10) 36 / 203 (8) 02 / 152 (6) 04 / 64 (2.5) 02/ 38 (1.5) |

  - Tolerances for Spools and Adaptors

The tolerance of spool & adaptor will be as per the reference document mentioned in annexure A.

  - Interfaces

Spools and adaptors inter-connect chambers, cryo-pumps and gate valve. They are also used for optical alignment purpose, installation of vacuum equipment and also help in expansion during initial baking. Location of spools and adaptors are shown schematically in fig. 1.4.

    - Beam tube

  - Purpose:

Beam tube (annexure A) is one of the important part of the entire vacuum system. It is a cylindrical structure having 4 km length in each perpendicular direction as shown in fig. 2.7.

- (b)

Fig. 2.7: (a) Orthogonal beam tubes at Hanford. (b) Section of beam tube at 			         Hanford observatory

Hot rolled annealed and pickled (HRAP) SS304L sheet will be used for the fabrication of 8 km beam tube of the interferometer. As hydrogen outgassing at UHV is a concern during the operation of the detector, it is planned to control the ppm level of hydrogen in the SS304L sheet to ~1.0 ppm (LIGO-P930020) at the stage of sheet manufacturing. The steel roll again will go through the process of air baking at 450oC for 36 hours to ensure reduction of hydrogen to a level of ~1 ppm. Before the sheet is spiral welded into beam tube sections, a coupon from each steel roll (in the form of coil) will undergo outgassing measurement for hydrogen to certify an outgassing rate of <10-14 mbar l/s/cm2. Once the steel coupon passes this outgassing rate, the steel roll will be spiral welded into beam tube sections.

The beam tube is kept in ultra-high vacuum so that scattering of laser beam is minimized within the Fabry-Perot cavity.  Each beam tube has two (02) numbers of beam tube modules of 2.0 km length each. Each beam tube module will be fabricated from 20 m long beam tube sections of 100 numbers or 10 m long beam tube sections of 200 numbers. These beam tube sections will be reinforced with vacuum stiffeners as well as support stiffeners. Vacuum stiffeners are provided at a separation of 750 mm distance throughout the length so that the tube can withstand pressure difference between inside and outside the tube and avoid collapsing. Support stiffeners are provided at 20 m apart to stiffen and assemble the support structure for the beam tube.

Three types of support structures are used. They are (a) fixed supports, (b) guided supports and (c) termination supports are also attached with the support stiffeners as shown in fig. 2.8.

- (b)

(c)

Fig. 2.8: Different types of support structures that will be used to anchor the beam tubes. (a) Fixed support, (b) Guided support and (c) Termination support

These support structures are anchored on the ground with a provision for horizontal and vertical alignment. These support structures provide the following features:

- Support the tube and other applicable loads if any.

- Allow thermal expansion/contraction of the tube during baking and diurnal/seasonal temperature changes

- A convenient means of adjusting the alignment of the tubes at installation and after subsequent settling or shifting of the slab.

- Adequate thermal insulation at points of support to limit calculated tube wall temperature differential to less than 10 C with a bulk tube wall temperature of 150C (assuming heating by I2R losses in tube wall).

Fixed support will be fixed to the ground and connected at one end of beam tube so that it cannot slide during expansion and contraction of beam tube. Guided support is a movable type and connected at another end of beam tube so that it can slide during expansion and contraction of beam tube. Each guided support is provided along expansion joint to compensate thermal expansion/contraction of the tube during baking and seasonal temperature variations. It also aids alignment during installation and alignment adjustments during operation. Termination support is a fixed type and connected at the ends of beam tube module (2.0 km apart) to take care of complete axial load and all other loads of the beam tubes. The expansion joint (bellow) of 635 mm (E950026-06-B) will be provided in each 40 m beam tube section so that it compensates for thermal expansion/contraction of the tube during bake out (150 C), and seasonal temperature variations. It aids alignment during installation and alignment adjustments during operation. Each 2 km beam tube module will have 09 ports of OD 254mm. These ports will be used for initial pumping during baking and during normal pumping, connecting pressure gauges and for future use.

Two segment of beam tube section will be joined by butt welding. Out of two segments, one segment of beam tube section will have welded bellow at one end and other tube section will be without bellow.

  - Dimension:

Beam tube mainly consists of following components:

- Cylindrical shell

- Stiffeners

- Vacuum Stiffener

- Support Stiffener

- Support Structure

- Fixed Support

- Guided Support

- Termination Supports

- Expansion Joint

- Baffle

- Pumping Ports

The few features of Beam tube are as follows

- Total volume of Beam tube (each arm)	: ~ 9522 m3

- Total surface area of Beam tube (each arm)	: ~ 30717 m2

- Wall thickness	of Beam tube			: ~ 3 mm

- Total surface area of baffles in each arm	: ~ 225 m2

- Thickness of each baffle			: ~ 1 mm

The dimensional details for Beam tube are mentioned in table 2.9.

Table 2.9: Dimensions of Beam tube.

| Sr. No. |  | Parameters | Values [mm (inch)] |
| --- | --- | --- | --- |
| 1 |  | Inner diameter | 1240 |
| 2 |  | Length in each arm | 3942.5 |
| 3 |  | No. of Nozzles / OD Size | 36 / 254 (10) |
| 4 |  | No of baffles in each arm | 213 |
| 5 |  | ID / OD of each baffle | 1070 / 1240 |

  - Tolerances for beam tube:

The beam tube modules will be fabricated with tolerances mentioned in table 2.10. The technical requirements of expansion joint for beam tube are mentioned in table 2.11.

Table 2.10: Details of Tolerances Requirements for beam tube module.

| Sr. No. | Parameters | Values [mm (inch)] |
| --- | --- | --- |
| 1 | Prependicularity of the end of the tube to the axis of the tube | 0.25 (0.01") |
| 2 | Flatness of the tube end | 0.25 (0.01") |
| 3 | Circumference of the tube ends only | 1.3 ( 3/64") |
| 4 | Longitudinal straightness of tube over 20m | 3.2 ( 1/8") |
| 5 | Cut length over 20 m | 6.3 (+ 0.25") |
| 6 | Equivalent diameter as measured by circumference other than tube ends | + 7.1 (0.28"), - 0 |
| 7 | Concentricity of expanded end to tube axis (if ends are expanded) | 0.25 (0.01") |

Table 2.11: Details of Expansion Joints for beam tube.

| Sr. No. | Parameters | Values [mm (inch)] |
| --- | --- | --- |
| 1 | Expansion Joint type | Single, unreinforced |
| 2 | Length | 635 (25") |
| 3 | Maximum installation movement Axial: Lateral: Rotational: | 3.175 ( 0.125") 0 0.13 degree |
| 4 | Maximum operating movement Axial: Lateral: Rotational: | 6.65 (0.262") Contraction 40.89 (1.610") Extension 1.01 (0.04") - degree |

  - Interfaces

Beam tube interfaces with GV4, GV6 and GV7 in X-Arm and GV8, GV9 and GV11 in Y-Arm as shown in fig 1.4.There are 18 ports on each beam tube arm for connection of vacuum pump, gauge etc.. Blank flanges with manually operated gate valve would be provided when not in use.

    - Metal Baffles (beam tube):

  - Purpose

The Beam Tube Baffles (LIGO-E950083) are needed to reduce the amount of phase-modulated scattered light reaching the photodetectors in the length sensing subsystems of the interferometers. They prevent the guided reflection of light from one end of the beam tube to the other, and they mask the beam tube walls from direct viewing by the LIGO suspended test mass mirrors. A schematic of baffles laid inside the beam tube is shown in fig. 2.9. The baffles must have two key properties:

- they must have acceptably low backscatter;

- they must minimize the amount of forward diffraction which can reach the far test mass mirror.

- The number of baffles and their arrangement in the arm will be as per the document LIGO-E950083 (LLO configuration).

Fig. 2.9: Serrated baffles laid inside the 4-km beam tube

Baffle is in truncated conical shaped one whose bigger end is connected to the inner surface of beam tube and smaller end provides the required clear aperture of 1070 mm. It is used to reduce scattering of laser beam. A number of pumping ports at a right angle to the beam tube axis are provided to remove the outgassing products, primarily water and hydrocarbons during baking.

Optical properties of the baffles have to be taken care while placing inside the tube.

Material of the baffle used inside the tube shall not produce appreciable gas load so as to compromise the vacuum requirement of the interferometer. Table 2.12 shows the vacuum properties of baffles used.

Table 2.12 Vacuum properties of the beam tube baffles

| Species | Requirement |
| --- | --- |
| H2 | <2x10-12 mbar-liter/cm2/s |
| Hydrocarbons | <2x10-16 mbar-liter/cm2/s |

  - Dimension

The full serration baffles shall consist of nominally 0.036” (1 mm) thick SS 304L sheet metal conformed to a truncated conical surface as shown in fig. 2.10. The larger base of this cone is welded to a cylindrical band of similar material which will have the same diameter as the beam tube internal diameter; the smaller base of the cone shall protrude into the beam tube and define the clear aperture. This smaller base shall be serrated along the circumference. The serrations are designed to minimize diffraction.

Fig. 2.10: Section of serrated baffles inside the beam tube

  - Tolerances for Baffles

- Individual teeth shall have fixed base angles of B=70o +/- 3 ;

- The sides of the teeth shall be straight to +/- 2o ;

- The peaks and valleys shall have radii of curvature no greater than 0.4 mm [1/64”];

- The average peak-to-valley (P-V) shall be 8 mm +/- 0.5 mm;

- The heights of the peaks and the depths of the valleys shall be characterized by variations that are uncorrelated and are each normally distributed about their means with standard deviations P or V = 1.0 mm +/- 0.1 mm. The pseudorandom realization of this pattern shall not have excursions exceeding 2.8, and the pattern shall not repeat on a scale < 50cm;

- The coherence length of the pseudorandom pattern within one cycle shall be < 2 cm;

- The pitch of the teeth is variable and is determined by constraints introduced in items a., b., and c. above. The typical pitch is 6mm.

- The overlap region of the baffle where they are joined together shall be designed so that only one of the two baffle surfaces has serrations projecting into the aperture: an overlap of < 5 teeth is allowable.

- If the baffle is fabricated from segments which are joined, the random pattern of serrations at the joint must match up so that no discontinuous steps (< 0.010” or 0.25 mm) are present in the serration pattern.

- Successive baffles placed into the beam tube should be arranged so that the serration patterns from baffle-to-baffle are staggered randomly with at least a 5o scatter (this corresponds to rotational variations of 2” or 5cm at the tube wall).

  - Interfaces

Beam tube baffles are installed inside the beam tube at specified interval with the help of expansion joint (there is no welding between the beam tube wall and baffle).

    - Gate valves

  - Purpose

Once UHV is achieved inside the interferometer, it has to be maintained throughout the operation. For any modifications and changes in the interferometer components, the vacuum has to be broken in a given isolatable-volume while isolating other neighboring volumes. For this purpose, gate valves of different sizes will be used at different places. These large gate valves are either electrically (preferred) or pneumatically actuated and operated manually. Gate valves of standard sizes will be used in the interferometer. A typical gate valve used in LIGO-US is shown in fig. 2.11.

Fig. 2.11: A large gate valve used in LIGO-US

Large size gate valves will be used for isolation of cryopumps and some of them also serve to isolate the long sections of beam tube. Valve actuation or other intermittent device operation shall induce no more than 0.01 g peak to peak acceleration (E940002-02) at any point within 1 meter of any HAM or BSC chamber. Large size gate valves will have following features

- Gate valves are constructed using stainless steel and will be bakeable to 150 C  20 C.

- All internal mechanisms should be non-lubricated.

- Gate valves should be provided manual position lock for both open and closed condition.

- Gate valves shall be provided with double O-ring grooves in the mating flanges except the gate valves to be used for beam tube isolation. The gate of the gate value however requires only one O-ring.

The small size gate valves will be used for closing pumping ports and isolating ion pumps. These are UHV valves constructed of stainless steel with metal bonnet seals, Conflat flanges and vacuum-baked Viton gate seals. They are non-lubricated and bakeable up to 150 C. They are equipped with limit switches and padlockable devices to prevent inadvertent opening. These small sizes gate valves used at many locations (especially at beam tube sections) has to be pumped from other side before being opened. Gate valves should not be kept in close condition for long time duration to avoid the flattening of Viton seal.

    - 80 K cryopump

- Purpose

The 80K Cryopumps (ref. Anexure A) together with the ion pumps are designed to provide “Vibration free” pumping during normal operation of LIGO interferometer. The 80K Cryopumps will help pump out water and contaminating hydrocarbons which are emitted by the interferometer components. These Cryopumps will also protect the beam tube by effectively blocking water and other condensable entering the beam tube from the instrumentation kept inside the chambers.

The 80K Cryopumps of two sizes will be used  i.e. long cryopumps having flange to flange length of 5283 mm (208) while short cryopumps having flange to flange length of 2794 mm (110) as shown in fig. 2.12.

Long cryopump

Short Cryopump parts details

Short cryopump

Fig. 2.12: 80 K long and short cryopump shown with transparent outer shell

The enclosure of Cryopumps will be made of SS 304 / 304 L. The pump reservoir, which contains liquid nitrogen, is an aluminum inner vessel (SB209, 6061, T6) designed for the worst combination of pressure and temperature. Dissimilar metal joints of aluminium Cryo-pump to cryogen supply line made of stainless steel will ensure thermal fatigue requirements.

Stainless steel rod support the weight of the reservoir and restrain it against lateral and longitudinal seismic accelerations. Springs are included in the support rods and restraints to mitigate vibrations from the pump reservoir to the outer shell. Since gate valves are located at each end of the Cryopump, there is a possibility of over pressurization due to malfunction during the regeneration cycle.  To mitigate the consequences of this condition and to prevent a rupture, a pressure relief device has been designed for the outer shell.  This device will maintain the vacuum seal but it will open before substantial positive pressure is attained in the outer shell.

The 80K Cryopump is made up of shield, reservoir, LN2 inlet line, vent line, regeneration line etc. The purpose of shield is to reduce the heat load on the reservoir. The purpose of reservoir is to store Liquid nitrogen and provide cold surface area for pumping. The other auxiliary requirements are storage tank, air purge arrangement, evaporator, strainer etc. The storage tank should have storage capacity of LN2 for continuous operation of 90 days. The Cryopump is regenerated by hot GN2, when pump becomes saturated.

- Dimension:

80K Cryopump mainly consists of following components

- Outer Cryostat

  - Cylindrical shell

  - Conical shell

  - Lifting lug and

  - Stiffener

- Shield Assembly

- Reservoir

- Ports

- Flanges

- LN2 Inlet line

- GN2 Vent line

- Regeneration line, and

- Level control system etc.

The dimensional details for 80K Cryopump is mentioned in below Table 2.13.

Table 2.13:Dimensions of 80K Cryopump.

| Sr. No. | Parameters | Values [mm (inch)] |
| --- | --- | --- |
| 1 | Outer Cryostat (SS304L) |  |
|  | Cylindrical shell ID | 2019 (79.5) |
|  | Conical shell small end ID | 1133 (44.62) |
| 2 | Shield Assembly (Aluminum 5052 series) |  |
|  | Cylindrical shell ID | 1953 (76.875) |
|  | Conical shell small end ID | 1130 (44.5) |
| 3 | Reservoir Assembly(Aluminum 6061) |  |
|  | Outer cylinder ID | 1486 (58.5) |
|  | Inner cylinder ID | 1346 (53) |
| 4 | Flange ID (SS304L) | 1133 (44.62) |
| 5 | Pumping port | 01/254(10) |

- Tolerances for cryopump

The tolerance of Cryopump will be as per the reference document mentioned in Annexure A.

- Interfaces

The cryopump, interfaces with Spool and gate valve at corner and at end station.

  - Vacuum pumps for LIGO-India interferometer

An ultimate base pressure of ~ 10–9 mbar inside the vacuum enclosure is desired for detector’s sensitivity. To achieve UHV of such kind, the overall gas load and external leaks have to be taken care properly. External leaks could be avoided with proper welding scheme while gas load could be suppressed by baking the vacuum system and proper distribution of vacuum pumps. The complete vacuum system comprises of BSC chambers, HAM chambers, Inter-connecting Spool & Adaptors, cryo-pumps and beam tubes. Based on the locations of these components and their independent operation scenarios, the entire pumping system of LIGO-India detector is classified into three categories as discussed below:

  - LVEA Pumping and pump location

The LVEA contains all HAM, BSC (BSC-1 to BSC-3), Spools & Adaptors, MC tubes and Large 80 K Cryopumps as shown in Fig.2.13.

Fig.2.13: Major vacuum components of LVEA area.

HAM-1 and HAM-6 are isolated from rest of the components of LVEA by septum plates installed between HAM-1 & HAM-2 and HAM-6 & HAM-5 respectively. HAM-1 & HAM-6 will be provided with separate pumping facility. If it is considered that the total gas load and the diffusion from a single HAM chamber with its adjacent tube or chamber interconnecting flange joints to be 1.0  10–9 mbar l/s and if there are 100 joints of such kind associated to it, then the total leakage will be 1.0  10–7 mbar l/s. Hence the required pumping speed to hold 1.0  10–9 mbar will be about 100 l/s. Considering this aspect with additional precautions, each HAM-1 & HAM-6 will be facilitated separately with a scrolled dry pump (100 m3/hr.), a magnetic levitated turbomolecular (TMP) pump (500 l/s), an Ion pump (500 l/s) and one (01) no of (pirani & ion) gauge.

HAM-1 and HAM-6, each has four (04) numbers of large custom made flanges having double O-rings enclosing the annulus region in between. To reduce the permeation of hydrogen gas, all four annulus regions of each HAM-1 & HAM-6 will be pumped together by using 55 l/s ion pumps separately. After installation at site, initially HAM-1 & HAM-6 will be degassed under vacuum condition at a bake-out temperature of 150 C  20 C using heating pads till the reduction in water vapour content will be observed. In order to reduce the heat loss to the atmosphere during baking, these chambers will be insulated with thick glass fibre jackets. Since these two chambers may be vented frequently due to maintenance requirement, precautions must be taken during venting with air under class 100 environments.

From HAM-2 to HAM-5 and up to both large 80K cryopumps including BSC chambers, this space will be pumped using two (02) sets of positive displacement dry rotary pump of capacity 100 m3/hr) , four (04) numbers of magnetic levitated TMP (2000 l/s) backed by dry scroll pump of capacity 10 cfm and four (04) numbers of ion pumps (each has 2500 l/s pumping speed) distributed around both arms. These ion pumps will be installed on the top of the MC Tubes. Two (02) nos (pirani & ion) gauges at two arms and one (01) no (pirani & ion) gauge on BSC chamber will be used for vacuum measurement. The annulus spaces of double O-rings flanges of HAM-2, HAM-3, HAM-4, HAM-5, BSC-1, BSC-2, BSC-3, spools & adaptors, etc. will be pumped using numbers of 55 l/s ion pump. After installation at site, these components will be degassed under vacuum condition at a bake-out temperature of 150 C  20 C till the reduction in water vapour content will be observed.

During baking, controlled ramping of temperature will be used to prevent non-uniform temperatures or overheating at any location. In order to reduce the heat loss to the atmosphere during baking, these chambers will also be insulated with removable thick glass fibre insulation jackets. Once the base pressure inside the chamber reaches below 10–5 mbar, two (02) numbers of large size 80 K cryopumps will be used so that the required partial pressures for water and CO2 could be achieved. This space will be kept under vacuum all the time.

Each large 80 K cryopump has one (01) number of temporary TMP (500 l/s) facility for initial pumping followed with one (01) number of (pirani & ion) gauge. During cool down of cryopump, it will be kept isolated from other system by closing two numbers of gate valves installed on each ends. After commissioning of the cryopumps, the TMP will be removed and the openings will be sealed with manual 250 CF gate valves. The annulus space of the gate valve will be pumped with 55 l/s ion pump located on the gate valve bonnet. Nearly 20 nos of 55 l/s ion pumps will be used in LVEA. Scroll pumps will be used as backing pump for TMP so that contamination due to oil streaming shall be eliminated.

  - LVEA Interfaces with vacuum pumps

All vacuum pumps connected to the chambers of the interferometer are connected through compatible manual gate valves. All the gauges which monitor the vacuum pressure (except cold cathode gauge connected to the BSC chamber) are connected on the chambers through a gate valve.

  - LVEA Power requirement for vacuum pumps

Availability of single phase and three phase power requirement in LVEA for vacuum pumps and vacuum equipments at various locations is to be discussed with DCSEM and will be finalized later (annexure B).

  - VEA Pumping and pump location

Each VEA contains one BSC chamber, Spools & Adaptors and one short cryopump as shown in Fig.2.14.

Fig.2.14: Major vacuum components of VEA.

Each VEA pumping will have a positive displacement dry rotary pump (100 m3/hr), a magnetic levitated TMP (2000 l/s), a ion pump (1500 l/s) and one (01) no of (pirani& ion) gauge. Getter pump near the end stations shall be mounted at least 250 m away from the BSC chamber so that UV light effect and charge accumulation on test mass could be reduced. For commissioning & baking, similar procedure like LVEA pumping will be adopted.

The short cryopump at each VEA will be operated in a similar fashion to the large cryopumps. All the interspace of gate valves and closing flanges will be pumped using 55 l/s ion pumps. Nearly 08 numbers of 55 l/s ion pumps will be used in VEA.

  - VEA Interfaces with vacuum pump

Similar to LVEA, all the interfaces with pumps and gauges to the chambers will be through gate valves.

  - VEA Power requirement for vacuum pumps

Availability of single phase and three phase power requirement in VEA for vacuum pumps and vacuum equipments at various locations is to be discussed with DCSEM and will be finalized later (annexure B).

  - Beam tube Pumping system

Each arm of the beam tube is divided into two (02) numbers of beam tube modules of 2.0 km each. The beam tube module is separated using a gate valve and each module is connected with gauges (pirani & ion). During installation, the individual leak rate shall be maintained at 10–10 mbar l/s while the integrated leak rate shall be maintained at 10–8 mbar l/s. Since the outgassing rate of air baked SS 304L material used for beam tube fabrication is measured to be 5.0  10–14 mbar l s–1 /cm2 or better, considering the contribution due to gate valves, pumps and their assembly components, the total outgassing load in a beam tube module could be twice its empty value. Hence to achieve an ultimate base pressure of ~ 10–9 mbar in the beam tube, the ion pump with minimum pumping speed of 2500 l/s for N2 at 1.0  10–6 mbar will be installed at each 2.0 km distance along the tube length (either toward the end station or corner station side).

After installation of beam tube, the partial pressure of different gas species inside the beam tube should be maintained as mentioned in table 2.14. For this purpose, the outgassing of water vapour and hydrocarbons is reduced by baking each component at 150 C of each 2.0 km long beam tube module.

Table 2.14: Partial pressure of different gas species in the beam tube (Ref.T040001)

| Sr. No. | Parameters | Adv. LIGO Requirement | Goal |
| --- | --- | --- | --- |
| 1 | Vacuum | <1.0 10–9 mbar | < 1.0 10–10 mbar |
| 2 | Partial pressure for hydrogen gas | 1.0 10–9 mbar | 1.0 10–10 mbar |
| 3 | Partial pressure for water vapour | 1.0 10–10 mbar | 1.0 10–10 mbar |
| 4 | Partial pressure for nitrogen | 6.0 10–11 mbar | 6.0 10–11 mbar |
| 5 | Partial pressure for carbon monoxide | 5.0 10–11 mbar | 5.0 10–11 mbar |
| 6 | Partial pressure for carbon dioxide | 2.0 10–11 mbar | 2.0 10–11 mbar |
| 7 | Partial pressure for methane | 3.0 10–11 mbar | 3.0 10–11 mbar |
| 8 | Partial pressure for AMU 100 hydrocarbon | 2.0 10–12 mbar | 7.0 10–13 mbar |
| 9 | Partial pressure for AMU 300 hydrocarbon | 2.2 10–13 mbar | 5.0 10–14 mbar |
| 10 | Partial pressure for AMU 500 hydrocarbon | 9.0 10–14 mbar | 1.0 10–14 mbar |

The resistive baking of the tube module will be carried out using high power DC power supply. In order to reduce the heat loss to the atmosphere during baking, the beam tubes will be insulated with thick glass fibre jackets. The beam tube bakeout temperatures will be ramped up in a step of 10 C to a nominal bakeout temperature of 150 °C by manually setting the power supply currents. The temperature sensors mounted at every 250 m distance will be used as feedback signal for maintaining the beam tube at 150 °C under ambient condition.

Fig 2.15: Schematic of RGA and pumping during bakeout of beam tube module.

At every 250 m of the beam tube segment, 250 CF port openings are provided for pumping during bakeout time. The TMPs (2000 l/s) and commercially available Cryosorption pumps (4000 l/s for H2O) will be connected to beam tube module through gate valves. There will also be a provision on each port for set of vacuum gauge, leak detector and residual gas analyser as a vacuum diagnostic tool. The pumping system arrangement as shown in Fig. 2.15 is proposed for twosets of beam modules to be tested simultaneously at site. Initial pumping of the beam tube module will be carried out using a Roots pump backed by a single dry rotary pump. After commission, this pumping set will be used for LVEA pumping. Partial pressures will be monitored during the bakeout and cool down by a single RGA located at the module mid-point. After bakeout, these pumps will be removed and the openings will be closed using gate valves. All the annulus spaces of double O-rings will be pumped by small annulus ion pumps. The beam tubes after commissioning will be maintained under vacuum continuously.  During normal operation, the beam tube will be pumped only by ion pumps and 80K Cryopumps.

  - Beam tube interfaces and location of vacuum pump

The location of vacuum pump along the beam tube module is as shown in fig 2.15.

  - Beam tube Power requirement for vacuum pumps

Availability of single phase and three phase power requirement along the Beam tube for vacuum pumps and vacuum equipments at various locations is to be discussed with DCSEM and will be finalized later (annexure B).

  - Evacuation of assembled vacuum system

The vacuum system of LIGO-India consists of all the chambers, mode cleaner tubes, couplers, 80K cryopumps and beam tube modules, all assembled and aligned. As discussed in above and as shown in Fig. 2.16, the entire vacuum assembly of the interferometer is divided into three sections. Section 1 contains beam tube and 80K Cryopumps while Section 2 has the corner station and Section 3 the end stations. Each Section will be evacuated to ~ 10–9 mbar separately. After assembly of the vacuum enclosure, its evacuation will be carried out as per the following steps

Pump down after initial assembling of vacuum enclosure

- Initially, a dry rough pumping system will be used to obtain ~ 10–2 mbar, before changeover to TMP.

- After achieving ~ 10–2 mbar, TMP stations with ~ 2,000 l/s pumping speed will be used for further pump down.

- After achieving ~10–7 mbar with TMP stations, the cryopumps will be used to reduce water load in the beam tube. Finally ion pumps are used to reach the desired vacuum for the interferometer.

- The continuous pumping will be carried out for days to achieve final vacuum ~ 10–9 mbar. The gate valves separating Section 1, Section 2 and Section 3 will be opened to allow for complete pumping of the interferometer.

- Finally, the system will be pumped only by the ion pumps and online 80K Cryopumps with possible addition of non-evaporable getters. All these pumps are vibration free.

Fig 2.16: Schematic showing individual stations and all the vacuum enclosures.

Pump down after final assembly of vacuum enclosure along with optical components

After achieving the required vacuum, the beam tubes will be isolated from the chambers by closing the gate valves to mount optics and other components inside the chambers. The chambers will be closed after mounting components and structure and evacuation of the chambers will be carried out as follows:

- Initially, a dry rough pumping system will be used to obtain ~ 10–2 mbar, before changeover to TMP.

- After achieving ~ 10–2 mbar, TMP stations with ~ 2,000 l/s pumping speed will be used for further pump down.

- After achieving ~10–7 mbar with TMP stations, the cryopumps are used to reduce the water load in the beam tube and later ion pumps are used to reach the ultimate vacuum in the interferometer.

- Continuous pumping will be carried out for long duration to achieve final vacuum ~ 10–9 mbar after which the gate valves separating Section 1, Section 2 and Section 3 will be opened to allow for the complete pumping of the interferometer.

- Finally, the system will be pumped only by the ion pumps and online 80K cryopumps, with possible addition of non-evaporable getters.

  - Vacuum system basic design and fabrication requirement: Other than the beam tube.

The design of vacuum systems will be based on the codes & standards, defining the 	loads, manufacturing processes, surface conditions, choosing the materials, defining the 	welding process and the boundary conditions.

    - Codes & Standards

The codes & standards to be followed for design, welding and testing of different vacuum components are defined in table 2.15.

Table 2.15: Details of codes & standards for design and testing.

| Equipment | Process | Codes & Standard |
| --- | --- | --- |
| Beam Splitter Chamber / Horizontal Access Module / Spools and Adaptor | Design | ASME BPVC, Section VIII, Division 1 |
|  | Welding | ASME BPVC Welding and Brazing Qualifications, Section IX. |
|  | Testing | ASTM E498-Standard Test Methods for Leaks Using the Mass Spectrometer Leak Detector. |
| Expansion Joint | Design | EJMA, Standards of Expansion Joint Manufacturer's Association |
| Flange |  | ISO Standard 2861-Flange standards. |
| Acoustical Measurements and Noise Control |  | Noise Criteria for Heating, Ventilating, and Air Conditioning Systems. |
| Safety |  | National Fire Protection Association (NFPA) Standards, No. 70-National Electrical Code |

    - Materials & Design loads

The vacuum boundary shell material shall meet all the requirements as per standard followed by LIGO-US such that the ultimate pressure of 10-9 mbar can be achieved. This would mean that the material has a very low outgassing rate for hydrogen and can tolerate the loads due to dead mass, bake-out temperatures etc. Following criteria are adopted for LIGO-India system

    - Materials

The materials to be used for the fabrication of different vacuum components of LIGO-India interferometer are as follows:

- All fabricated components exposed to vacuum shall be made from type SS 304L using low carbon weld filler wire.

- Copper, Aluminum and prebaked Viton may be used for seals.

- Vacuum feed-through may utilize UHV compatible glass or ceramic.

- All surfaces shall be free from cracks, pits and porosity.

    - Design loads

All vacuum components shall be anchored to the floor or to each other in such manner as to restrict all motions to bellow units. The design of the vacuum enclosure shall allow for strains and stresses due to the following

- Normal cycling of the station HVAC (heating, ventilating and air conditioning) system expected to maintain temperature within a range of  1 °C

- Vacuum cycling of other sections of the vacuum enclosure.

- Bake-out of any vacuum section up to 150 C.

    - Fabrication & Technical requirements

The fabrication & technical requirements will determine material handling & procedure, welding process, workmanship and special environmental requirements required to ensure the quality of the vacuum system and beam tubes.

    - Material handling & procedure

- Carbon steel rollers shall be covered with heavy (paper or carpet) or SS during the rolling process to prevent carbon steel contamination of the stainless steel.

- The seam edges of plates to be rolled are to be pre-worked to assure roundness of the final cylinder.

- No grinding with abrasive wheels, cloth or stones is allowed on the internal vacuum surface unless specified in the specification.

- Potential hydrocarbon contamination shall be prevented.

- The material shall be wrapped and covered at all times the material is not being processed to minimize possible exposure to contaminants.

- No iron, carbon steel or other contaminants (such as grease, oil or hydrocarbons) are to come in contact with the vessel interior surfaces during material handling and assembly.

- Machining fluids shall be water soluble and free of oil and sulfur.

- During fabrication process, hydrogen gas is not permitted.

- External support structures may be welded using Gas metal arc welding (GMAW) process.

- All attachments to the vessel shall be tungsten arc welding (GTAW).

- All weld wire and weld preparation areas shall be cleaned with CO2 scrubbing prior to welding.

Steps to be followed for the realization of chambers are schematically shown in fig. 2.17.

Fig.2.17: Major steps to be adopted for realization of chambers

    - Welding

Welding of different components of the vacuum system and the connecting flanges shall be carried out as follows

- All welding exposed to vacuum shall be done by the tungsten-arc inert-gas (TIG) process or vacuum compatible welding.

- Before welding, the surfaces should be cleaned according to the UHV cleaning procedure.

- Full penetration welds shall be used and wherever possible the vacuum welds shall be internal and continuous.

- All external welding for stiffeners / structural purposes shall be intermittent to eliminate trapped volumes.

- Defective welds shall be repaired by removal to sound metal and re-welding.

- All vacuum weld procedures shall include steps to avoid contamination of the heat affected zone with air, hydrogen, or water.

- Welds to be smooth but not flush & not ground.

- Shell should be welded on full thickness without leaving any gap from inside or outside.

- All welding procedures, weld and welder qualification followed as per ASME section IX except some special welding requirement.

    - Workmanship

- The finished product shall be free of weld spatter, cutoff spatter, free iron, weld oxidation and defects.

- There shall be no grinding or abrasion of completed welds or internal vacuum surfaces.

- Helium leak test will be carried out.

    - Special environmental requirements

The LIGO vacuum equipment and laser instrumentations are potentially highly sensitive to vibration and acoustic noise, shock-induced damage or misalignment, electromagnetic interference and contamination. Following criteria should be followed in the final phase of pumping i.e. after rough pumping phase:

- Shock -To eliminate shock due to large gate valves, electrical actuation based components will be used.

- Acoustic Noise -Acoustic noise from all simultaneously operating equipment in normal operation shall not exceed NC-20 (Noise Criterion 20) at any location within LIGO vacuum equipment and laser areas.

- Vibration -The vacuum equipment shall be designed in such a way that vibration from all simultaneously operating vacuum equipment, in the absence of vibration from other sources, shall not induce motion of the walls of any vacuum chamber or of the facility floor within 1.0 m of any chamber which exceeds the following spectral density limits as listed in table 2.16.

Table 2.16: Maximum allowable spectral density of chamber or floor vibration induced by operation of LIGO vacuum equipment.

| Sr. No. | Frequency band (Hz) | Vibration limit (m/s2/√Hz) |
| --- | --- | --- |
| 1 | 0.1-10 | 3 10–11 |
| 2 | 10-1000 | 3 10–9 (Hz/f)2 |
| 3 | 1000-10000 | 3 10–15 |

- Particulates and other contaminants –Travelling clean room will have ISO class 5 or better conditions, especially when the chambers are open

    - Mitigation of Shock, Vibration & Acoustics

The technical approach which mitigate risk to the interferometer due to shock, vibration & acoustics caused by few main components are described as follows

- The first stage blower and backing pumps used for main roughing pump system will be kept in the Mechanical Equipment Room using vibration isolators, rubber bushing and washers.

- The turbomolecular pump along with its backing pump shall be connected using suitable interface.

- The power supplies of ion pumps will be located in the close vicinity of the ion pumps.

- All large gate valves will be completely supported from below to isolate them from the facility floor.

- The vent and purge system will be placed inside the Mechanical Equipment Room. The discharge and suction of the system in the corner station have mufflers or sound attenuators. The mid and end stations systems are not operated during normal operation.

- The 80K pumps will produce vibrations due to the formation and collapse of bubbles in the liquid nitrogen. The generation of large bubbles via the inlet pipe has been reduced by bringing the stratified flow from the inlet pipe above the liquid reservoir. The incoming liquid flows gently down a chute into the reservoir while the gas escapes without bubbling through the liquid. The bubbles generated from the boiling liquid in the reservoir are smaller and generate higher frequencies. Vibration transmission into the interferometer resulting from this action is reduced by low frequency isolators. An additional source of vibration from the 80K pump operation is due to vibration in the supply and return lines which are attenuated by the use of flex lines.

    - Vent & Purge systems

In order to carry out the required modifications and changes in the interferometer components, the vacuum chamber shall be vented using clean air. For this purpose, one 200 CFM clean air supply system will be facilitated at Laser Vacuum Equipment Area (LVEA) while one 50 CFM system will be facilitated at each Vacuum Equipment Area (VEA). Each system is located in the mechanical equipment room to minimize the noise and vibration. These clean air systems should possess the following features

- The compressor draws ambient air from the LVEA into the system through piping which is having an intake silencer & air intake filter.

- The compressor should be oil-free type to preclude the addition of hydrocarbons to the air stream.

- The compressor should be extremely quiet and have low vibration.

- The process air should be directed through one of the twin tower desiccant dryers for removal of water vapor to a dew point of – 40 C (measured at atmospheric pressure).

During venting of vacuum chambers, the portable soft-walled cleanrooms (PSC's) are used to isolate sections of the vacuum envelope from the building environment.  They will provide a Class 100(ISO Class 5) environment for the opening of chamber ports. LVEA will have a cleanroom system (Ref: LIGO-E0900047-v23) consisting of two BSC cleanrooms, one HAM cleanroom, and a gowning room. These cleanroom units can also be joined to form the desired work space. Each VEA will have one BSC cleanroom and a gowning room. A total of eight (08) PSC's will be provided. Each mid-station will also have integral gowning room. The size of each BSC cleanroom will be is 5700 mm (L)  5700 mm (W)  4900 mm (H). This cleanroom will have an opening in the ceiling to accommodate the 2743.2 mm diameter chamber dome, which can be lifted by a crane up through the cleanroom ceiling approximately it's entire length without compromising the Class 100 (ISO Class 5) environment. Similarly, the size of each HAM cleanroom will be is 5150 mm (L)  3385 mm (W)  3650 mm (H). It is similar in design to the BSC cleanrooms, but without the dome accommodation or integral gowning.

Class 100 (ISO Class 5) environment will be achieved in all cleanrooms by utilizing blowers with pre-filters ducted to a HEPA filtering system.  The filters will have dampers for initial flow balancing. The blowers will have variable speed controllers to reduce the power consumption during periods of inactivity, or will be single speed ducted to allow half to be turned off during such times. All cleanroom frames will be constructed with painted tubular steel.  Trusses for overhead support will be used so that the cleanrooms can be lifted and moved into place by the building overhead crane. The cleanrooms will also be equipped with a removable six foot square window on at least three sides. The vinyl sections provided for this window will seal around the tube or nozzle entering the cleanroom. An additional two-inch diameter hole will be provided in this window to be used for laser alignment during equipment installation. The bottom of the six foot window will have a velcro-sealed slit directly below it, extending to the bottom of the wall to allow lowering the cleanroom into place over the tube.

    - Reliability and Quality Assurance

- No material shall be used for fabrication until its acceptability is assured.

- Fabrication process, testing, handling etc. shall be assured before adapting for fabrication.

- Helium leak rate of any UHV component shall not exceed 1.0  10–9 mbar l/s.

- The average hydrogen outgassing rate from all surfaces exposed to vacuum shall be <1.0  10–11 mbar l s–1/cm2 at room temperature ( 300 K) for BSC and HAM chambers.

- After fabrication of any components involving a large number of welding (specially lower parts of BSC / HAM chambers), the baking at 400 – 450 C in nitrogen / argon environment will be carried out so that stress gets released.

- The inner surfaces of baked components (BSC / HAM chamber) if oxidized then electrochemical buffing will be carried out to remove oxidation.

- Detailed plans (including descriptions of the test equipment and procedures) for the qualification, screening, contamination control and acceptance tests shall be worked out prior to use.

    - UHV Cleaning

All surfaces exposed to vacuum shall be cleaned in accordance with procedures approved by LIGO-India prior to fabrication and installation. The following steps shall be adopted for the use in ultra-high conditions

    - General procedures

- No visible contaminant (viewed with the naked eye, under both natural and ultraviolet light) of any form shall be left within the vacuum enclosures when installed (for example: water, dust, sand, hydrocarbon film, etc.).

- Personnel shall wear full clean-room garb (T1400236) when walking on / working in the interior surfaces.

- Cleaning will be applicable to any fabricated stainless or aluminum component that is exposed to UHV service.  It applies to the cleaning of these components subsequent to completion of all machining and welding operations. Drying will be accomplished by blowing clean air over the component before allowing it to cool. Inspect and bag the component (or close up the vessel) as soon as possible after drying. Avoid contamination from unfiltered shop air through open doorways.

- The presence of any hydrocarbon or fingerprints on any interior surface or flange face shall be cause for rejection.

- Surface recontamination shall be prevented during all subsequent processes.

- Immediately after drying and inspection, double bag the component using clean, oil-free polyethylene bags or wrap and seal using the same material. Avoid the use of materials that shed particulates such as newspaper. Bubble wrap is recommended for many parts.

    - Special procedures

- All components shall be thoroughly cleaned to remove all ink, oil, grease, dirt and chips. All inks (stamping, Sharpie, etc.) should be removed with solvent such as acetone or isopropanol.

- Use Soap and water to remove machining fluids.

- Thoroughly rinse soap with clean water (DI / distilled preferred).

- Parts should be dried thoroughly with clean air, nitrogen, or lint-free cloth.

- Wrap parts to protect from damage and to maintain cleanliness during shipment.

- Stainless steel parts shall be pickled and passivated (citric acid preferred) at room temperature. The pickling and passivation process formulation, time and temperature should be chosen to achieve slight chemical etching (< 0.0005 inch). (Ref. ASTM A380 for pickling and ASTM A967 for passivation).

- Rinse parts with DI or distilled clean water. After rinsing, visually inspect the surfaces for cleanliness.

- Dry Parts using dry nitrogen blow-off or bake at approximately 60 °C.  Re-inspect for cleanliness.

- Electro-polishing of Stainless Steel parts is allowed. If electro-polishing is carried out on a part, then Pickling and Passivation of Stainless Steels have to be carried out.

- Aluminum parts shall be washed with Alkaline Soak Cleaner and may be electro-polished to achieve a material removal of .001 inch per surface.

    - QA / QC plans

In order to monitor the quality during fabrication, the quality assurance / control (QA / QC) shall be used as a guideline document during the entire fabrication and acceptance process. QA / QC at different stages can be elaborated as follows

    - Before fabrication

- Design specification – (a) verify, review and acceptance of fabrication drawing and its specification, (b) preparation of design calculation sheet and its approval from engineering section, (c) ensuring implementation method of design change and re-approval.

- Material control – (a) verify specification of ordered material, (b) identification and testing for compliance with specification.

- Examination & Inspection – (a) verify factory test certificates, (b) establish procedures for examination.

- Process control – (a) fabrication sequence, (b) welding procedure, (c) welder qualification.

- During fabrication

- Inspection records of processes like (a) cutting and rolling plans, (b) weld inspection reports, (c) heat treatment, (d) visual check, UT, RT etc. reports.

- Surface finish of the material and sealing surfaces.

- Dimensional records

- Calibration of the instruments / equipment used for testing & acceptance standards for testing.

- Reporting of non-conformities and their rectification.

    - After fabrication

- Leak tightness test of the component by (1) Sniffer mode and (2) Vacuum mode using helium gas.

- Performance test of the complete system to demonstrate its design capability by ultimate vacuum test.

- Proper handling and careful shifting to avoid damage during transportation.

An overview of typical QA / QC for UHV component fabrication can be represented as shown in table 2.17.

Table 2.17: Quality assurance plan for UHV component fabrication.

| Sr. No. | Process description | Format for report | Action list | Acceptance standard |
| --- | --- | --- | --- | --- |
| 1 | Fabrication drawing | Approval | Review by QAC | UHV practice |
| 2 | Raw material | Test report | Chemical / Physical / Visual / UT | ASTM / ASME Boiler & Pressure vessel |
| 3 | Fabrication sequence | Flow chart | Review by QAC | UHV practice |
| 4 | Welding | WPS / PQR / WPQ | Soundness / Ductility / Strength | ASME Boiler & Pressure vessel |
| 5 | Fit-up / welding | Weld inspection | Root LP, Visual / Final RT / UT | ASME Boiler & Pressure vessel |
| 6 | Port machining | Inspection | Machine accuracy | ISO / Drawing |
| 7 | Flanges | Inspection | Measurement & weld | Drawing / ASME code |
| 8 | Leak testing | Inspection | Leak rate | 1.0 10–10 mbar l/s |
| 9 | Cleaning / polishing | Inspection | Process / Surface finish | Vacuum practice |
| 10 | Ultimate | Inspection | Pressure | 1.0 10–9 mbar |

  - Vacuum system basic design and fabrication requirement: Beam tube.

The design of the vacuum components are based on the codes & standards, defining the loads, manufacturing processes, surface conditions, choosing the materials, defining the welding process and the boundary conditions.

  - Codes & Standards

The codes & standards to be followed for design, welding and testing of different vacuum components are defined in table 2.18.

Table 2.18: Details of codes & standards for design and test.

| Equipment | Process | Codes & Standard |
| --- | --- | --- |
| Beam tube | Design | ASME BPVC, Section VIII, Division 1 |
|  | Welding | ASME BPVC Welding and Brazing Qualifications, Section IX. |
|  | Testing | ASTM E498-Standard Test Methods for Leaks Using the Mass Spectrometer Leak Detector. |
| Expansion Joint | Design | EJMA, Standards of Expansion Joint Manufacturer's Association |
| Flange |  | ISO Standard 2861-Flange standards. |
| Safety |  | National Fire Protection Association (NFPA) Standards, No. 70-National Electrical Code |

  - Materials & Design loads

Material used for beam tube will be SS304L with low hydrogen content to achieve ultimate pressure less than 10-9 mbar.

    - Materials

The materials to be used for the fabrication of different vacuum components of LIGO-interferometer are as follows

- All fabricated components exposed to vacuum shall be made from type SS 304L using low carbon weld filler wire wherever required.

- Copper seals will be used at beam tube vacuum pump interfaces.

- All beam tube surfaces shall be free from cracks, pits and porosity.

- Air baked raw material with acceptable outgassing levels for hydrogen shall be used.

- HRAP (hot rolled, annealed and pickled) surface is acceptable.

- Any alternate finish is subject to approval.

    - Design loads

Beam tube supports shall be anchored to the floor or to each other in such manner as to restrict all motions to bellow units. The design of the tube shall allow for strains and stresses due to the following:

- Variations in atmospheric pressure.`

- Vacuum cycling of other sections of the vacuum enclosure.

- Bake-out of tube to 150 C.

In case of beam tube fabrication, the following additional design loads shall be considered as shown in table 2.19.

Table 2.19: Details of Design Load Requirements.

| Sr. No. | Parameters | Values |
| --- | --- | --- |
| 1 | Dead load | ~ 2.2 ton / 20 m (for thickness – 3.2 mm); It depends on the length of beam tube & thickness available. |
| 2 | Wind load | Depends upon the site (as per wind speed). |
| 3 | Seismic load | Depends upon the site (Seismic zone &vibration measurement). |
| 4 | Vacuum load | Outside – Atmospheric pressure. Inside – UHV (~ 0 pressure). |
| 5 | Expansion joint force | As per baking adopted methodology. |
| 6 | Differential settlement load | 25.4 mm for differential between fixed and guided support. |
| 7 | Insulation load due to wrapped insulation | Depends on the selected insulation material |

  - Fabrication & Technical requirements

Fabrication and technical requirement for beam tube will be similar to that of the vacuum chambers other than the points mentioned below.

    - Material handling & procedure

Due to transportation and handling difficulties, the complete beam tube will be fabricated using small tube sections. The technical requirements for the beam tube is mentioned in table 2.20 and the steps to be followed for the realization of beam tube is shown in Fig.2.18.

Table 2.20: The basic technical requirements for beam tube.

| Sr. No. | Parameters | Values |
| --- | --- | --- |
| 1 | Clear Aperture | 1070 mm |
| 2 | Inner Diameter | 1240 mm |
| 3 | Raw Material | SS 304L (sheet or coil) |
| 4 | Air Bake (Raw material) | 440 8 C for 36 hours (to reduce hydrogen content) |
| 5 | Baffle | 213 nos. per 4-km, Installed without welding |
| 6 | Welding | TIG/Plasma (No filler wire) with inert purge gas |
| 7 | Vacuum baking after installation | Up to 150 °C |
| 8 | Support structure adjustment range | ± 75 mm in both the vertical and horizontal |
| 9 | Tube axis height above the nominal (mid variation range) slab plane | approximately 1100 mm |

Fig.2.18: Major steps to be adopted for realization of beam tube.

    - Welding

The welding of beam tube and the connecting flanges shall be carried out as follows

- The beam tube will be either spiral welded or seam welded.

- All welding exposed to vacuum shall be done by the tungsten-arc inert-gas (TIG) process or vacuum compatible welding.

- Before welding, the edges preparation will be carried out.

- Full penetration welds shall be used and wherever possible the vacuum welds shall be internal and continuous.

- All external welding for stiffeners / structural purposes shall be intermittent to eliminate trapped volumes.

- Defective welds shall be repaired by removal to sound metal and re-welding.

- All vacuum weld procedures shall include steps to avoid contamination of the heat affected zone with air, hydrogen, or water.

- Welds to be smooth but not flush & not ground.

- Beam tube should be welded on full thickness without leaving any gap from inside to outside.

- All welding procedures, weld and welder qualification followed as per ASME section IX except some special welding requirement.

    - Workmanship

- The finished product shall be free of weld spatter, cutoff spatter, free iron, weld oxidation and defects.

- There shall be no grinding or abrasion of completed welds or internal vacuum surfaces.

- Helium leak test will be carried out.

    - Special environmental requirements

- Particulates and other contaminants - Traveling clean rooms to reach ISO Class 5 will be used during installation of beam tube.

- Environment -The beam tube module shall be designed to withstand ambient conditions existing at the sites (10 - 50 C). Tube enclosures will enclose the tube at a planned interval after the beam tube sections are welded and aligned as per the requirement.

  - Vent & Purge systems

In case of emergency dry nitrogen will be used and venting process similar to the chamber will be adopted.

  - Reliability and Quality Assurance

- No material shall be used for fabrication until its acceptability is assured.

- Fabrication process, testing, handling etc. shall be assured before adapting for fabrication.

- Helium leak rate any UHV component / each beam tube module shall not exceed 1.0  10–10 mbar l/s.

- The average hydrogen outgassing rate from all surfaces exposed to vacuum shall be <1.0  10–14 mbar l s–1/cm2 at room temperature (RT  300 K) for the beam tubes.

- The stainless steel sheets used for fabrication of UHV beam tubes will be pre-baked at 450 °C in dry air for 36 h using dry air and hydrogen free weld-wire during welding to reduce hydrogen concentration significantly.

- Samples from all air baked raw materials shall be tested to confirm acceptable outgassing level for hydrogen 5.0  10–14 mbar l s–1/cm2 at RT.

- Each beam tube module shall have a minimum clear aperture of 1.07 m.

- The beam tube sections shall be fabricated, assembled and tested in accordance with a procedure adopted for measuring the alignment, outgassing rates and leak rate.

- The installation of beam tube segments at site shall be carried out using portable clean room environment.

- Detailed plans (including descriptions of the test equipment and procedures) for the qualification, screening, contamination control and acceptance tests shall be worked out prior to use.

  - UHV Cleaning

All surfaces exposed to vacuum shall be cleaned in accordance with procedures approved by LIGO-India prior to fabrication and installation. The following steps shall be adopted for the use in ultra-high conditions

    - General procedures

- No visible contaminant (viewed with the naked eye, under both natural and ultraviolet light) of any form shall be left within the vacuum enclosures when installed (for example: water, dust, sand, hydrocarbon film, etc.).

- Personnel shall wear full clean-room garb (T1400236) when walking on / working in the interior surfaces.

- Cleaning will be applicable to any fabricated stainless or aluminum component that is exposed to UHV service.  It applies to the cleaning of these components subsequent to completion of all machining and welding operations. Drying will be accomplished by blowing clean air over the component before allowing it to cool. Inspect and bag the component (or close up the vessel) as soon after drying as possible. Avoid contamination from unfiltered shop air through open doorways.

- The presence of any hydrocarbon or fingerprints on any interior surface or flange face shall be cause for rejection.

- Surface recontamination shall be prevented during all subsequent processes.

- Immediately after drying and inspection, double bag the component using clean, oil-free polyethylene bags or wrap and seal using the same material. Avoid the use of materials that shed particulates such as newspaper. Bubble wrap is recommended for many parts.

    - Special procedures

- All components shall be thoroughly cleaned to remove all ink, oil, grease, dirt and chips. All inks (stamping, Sharpie, etc.) should be removed with solvent such as acetone or isopropanol.

- Use Soap and water to remove machining fluids.

- Thoroughly rinse soap with clean water (DI / distilled preferred).

- Parts should be dried thoroughly with clean air, nitrogen, or lint-free cloth.

- Wrap parts to protect from damage and to maintain cleanliness during shipment.

- Stainless steel parts shall be pickled and passivated (citric acid preferred) at room temperature. The pickling and passivation process formulation, time and temperature should be chosen to achieve slight chemical etching (< 0.0005 inch). (Ref. ASTM A380 for pickling and ASTM A967 for passivation).

- Rinse parts with DI or distilled clean water. After rinsing, visually inspect the surfaces for cleanliness.

- Dry the parts using dry nitrogen blow-off or bake at approximately 60 °C.  Re-inspect for cleanliness.

  - QA / QC plans

In order to monitor the quality during fabrication, the quality assurance / control (QA / QC) shall be used as a guideline document during the entire fabrication and acceptance process. QA / QC at different stages can be elaborated as follows

    - Before fabrication

- Design specification – (a) verify, review and acceptance of fabrication drawing and its specification, (b) preparation of design calculation sheet and its approval from engineering section, (c) ensuring implementation method of design change and re-approval.

- Material control – (a) verify specification of ordered material, (b) identification and testing for compliance with specification.

- Examination & Inspection – (a) verify factory test certificates, (b) establish procedures for examination.

- Process control – (a) fabrication sequence, (b) welding procedure, (c) welder qualification.

    - During fabrication

- Inspection records of processes like (a) cutting and rolling plans, (b) weld inspection reports, (c) heat treatment, (d) visual check, UT, RT etc. reports.

- Surface finish of the material and sealing surfaces.

- Dimensional records

- Calibration of the instruments / equipment used for testing & acceptance standards for testing.

- Reporting of non-conformities and their rectification.

- After fabrication

- Leak tightness test of the component by (1) Sniffer mode and (2) Vacuum mode using helium gas.

- Performance test of the complete system to demonstrate its design capability by ultimate vacuum test.

- Proper handling and careful shifting to avoid damage during transportation.

An overview of typical QA / QC for UHV component fabrication can be represented as shown in table 2.21.

Table 2.21: Quality assurance plan for UHV component fabrication.

| Sr. No. | Process description | Format for report | Action list | Acceptance standard |
| --- | --- | --- | --- | --- |
| 1 | Fabrication drawing | Approval | Review by QAC | UHV practice |
| 2 | Raw material | Test report | Chemical / Physical / Visual / UT | ASTM / ASME Boiler & Pressure vessel |
| 3 | Fabrication sequence | Flow chart | Review by QAC | UHV practice |
| 4 | Welding | WPS / PQR / WPQ | Soundness / Ductility / Strength | ASME Boiler & Pressure vessel |
| 5 | Fit-up / welding | Weld inspection | Root LP, Visual / Final RT / UT | ASME Boiler & Pressure vessel |
| 6 | Port machining | Inspection | Machine accuracy | ISO / Drawing |
| 7 | Flanges | Inspection | Measurement & weld | Drawing / ASME code |
| 8 | Leak testing | Inspection | Leak rate | 1.0 10–10 mbar l/s |
| 9 | Cleaning / polishing | Inspection | Process / Surface finish | Vacuum practice |
| 10 | Ultimate | Inspection | Pressure | 1.0 10–9 mbar |

  - Vacuum Control and Monitoring System (VCMS)

Initial pump down of the beam tube will be carried out using standalone pumping system. Control will be carried out manually for the gate valves as well as for the standalone vacuum pumps connected to the chambers. The complete Vacuum Control and Monitoring System (VCMS) will only be functional after integration of all vacuum sub-systems constituting the detector. VCMS will monitor the status of vacuum and will also control the vacuum system to maintain UHV. The different functions of VCMS is described as follows

Functions of VCMS

The primary function of VCMS is to provide remote operation and monitoring of LIGO-India Vacuum Equipment (VE) during detector operation. Other functions involve:

- Monitoring of all sensors associated with the vacuum equipment.

- Close loop control of some specific equipment such as Cryopump, etc.

- Appropriate interlocking of the vacuum systems.

- Local operation of all vacuum equipment.

- Operator interfaces, as necessary, to support both local and remote operation.

- Operational support services such as data archiving (short and long term) and retrieval and alarm management.

The VCMS functionality depends upon vacuum system requirements and vacuum physical parameters.

The VCMS is to be a subsystem of larger LIGO-India Control and Data System (CDS) and will be interfaced to other subsystem for operational/functional point of view.  The design and implementation of VCMS shall be in compliance with the CDS control 
and monitoring requirements. VCMS will use the CDS resources for interfacing and data communication. Fig. 2.19 shows the relationship of VCMS with LIGO-India CDS and Vacuum Equipment.

Fig.2.19: VCMS interface overview in LIGO CDS

  - aLIGO-India detector upgrade

At the time of writing this report, there are plans to upgrade the Adv-LIGO-USA detector of its performance and sensitivity by a factor of 1.7 (LIGO-G1901072). The upgradation will recoin the existing aLIGO detector as A+ detector. To achieve the said improvements one of the changes to the Advanced LIGO interferometer configuration is the addition of a filter cavity of length 300 meters. A schematic of the addition of the filter cavity to the existing LIGO detector is shown in fig. 2.20.

Fig. 2.20: Schematic of A+ configuration. Two additional HAM chambers namely, HAM7 and HAM8 will be installed for A+. HAM7 will be integrated with HAM5 and HAM8 will be integrated with HAM7 using long vacuum tube. HAM7 will be in the LVEA area while, HAM8 will be installed 300m away in an enclosure similar to the end station. The region of the tube outside LVEA will have an enclosure similar to 1.2m diameter beam tube enclosure.

Additional two HAM chambers will be installed for A+. HAM7 will be integrated with HAM5 with a cylindrical vacuum tube whose length will be defined by the accessibility of opening the HAM5 door. A 300 meter long vacuum tube constituting the filter cavity will be connected between HAM7 and HAM8. The 300 meter long tube will be made of SS304L and will have a diameter of 250mm. This tube will be a seamless tube and have an average partial pressure of hydrogen less than 10-9 mbar. The details of the vacuum requirement for the A+ upgradation will be addressed along with its interfaces as per annexure B and C.

  - ANNEXURE

Annexure A

Details of references used for the VSRD

| Component | Reference drawing | Description |
| --- | --- | --- |
| Basic Symmetric Chamber (BSC) | V049-4-003 | BSC shell Weldment/machining |
|  | V049-4-101 | BSC Overall assembly |
|  | V049-4-A11 V049-4-014 | BSC End Cover |
|  | V049-4-023 | Chamber support assembly |
|  | V049-4-025 | BSC annulus tubing assembly |
|  | V049-4-019 | 60.5inch ID flange (Grooved) |
|  | V049-4-022 | 104.5inch ID flange (Grooved) |
|  | V049-4-041 | 104.5inch ID flange (Flat faced) |
|  | V049-4-122 | Annulus Ion Pump Support |
|  | V049-4-036 | BSC Floor assembly & Details |
| Horizontal Access Module (HAM) | V049-4-128 | HAM shell Weldment |
|  | V049-4-002 | HAM assembly |
|  | V049-4-040 | HAM Bellows tie rod assembly |
|  | V049-4-053 | 60.5inch ID Metal bellow |
|  | V049-4-054 | Flange/annulus tubing assembly |
|  | V049-4-127 | 84inch ID Access Cover |
|  | V049-4-A4 | 60 inch HAM Cover, Grooved, Type A-4 |
|  | V049-4-021 | 84.25inch ID flange (Grooved) |
|  | V049-4-123 | Annulus Ion Pump Support |
| MC Tube | C1001348 | MCA |
|  | C1001349 | MCB |
|  | T0900653 | MC Tube ports location |
| Spool | V049-4-BE1 | Drawing of spool BE1 |
|  | V049-4-BE2 | Drawing of spool BE2 |
|  | V049-4-BE3A | Drawing of spool BE3A |
|  | V049-4-BE3 | Drawing of spool BE3 |
|  | V049-4-BE4 | Drawing of spool BE4 |
|  | V049-4-BE5 | Drawing of spool BE5 |
|  | V049-4-BE6 | Drawing of spool BE6 |
|  | V049-4-B1 | Drawing of spool B1 |
|  | V049-4-B9 | Drawing of spool B9 |
| Adaptor | V049-4-A1 | Drawing of adaptor A1 |
|  | V049-4-A2 | Drawing of adaptor A2 |
|  | V049-4-A3 | Drawing of adaptor A3 |
|  | V049-4-A7 | Drawing of adaptor A7 |
| Septum Plate | C1001599-v1 | GNB FDR Submittal Septum Plates |
|  | D1001661-v2 | H1, L1 Input Septum Plate Port Locations |
|  | D1001662-v2 | H1, L1 Output Septum Plate Port Locations |
|  | D1002462-v2 | AdvLIGO Systems, H1 & L1 Output Septum Plate Assembly |
| 80K Cryopump | V049-4-004 | General Arrangement-Long-Left |
|  | V049-4-005 | General Arrangement-Short-Right |
|  | V049-4-006 | General Arrangement-Long-Right |
|  | V049-4-007 | General Arrangement-Short-Left |
|  | V049-4-090 | Reservoir-Short-Right |
|  | V049-4-091 | Reservoir-Short-Left |
|  | V049-4-092 | Reservoir-Long-Left |
|  | V049-4-093 | Reservoir-Long-right |
|  | V049-4-094 | Reservoir support assembly short left and right |
|  | V049-4-095 | Reservoir support assembly long left and right |
|  | V049-4-114 | Shield detail assembly Long right hand and Left hand |
|  | V049-4-117 | Short Shield detail assembly right hand and Left hand |
|  | V049-4-118 | Weldment long left |
|  | V049-4-119 | Weldment long right |
|  | V049-4-120 | Weldment short left |
|  | V049-4-121 | Weldment short right |
|  | V049-4-132 | 44 5/8inch I.D Flange (Grooved) |
|  | V049-4-140 | Pump Leg assembly long |
|  | V049-4-141 | Pump Leg assembly short |
|  | V049-4-146 | Head /nozzle details ,Long left -/Short right |
|  | V049-4-147 | Head /nozzle details ,Long right -/Short left |
|  | V049-4-148 | Shield support |
|  | V049-4-161 | 2inch jacketed line |
| Beam tube | E950025-02-B | Tube sections construction option |
|  | E950026-06-B | Expansion joints construction option |
|  | D950031-06-B | Beam tube related |
|  | D950032-05-B |  |
|  | D950033-05-B |  |
|  | D950034-07-B | Beam tube section |
|  | D950035-08-B | Fixed support |
|  | D950036-09-B |  |
|  | D950037-10-B |  |
|  | D950039-03-B | Ports details |
|  | D950040-05-B | Vacuum stiffener details |
|  | D950041-06-B | Support rings |
|  | D950044-09-B | Guided support |
|  | D960511-03-B | Big Gate valve soft support |
|  | D960512-04-B | Termination support |
|  | D960513-02-B | Termination support ring |
| Vacuum component assembly | D0901859-v2 | Advance LIGO corner station vacuum component arrangement |
|  | V049-5-003 | Initial LIGO corner station vacuum component arrangement |
|  | D1600164 |  |
|  | V049-5-008 | Vacuum components at end station |
|  | V049-5-009 | Vacuum components at end station |
|  | T040001 | Vacuum outgassing Requirements |
|  | E950083 | Beam tube baffle |
|  | G970217-00-E | Beam tube baking |
|  | T970148-00-E |  |
|  | E0900047-v23 | aLIGO Contamination Control Plan |

Annexure B

Interface Documents with other organisation identified

      - RRCAT:

- Interface between the vacuum chambers, interfacing various ports and detector components.

- Clear aperture of baffles and tolerances for the beam tube

- Coordinates marking for BSC, HAM, large gate-valves, tubes and supports in LVEA and VEA

      - DCSEM:

- Utility (crane, power, HVAC) requirement

- Space requirement around vacuum components

- Beam tube enclosure

- A+ enclosure

- Interface between support, concrete slab and load support to the ground

- Power requirement for vacuum pumps with locations

Annexure C

Documents which will be prepared separately

- Detail QA/QC document of vacuum systems (IPR)

- Baffle details and position (RRCAT-IPR)

- Identifying dimension/specification which cannot be compromised (RRCAT-IPR)

- VCMS details (IPR)

- Roles and responsibilities of different institutions (PMB)

- Documentation trees for system engineering requirements and Integrated project scheduling  (RRCAT-IPR-DCSEM-IUCAA)

- Material, fabrication, cleaning, baking, leak testing, packaging, transportation and commissioning (IPR, DCSEM).

- A+ layout and vacuum location in LVEA and at the termination of A+

- Documents listing deviations between LIGO-US and LIGO-India

- Failure Mode and Effect Analysis (FMEA) (E960111-01-V)