## Slide 1: ITM03; LHO ITMx Post O2 inspection
The Hanford ITMx exhibited high absorption. The optic was examined and cleaned in-chamber, May 2017.
The absorption persisted, the optic was removed after O2
A feature was found at the location predicted by TCS. This feature is highly absorbing. A feature at this location was present in the scatter scans of the optic as received from the coater. Several other post-coating absorbers over 20 ppm were found in the center 60 mm of the optic.
G. Billingsley, L. Zhang
LIGO-G1800345-v1

## Slide 2: Experimental process
Scatter measurement - The map before installation: measured after First Contact cleaning with ion gun. The map after O2: measured as-returned from LHO, no cleaning
Cleaning with ion gun because there was some dust visible when checking with a flash light.
Gentle absorption: the measurement was carried out with a beam of 0.3mm in diameter and 0.9W, i.e. 12.7W/mm^2. The lower limit of sensitivity for this scan was about 20ppm.
High resolution probe of the high absorption spot. The measurement was done with a beam of 60 um in diameter at 25 mW, i.e. 8.8 W/mm^2.
All scatter and absorption data are referenced from LIGO-E1000766

## Slide 3: “Gentle” Absorption Scan
Absorber found!
Lower power and larger beam than the usual absorption scan: Noise floor ~20 ppm

## Slide 4: Probe absorption with smaller beam

## Slide 5
absorber: 155 µm across bright center

## Slide 6
absorber: 155 µm across bright center

## Slide 7: “Gentle” Absorption Scan with bottom noise chopped off
Note there was no first contact cleaning before this scan.
Possible dust added during chamber removal or installation on absorption test stand

## Slide 8: Dark Field Microscopy
Features found within a radius of ~60 mm of center.
1 mm ∅ field
Center scale squares are 50 µm

## Slide 9
Features found within a radius of ~60 mm of center.
1 mm ∅ field
Center scale squares are 50 µm
Absorbing
Interesting 2
Interesting 1
Interesting 3
(0.4, -4)
(-21, -13)
(11, 21)
(-2, 23)
(-23, -27)
(8, -30)
(-8, -34)
(11, -26)
(-60, -43)
(5, -27)
(6, -27)
(-4, -43)
(-2, -50)
(15, -41)
(27, -47)
(-2, -46)
(40, -56)
(32, -59)
(11, -61)
(-1.4, -62)
(47, -43)
(-41, -52)
(1, -44)
(-3, -33)
(8, -30)

## Slide 10: Interesting 1 (5, -27)
Branch-like center structure

## Slide 11: Interesting 2 (-8, -39)
Shown in bright field
~80 µm in diameter

## Slide 12: Interesting 3 (-60, -43)
Center is filled

## Slide 13: Distribution

## Slide 14: ITM03 scatter In chamber view, before install, after removal
Scatter before installation: 5.6 ppm. Scatter after installation 6.3 ppm
Spot # 2 found to be the high absorber,
Present in “before” and “after” scatter scans

## Slide 15: Absorption spot seen in figure data? Or dust?
Polished and coated data, 60 mm mask (orange circle) 0.06 nm rms polished, PV 0.49 nm 0.08 nm rms coated, PV 1.06 nm - 1 pixel = 400 µm
"Interesting point" #1 (page 10) is in a similar, though not exact location to the point seen low right.

## Slide 16: Next
Scanning available ITM spares to screen for possible high absorption spots using the “gentle” absorption method.
See E1000766 for ITM03 Characterization data
See T1700193 for the in-chamber inspection at LHO