## Slide 1: OFI shroud design review
Alena A., Calum T., Corey A., Eddie S., Stephen A.

## Slide 2
Robert Schofield LIGO-G1800717
Scatter from OFI during operation
Vacuum enclosure wall – speculated numbers
DLC on stainless steel Super#8 (Duralar)
Black Nickel on stainless steel super #8 (Anoplate)
DLC on stainless steel Super#8 (Richter)
Scatter from a moving vacuum enclosure wall (HAM5 ) aLog41449

## Slide 3
As a quick check, the attached figure shows that the vacuum enclosure walls of LLO HAM5 move between about 10^2 and 10^3 times more than the table top in the 10-100 Hz band. Thus, even if baffles that were mounted on the table top were as good at back-scattering light as the HAM5 doors (unlikely), and they were only 10 cm away from an OFI scattering site, a linear estimate suggests they would cause less noise than reflections off of the more distant door. Of course, for the same geometric attenuation reason, the further the baffles are from the OFI, the better.
https://alog.ligo-wa.caltech.edu/aLOG/index.php?callRep=41449

## Slide 4
https://dcc.ligo.org/LIGO-D0900456
Existing
Septum baffle
HAM 5 post SLiC part A install (LLO) – Existing as of 4/18/2018
OFI
SR3
SRM
Table baffle
Table baffle
ZM2
SRM AWC
Squeezer (transmission) beam dump

## Slide 5
Faraday Isolator Refl Baffle removal
D0901134
FRS 9647
ECR E1700425
A ghost beam from septum window was found hitting the reflection baffle (uncoated) and going back to the window
Already removed at LLO
Pending at LHO (function handled by shroud)
Refl baffle in NOT part of the new shroud!

## Slide 6
Purpose of the shroud:
Block scatter paths from OFI to vacuum enclosure and absorb the scatter with minimum bounces (Super #8 * side of panels should face the inside of the shroud)
Minimize new stray light scatter from added panels with apertures (Super #8 * side of panels should face the laser input)
Septum window ghost beam dumping
* Super #8 is only finished on one side, the other side is ~ mill finish

## Slide 7
Inspiration:
excising structure, angled walls, apertures
Every house needs a roof!
Bar / restaurant in Seattle, WA

## Slide 8
#1 High roof concept
#2 Low roof concept
Existing
Septum baffle
Squeezer input
Squeezer input
- 13 panels!
- 8 panels
Minimizes uncoated surfaces inside the shroud
Existing
Septum baffle
OFI input baffle

## Slide 9
OFI input side (panels are not yet detailed)
OFI input baffle
Aperture:
- same as output side baffle and old reflection baffle D0902845
- additional vertical and horizontal adjustment
Same bracketry as SLiC part A
Non-polished side
Trapezoidal panel pitched 10°
Polish – on the inside of the shroud
Stray light on the outside – dumping on SRM baffle
Existing
Septum baffle
Squeezer input

## Slide 10
Super #8 side
Non-polished side
Existing Septum baffle
OFI output side (panels are not yet detailed)
Ghost beam baffle
OFI output trapezoidal baffle is pinched 10°
Ghost beam baffle pitched 10° and optionally yawed with a 5° wedge
OFI output baffle
Aperture:
- same as input side baffle and old reflection baffle D0902845
- additional vertical and horizontal adjustment
Same bracketry as SLiC part A
Two panels:
OFI output baffle
Septum window ghost beam baffle
ghost beam

## Slide 11
Two panels on squeezer input side:
Table baffle (stop stray light purpose)
10° pitched rectangle
polished side facing the input
small aperture
optional yaw
Squeezer input side (panels are not yet detailed)
OFI input baffle
Existing
Septum baffle
Squeezer input
Super #8 side
Non-polished side
2) Inside baffle (stop scatter purpose)
10° pitched parallelogram
polished side inside
large aperture
same bracketry as SLiC part A
Aperture of the inside baffle

## Slide 12
Panel opposite to squeezer input
Non-polished side
10° pitched parallelogram
polished side inside
same bracketry as SLiC part A
clearance for the cable

## Slide 13
The roof
Roof extends outside the OFI for better coverage
Wide slots
Polished side on inside
No pitch required
Extended coated bracket for better coverage

## Slide 14: OFI TFP Beam Dump
Alena A., Calum T., Corey A., Eddie S., Stephen A.

## Slide 15
Koji’s Optics Lab Measurements (LHO #39436)
Sheila’s Single Bounce Measurements (LHO #41229)
How much scattered power in SQZ path?
Where does this beam go?

## Slide 16
OFI TFP AR Ghost Beam – adding V-shaped beam dump
https://dcc.ligo.org/LIGO-G1800771
This beam is an artifact of Zemax model – not real

## Slide 18
TFP Front Side Reflection
TFP AR Ghost Beam
Cartoon Beam Dump. Location TBD.
OFI TFP AR Ghost Beam – adding V-shaped beam dump

## Slide 19
OFI TFP AR Ghost Beam – adding V-shaped beam dump
SYS black glass beam dump
  Based on D1400222
For HAM5 beam heights
  L1 (9.11”)
  H1 (8.35”)
Flexible positioning in front of ZM2
All dirty parts will be in CIT’s hands by end of this week
Related to Ticket 10407
Representation of design, not in final location.
Pending: FEA on beam dump and possibility of adding damping (viton)