## Slide 1: The LIGO 4km Vacuum System: Design, Construction, Features & Faults
Mike Zucker, LIGO Caltech & MIT
NSF Workshop on Large UHV Systems for Frontier Scientific Research LIGO Livingston Observatory
29-31 January, 2019
LIGO-G1900137
Image Credit: Aurore Simmonet/SSU

## Slide 2: 30 years ago…

## Slide 3
Vacuum Equipment vs. Beamtubes
VE construction
Tube construction & Bakeout
Performance
Gotchas & Lessons Learned
Topics

## Slide 4
Vacuum Equipment
Chambers, pumps, instruments-
“Vacuum Equipment”
  Houses detector apparatus
  Isolation (valves), access (doors)
  Electrical, mechanical, optical penetrations/interfaces
  Pumping & instrumentation
  Somewhat “conventional”
  F:A ~ 10-2 ls-1cm-2
Beam tubes
  Just a long hole in the air;
      Never to be vented
  Highly “unconventional”
    10 million liters (per site)
    300 million cm2 (per site)
    200 l/s char. conductance
  F:A ~ 10-5 ls-1cm-2
LIGO is Really Two Vacuum Systems (at each site)

## Slide 5
LIGO Vacuum Equipment, LHO Corner Station

## Slide 6: Modular Vacuum Equipment Design
50 m
Corner station (Hanford)
50 m

## Slide 7
BSC chamber
(Boring Symmetric Chamber)
2.8m Ø x 5.5m h
Upper third removable dome
Thin (10-15mm) 304L SS shell with welded stiffeners, F&D heads
Combination of GTAW and plasma welding
Ports < 35cm Ø: ConFlatTM
Ports > 35cm Ø: Dual O-ring
  Treated Viton elastomer
  Isolated pumped annulus between inner and outer seal
  Permeation and damage tolerant

## Slide 8: BSC Equipment Installation

## Slide 9
Livingston ETM Y BSC “Cartridge” Installation
21 Feb ‘14

## Slide 11
HAM chamber
(Horizontal -Axis Module)
House complex input/output optics
2.1m Ø x 2m w
More than 70% of area is removable access doors

## Slide 12
L1 ISI 1/31/08
Notes: HAM ISI insertion in HAM6 at Livingston with guidance from Jeff Kissel (LSU, left) and Brian O’Reilly (LLO, right). March 2008.

## Slide 15: In Situ Chamber Bakeout

## Slide 16: Passive 80K cryotraps pump water vapor released by detector components
3D model
E. Sanchez
photo
G. Moreno

## Slide 17: Gate Valves
40” & 44” ID valves isolate beamtubes from instrumentation

## Slide 19: Pumping for Non-condensibles
Maglev turbos for initial evacuation only
Ion pumps assisted by NEGs in normal operation
NO rotating or vibrating machinery allowed during interferometer operation

## Slide 20
LIGO Beamtube
  9000 m3 volume/site
  30000 m2 area/site
  50 km of spiral welds
  ~10-9 torr
  budget ~ $40M (1997)
    $2500/m
    $50/lb

## Slide 22: Residual Gas Pressure Limits in Beam Tubes

## Slide 23: Beam Tubes
304L SS
3.2 mm thick with external stiffeners
Raw stock air baked 36h @ 455C
  Final JH2 < 1e-13 Tl/s/cm2
coil spiral-welded into 1.2m tube 16m long
  method adapted from sewer pipe industry
16m sections cleaned, leak checked
FTIR analysis to confirm HC-free
sections field butt-welded together in travelling clean room
Over 50 linear km of weld

## Slide 24: Depleting Hydrogen from raw SS before tube fabrication: An economical alternative to high T vacuum bakeout
SS sheet from mill is baked in air 36 hours at 455 ºC
(Hotter treatment deemed inadvisable due to carbide formation)
Total dissolved hydrogen is reduced ~ 3x
Remaining H is tightly bound, high activation T
Care is required in welding to avoid re-introduction of H
Sample Temperature (ºC)
H2 Flux (A.U.)
Raw SS
After
Airbake
data courtesy of Virgo

## Slide 25
Leak Test “Coffin”

## Slide 26
position
butt weld
Beamtube Field Assembly
leak check
field fitup
next section
transport

## Slide 27: Unbaked Water Outgassing (norm. to 1000 hours)
Saito et al (KAGRA, 2011)
ECB 304L, 200C conditioning bake
-40C dewpoint (127 ppm) re-exposure
J(H2O) ~ 2 x 10-13 T l s-1 cm-2 x (1000 h)/t
Li and Dylla (1993)
Electropolished 304L
10 ppm water content air re-exposure
J(H2O) ~ 4 x 10-12 T l s-1 cm-2 x (1000 h)/t
Weiss et al, T970111
LHO beamtube
J(H2O) ~ 8 x 10-12 T l s-1 cm-2 x (1000 h)/t
Weiss et al, T940090
BTD at CB&I
J(H2O) ~ 3 x 10-12 T l s-1 cm-2 x (1000 h)/t
(~ 10-16 T l s-1 cm-2 post-bake)
Tolerable pressure for H2O ~ 1/10 that for H2
Passive 1/t desorption with time too weak
Low-temperature bakeout was required
  LIGO used 1-shot bakeout to save cost
  Tubes can never be re-exposed to atmosphere

## Slide 28: I2R Bakeout to Desorb Water
IDC = 2,000 A
3 weeks @ 160ºC
Final JH20 < 2e-17 Tl/s/cm2

## Slide 29: Bakeout Results (Hanford)

## Slide 30
Sounds great- but
WHAT WOULD WE DO DIFFERENTLY?
Photo credit: Hurricane Harvey and Richard Oram

## Slide 31: 20:20 Hindsight– What Worked Well for LIGO
Low-hydrogen steel airbake process
Spiral-welded tube construction (with aggressive QA)
I2R bakeout heating
Modular chamber & spool concept
Soft-wall modular cleanrooms for installation
Dual differentially-pumped O-rings on large flanges
Passive LN2 cryotraps
Ion pumps
Two highly proficient and cooperative subcontractors
  Chicago Bridge and Iron (beamtubes)
  Process Systems International (vacuum equipment)

## Slide 32: 20:20 Hindsight— Not so much…
Large gate valves
  Have had two develop bizarre leakage through stem
  Serious and increasing concerns about fragility of the mechanism
Microbial-induced corrosion
  Leaks caused by SS-eating bacteria in Louisiana
  Tube environment was not climate-controlled
  Vermin and moisture “disease vectors”
But our biggest regret of all is…

## Slide 33: 20:20 Hindsight– If I could turn back time…
Budget constraints left us no standing provision to vent, re-evacuate and re-bake beamtubes in case of future contingency

## Slide 34
Thanks for your attention

## Slide 35
REFERENCE SLIDES

## Slide 36: I2R Bakeout to Desorb Water (cont’d)

## Slide 37: “Dubinin-Radeschevich Isotherm” desorption model
Weiss et al, LIGO-T970111

## Slide 39: End Station Pressure Evolution after Backfill

## Slide 40: Effects of Pump Distribution on Required J(H2O)
Consider 4km aLIGO arm (1.2m ID) with 2 or 15 distributed pumps + cryotraps at ends
Assign Sh1/2(H20) < 0.1 x min aL curve1
Sh1/2 (H20) < 2.6 x 10-25 Hz-1/2
Then2, for 2 pumps
J(H2O)2p < 2.0 x 10-15 T l s-1 cm-2
and for 15 pumps
J(H2O)15p < 4.3 x 10-15 T l s-1 cm-2
(no leaks, just autogenous water; Fi = 534 l/s)
1 T1400177, Fig. 3, “BNS Optimized 125W”
2 T1500202
2 pumps
15 pumps
NB: 1 T l s-1 cm-2 = 1.33 x 103 W m-2