## Slide 1: LIGO Laser Power Calibration Standards
Richard Savage
for
Yannick (Niko) Lecoeuche
LIGO Hanford Observatory
March 15, 2019
GW Metrology Workshop , NIST, Boulder, CO

## Slide 2: Power Calibration Standards
Transfer power calibration from NIST to GW interferometers
  Establish relative (and absolute) calibration of detectors in network.

## Slide 3: Earlier Power Sensor Configuration
Two spacers with apertures between integrating sphere and photodetector.
  Less sensitive to radial position.
  Increased laser speckle amplitude.
  Less robust mechanically

## Slide 4: Mechanical Robustness
Lateral dependence
  1 % per 16 μm offset
Radial dependence
  No spacers: 4 %/mm
  One spacer: 1 %/mm
Measurement with
M. Szczepanczyk

## Slide 5
Laser Speckle
Increases with decreasing viewing area
  Less detector spacers better

## Slide 6
Photodetector mounting
Monolithic photodetector housing front plate
Two-piece assembly draws PD assembly against sphere port.

## Slide 7: Pneumatic Sliders for Automated Measurements
(Thanks M. Spidell)

## Slide 8: Working Standard Calibration Procedure
WS/GS calibrations use:
  Simultaneous measurements eliminate laser power variations
  Swapping positions eliminates beamsplitter variations

## Slide 9
Responsivity ratio measurements
100-second samples; 1000 measurements over weekend

## Slide 10: Temperature Dependence
Temp. dependence ~ 0.07 % / ºC.
Heat one WS to 35 ºC in oven
Measure resp. ratio as temp. relaxes

## Slide 11: Gold Standard Stability Before/After NIST Calibration
Finally,

## Slide 12: EXTRA SLIDES

## Slide 13: Gold Standard Stability Before/After NIST Calibration

## Slide 14: Temperature Dependence of Photodiodes
Installed temperature sensors by photodiodes to investigate trends
Thermalized WS at 35ºC, measured against WS at ~20ºC
Ratio changes by ~0.07%/ºC difference in WS temperature.

## Slide 15: Responsivity ratio measurements
Believed to have stable enough measurements to see laser speckle in timeseries ratios