|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **APPROVALS** | **DATE** | **REV** | **DCN NO.** | **CHECK** |
| **AUTHOR: K. Strain** | **01-17-2020** | **1** |  |  |
|  |  |  |  |  |
|  |  |  |  |  |
| **APPROVED:** |  |  |  |  |
| **DCC RELEASE** |  |  |  |  |

# Description

A list of the coating specifications for the A+ Output Mode Cleaner

# General Specifications

Wavelength: 1064nm

Polarization: P

Coating Scatter: < 5 ppm

Type: low absorption, ion beam sputtered deposition

# Specific Coatings

**Coating A:** Input/Output coupler

**Side 1**

HR T = 8300 ±800 ppm @ 4 degrees AOI (best effort for ±400ppm)

**Side 2**

AR R < 500ppm, best effort < 100 ppm @ 4 degrees AOI

**Coating B:** Beam splitter

**Side 1**

50/50 T = 50 ± 2% @ 45 degrees AOI

**Side 2**

AR R < 500ppm @ 45 degrees AOI

**Coating C:** High reflector

**Side 1**

HR T = 10 +2/5ppm @ 4 degrees AOI

**Side 2**

AR R < 500ppm, best effort < 100 ppm @ 4 degrees AOI

# Metrology

Coating vendor to provide:

1. Two 1” witness samples from each coating run

2. Spectrophotometer graphs of the reflectance and transmittance of the HR

3. Spectrophotometer graphs of the reflectance of the AR coating