## FC expansion requirements
FCT Layout parameters:
| 1 | -- | # of bellows in each section between anchor supports
| 3 | -- | maximum # 20 ft sections between bellows
| 60 | ft | maximum length of pipe between bellows
| 2 | -- | maximum # of 6-Way Cross elements between bellows
| 14 | in | length of 6-Way Cross elements between 10" CF flanges
| 748 | in | maximum length of pipe & crosses between bellows
| 304 | -- | pipe & cross material
| 0 | C | assembly temperature & bellows zero-strain temperature
| 9.9e-06 | in/in/F | CTE (coefficient of thermal expansion)
| 1.73e-05 | 1/C | CTE (coefficient of thermal expansion)
Cold Temperature Extreme
| -30 | C | minimum recorded temperature @Richland, WA
| -10 | C | minimum recorded temperature @Baton Rouge, LA
| 0.38821 | in | bellows extension due to cold excursion
Diurnal maximum temperature:
| 48 | C | Max @LLO in BTE
| 55 | C | Estimate of max @LHO (desert)
| -0.71172 | in | diurnal maximum induced bellows compression between anchors
Bake-Out (< 10 cycles):
| 200 | C | maximum bakeout temperature
| -2.5881 | in | bake-out induced bellows compression between anchors
Fabrication & Assembly Tolerances:
Longitudinal:
| 0.188 | in | + tolerance on flanged pipe length (D1900443)
| 0 | in | - tolerance on flanged pipe length (D1900443)
| 0.03 | in | + tolerance on 6-Way Cross (D1900485)
| -0.03 | in | - tolerance on 6-Way Cross (D1900485)
| 0.03 | in | + tolerance on Bellows length (D# TBD)
| -0.03 | in | - tolerance on Bellows length (D# TBD)
| 0.11811 | in | + tolerance on Anchor Support axial position (+/- 3mm monument survey tolerance per T1000230, section 5)
| -0.11811 | in | - tolerance on Anchor Support axial position (+/- 3mm monument survey tolerance per T1000230, section 5)
| 0.4 | in | + intentional bellows extension during assembly
| -0.77211 | in | Bellows initial compression due to tolerances
| 0.60811 | in | Bellows initial extension due to tolerances
Angular (bending):
| 1 | mrad | +/- angular (parallel) tolerance on flanged pipe (D1900443)
| 1 | mrad | +/- angular (parallel) tolerance on 6-Way Cross (D1900485)
| 5 | mrad | +/- total angular (bending) tolerance
Lateral:
| 0.0625 | in | +/- centering tolerance on each flanged pipe (D1900443)
| 0.0625 | in | +/- centering tolerance on each 6-Way Cross (D1900485)
| 0.11811 | in | +/- tolerance on Anchor Support lateral position (+/- 3mm monument survey tolerance per T1000230, section 5)
| 0.4 | in | +/- lateral error due to angular tolerances if the tube sections are bent to compensate
| 0.83061 | in | +/- Bellows lateral deflection
Combined Bellows Movements:
| -3.3602 | in | Xc, maximum axial compression - very few cycles (bake-out)
| -1.4838 | in | Xc, maximum axial compression - diurnal cycles
| 0.99632 | in | Xe, maximum axial extension
| 0.83061 | in | Y, maximum lateral deflection
| 5 | mrad | θ, rotation (bending) |  |  | mrad | lateral (in) | lateral (in) w/ bending correction
|  |  |  | pipe 1 | end 1 | 1 | 0
[5 empty] | end 2 | 1 | 0.24 | 0.08
|  |  |  | cross 1 | end 1 | 2 | 0.24
[5 empty] | end 2 | 2 | 0.24
|  |  |  | pipe 2 | end 1 | 3 | 0.24
[5 empty] | end 2 | 3 | 0.96 | 0.24
|  |  |  | cross 2 | end 1 | 4 | 0.96
[5 empty] | end 2 | 4 | 0.96
|  |  |  | pipe 3 | end 1 | 5 | 0.96
[5 empty] | end 2 | 5 | 2.16 | 0.4

## A+ Bellows
Chandra Romel's email 11/19/2019:
About the bellows, the diameter of convolutions is 11.3” so we will need to specify a neck to that is at least as long as the stud used to accommodate 0.5” GV blind hole + 1.12” CF thickness + 0.5” for two nuts (in case we have to double nut when studs get stuck) = 2”. Total length of bellows (flange to flange) should be 20” and here is the break down in order from one end to other:
Flange (rotatable): 1.12”
Neck: 2” (distance between back side of flange and first convolution)
Convolutions: 15.26” live length (for +/-1.5” stroke)
Neck: ½”
Flange (non-rotatable): 1.12”
| 20 | in | total length
MDC nearest equivalent (without CF flanges):
Note that the axial deflection for a 6" active (convolution) length is +1.50/-.60
Overall length for an MDC 4700224 with CF flanges would be ~10.5 in
Huntington Vacuum Nearest:
-3.36
0.596
0.33333
-10.08
1.788
Hyspan nearest equivalent (with Flanges, but not CF flanges):
type 2509
10 NPS, 14 in overall length
| 5.31 in compression |  |  | https://www.hyspan.com/2500RoundedEj.html
2.66 in extension
Nor-Cal nearest equivalent:
8" ID (10 CF), not the 10" ID (12 CF) that we seek
max. compression is 7.00 - 6.2 = 1.8"
HighVac Co. nearest equivalent:
https://highvacco.com/products/cf-flanges-fittings/cf-flanges-adaptors/cf-flanges-to-stainless-bellows/
CF1325-F has 3 in compression range with a overall length (free) of 14 in

## Guide Spacing
10 | in | Dp | pipe outer diameter (per D1900443-v5)
40 | in | 4 Dp | G1
3.3333 | ft | 4 Dp | G1
140 | in | 14 Dp | G2
11.6667 | ft | 14 Dp | G2
0.25 | in | t | pipe wall thickness (per D1900443-v5)
9.5 | in | Di | pipe inner diameter
28000000 | psi | Ep | elastic modulus of pipe material (304L per D1900443-v5)
91.054 | in^4 | Ip
-14.7 | psig | Pd | pressure (inside minus outside)
88.8311 | in^2 | Ae | bellows effective area, corresponding to the mean diameter of the convolutions, Dm
1567.8599 | lb/in | fi | theoretical, initial, axial elastic spring rate per convolution
0.21001 | in | exc | equivalent axial compression per convolution
NA | in | Lmax | maximum guide spacing per eqn. B-9
NA | ft | Lmax | maximum guide spacing per eqn. B-9
364.5225 | ft | Lmax-v | max guide length, vented while baking

## EJMA unreinforced circular
Standards of the Expansion Joint Manufacturers Association (EJMA), 6th ed.
input parameters
DESIGN PARAMETERS
TEMPERATURE | 68 | F | T | BELLOWS MATERIAL | A240-304 |  |  | BAND MATERIAL | N/A
PRESSURE | 15 | psig | P | BELLOWS MODULUS | 29025328.9627 | psi | Eb | BAND MODULUS | N/A | psi
INSIDE DIA. | 10 | IN | Db | OUTSIDE DIA. | 11.25 | IN |  | CON. HGT | 0.625 | IN | w
NO. OF PLIES | 1 |  | n | BELLOWS THK. | 0.01 | IN | t | BAND THK. | N/A | IN
NO. OF CONS | 16 |  | N | PITCH | 0.75 | IN/CONS. | q | MAX OPE. PITCH | 0.81227 | IN | qo
TANGENT LENGTH | 1.25 | IN | Lt | TUBE LENGTH | 14.5 | IN |  | BELLOWS CONV. LENGTH | 12 | IN | Lb
OVERALL LENGTH | 16.74 | IN | L | FLANGE THICKNESS | 1.12 | IN
Note: Modulus, Eb(T) is temperature dependent; Change T for Eb
CONSTANTS |  |  |  | A+ Filter Cavity Tube, Bellows (D# pending)
Dm, Mean diameter of bellows convolutions
Dm = Db + w + n t
Dm | 10.635 | IN
tp, Bellows material thickness for one ply, corrected for thinning during forming
tp = t sqrt(Db/Dm)
tp | 0.0096969 | IN
Cp, Factor used in specific design calculations to relate U-shaped bellows convolution segment behavior to a simple strip beam
Cp is read from Figure C24 based on the following two constants:
q/(2 w) | 0.6 | --
q/(2.2 sqrt(Dm tp)) | 1.0616 | --
Cp | 0.53 | --
Cf, Factor used in specific design calculations to relate U-shaped bellows convolution segment behavior to a simple strip beam
Cf is read from Figure C25 based on the above two constants
Cf | 1.25 | --
Cd, Factor used in specific design calculations to relate U-shaped bellows convolution segment behavior to a simple strip beam
Cd is read from Figure C26 based on the above two constants
Cd | 2 | --
k, a factor which considers the stiffening effect of the attachment weld and the end convolution on the pressure capacity of the bellows tangent
k = Min{ Lt / (1.5 sqrt(Db t)), 1 }
k | 1
EFF. AREA |  | IN^2
ANNEALED | NO
DESIGN MOVEMENTS
EXTENSION | 0.99632 | IN | xe | exe | 0.06227 | in | total, equiv., axial per conv.
COMPRESSION | 3.3602 | IN | xc | exc | 0.21001 | in | eqn C-1
LATERAL | 0.83061 | IN | y | ey | 0.1917 | in | eqn C-5
ANGULAR (bending) | 0.005 | radians | θ | eθ | 0.0016617 | in | eqn C-3
TORSION, f | 0 | radians | (one flange is rotatable to take out any torque during assembly)
COMBINED COMPRESSION | 0.40338 | IN | ec | per convolution
COMBINED EXTENSION | 0.1311 | IN | ee | per convolution
PRESSURE STRESSES
S1, Bellows Tangent Circumferential Membrane Stress Due to Pressure:
S1 = P (Db + nt)^2 Lt Eb k / {2 (n t Eb Lt (Db + n t) + tc k Ec Lc Dc)}
N.B.: 'c' subscrpts refer to reinforcing collar which we don't have in the LIGO bellows
S1 | 7507.5 | psi
S2, Bellows Circumferential Membrane Stress Due to Pressure:
S2 = P Dm / {2 n tp (0.571 + 2 w/q)}
N.B.: in C960015 Hyspan uses qo (pitch at maximum operating expansion) rather than nominal pitch, q, per EJMA eqn. C-22
S2 | 3898.5784 | psi
S3, Bellows Meridional Membrane Stress Due to Pressure
S3 = P w / (2 n tp)
S3 | 483.4038 | psi
S4, Bellows Meridional Bending Stress Due to Pressure
S4 = {P/(2 n)} {w/tp}^2 Cp
S4 | 16513.3301 | psi
Sab | 15700 | psi | allowable bellows material stress at design temperature. This value is for 304L @200F (per Larry Jones' note)
Cwb | 1 | -- | Longitudinal weld joint efficiency factor for the bellows
S1 < Cwb Sab | TRUE
S2 < Cwb Sab | TRUE
Cm | 3 | -- | Material strength factor, 1.5 for bellows in annealed condition, 3 for bellows in the as-formed (cold worked condition)
S3 < Cm Sab | TRUE
S4 < Cm Sab | TRUE
DEFLECTION STRESSES
S5, Bellows Meridional Membrane Stress Due to Deflection
S5 = (Eb tp^2 e) / (2 w^3 Cf)
S5 | 1803.7301 | psi
S6, Bellows Meridional Bending Stress Due to Deflection
S6 = (5 Eb tp e) / (3 w^2 Cd)
S6 | 242202.7494 | psi
Mt = p G n t Db^3 f/(4 Ld N) |  | eqn C-48
G | 10600000 | psi | for 304L, shear modulus (per Larry Jones' note)
Ld | 1.6783 | in-lb | developed length of one convolution, Ld = 0.571*q+2*w (section C-4.2.4)
Mt | 0 | in-lb
firwb | 3100409.851 | in-lb/rad | theoretical torsional spring rate for the bellows (not defined in the EJMA guidelines) | [5 empty] | 350299.1497 | N-m/rad
Ss = 2 Mt/(n t p Db^2) |  | eqn C-47, shear stress due to torsion
Ss | 0 | psi | Ss < 0.25 Sab per eqn C-47
Sab | 15700 | psi | allowable bellows material stress at design temperature. This value is for 304L @200F (per Larry Jones' note)
Sab/4 | 3925 | psi
allowable torsion, fab | 0.0019886 | radians | 0.11394 | degrees
SPRING RATE & Axial Force
fiu, Bellows Theoretical Axial Elastic Spring Rate
fiu = 1.7 Dm Eb tp^3 n / (w^3 Cf)
fiu | 1567.8599 | lb/in
fiub | 97.9912 | lb/in | theoretical axial spring rate for the bellows | [5 empty] | 17160.8997 | N/m
Fa | 329.2692 | lb
Lateral Force & Moment & Stresses
y | 0.83061 | in | imposed lateral movement
ey | 0.1917 | in | axial movement per convolution due to imposed lateral deflection y (Appendix C)
Vl | 184.9874 | lb | Lateral force at the ends resulting from imposed lateral deflection y
Ml | 799.1279 | in-lb | Moment at ends resulting from imposed lateral deflection y
fivb | 222.7128 | lb/in | theoretical lateral spring rate for the bellows (not defined in EJMA guidelines) | [5 empty] | 39002.9935 | N/m
S5, Bellows Meridional Membrane Stress Due to Lateral Deflection only
S5l = (Eb tp^2 e) / (2 w^3 Cf)
S5l | 857.2177 | psi
S6l, Bellows Meridional Bending Stress Due to lateral Deflection only
S6l = (5 Eb tp e) / (3 w^2 Cd)
S6l | 115106.1776 | psi
N.B.: According to the Standards of the Expansion Joint Manufacturers Association (EJMA), 6th ed., section C-4-1.4, "Deflection Stress":
"Typical calculated stress range values are 50,000 to 500,000 psi. These values are not true stresses, since they exceed the elastic limit of the material. They are meaningful when correlated with actual test results in evaluating fatigue life."
FATIGUE LIFE
Nc, Fatigue life, number of cycles to failure
Nc = {c / (St - b)}^a
St = 0.7 (S3+S4) + (S5+S6)
St | 255904.1931 | psi
From Figure C28:
a | 3.4 | --
b | 54000 | psi
c | 1860000 | psi
Nc | 1900.4212 | cycles

## Eb vs T
304 Stainless Steel, Elastic Modulus (Eb) vs Temperature
http://www.bssa.org.uk/topics.php?article=139
T (°C) | E (GPa) | Equation
20 | 200 | 200.1226
100 | 194 | 193.8201
200 | 186 | 186.2203
300 | 179 | 178.9185
400 | 172 | 171.903
500 | 165 | 165.1626
68 | T (°F)
20 | T (°C)
200.1226 | E (GPa)
29025328.9627 | E (psi)
Compare to 2.70E+07 psi used by Hyspan in C960015